## Slide 1
Status report
Design of a Multi Chip Module
for a general purpose readout electronics of MPGD:s
5.4.2017
Leif Jönsson
Lund Univeristy
(on behalf of the Lund group)

## Slide 2
The goal of the full project: Construct a TPC with excellent position resolution and two-track separation.
This requires: Small readout pads (<= 1x6 mm2).
The channel density of the readout electronics has to match the small pad size.
 The ’proof-of-princple’ has been demonstrated and the project is now in the ’feasibilty phase’, the aim
of which is to produce front-end electronics comparable in size to the readout pads.
 This readout electronics can be used in smaller MPGD-detectors for applications in various other fields like
medical diagnostics , material science at XFEL, and for investigations at ESS.
 The readout electronics is based on the SALTRO16 ASIC, which integrates the analogue and digital
processing in the same chip.
 The size of the die itself is 8.7x6.2 mm2, which corresponds to a channel occupancy of 3.37 mm2/channel.
 As this project was started it was not possible to find a company which could package the SALTRO ASIC
in a small enough capsule to satisfy our requirement.
We decided to go for a ’chip-on-board’ solution.
We developed a carrier board onto which the SALTRO ASIC was bonded and covered with an epoxy
glob to protect the ASIC and the bonding wires. The size of the carrier board is 12x8.9 mm2. Eight of
these carrier boards should be mounted on a Multi-Chip-Module (MCM) board, which also contained
a CPLD, connectors and passive components. This increases the channel occupancy to 6.4 mm2/channel

## Slide 3
Outcome: After some initial problems had been sorted out, the Swedish company managed to bond
one ASIC on a Carrier board, where all connections were verified, also after application of the epoxy
glob and the soldering balls.
However, as the carrier board had been soldered onto an adaptor board for characterization, some
areas had lost connection.
X-ray investigations showed the following:
Normal bonding loop
Bond wire without a significant loop

## Slide 4
Cracks in the epoxy (red arrows
in the right hand side picture)
and air bubbles (green arrows
in the right hand side picture).
Conclusion:
 There is nothing indicating an error in the design of the carrier board.
 The lost connections are due to the different expansions of the PCB and the epoxy material in the
soldering process, which causes breaks in those wires where the wire loop is not sufficiently high.
 Another disadvantage of this method is that the surface of the epoxy layer is not flat enough to
provide efficient cooling.

## Slide 5
 Through contacts with CERN we were directed to three companies , which are able to offer packaging in
small enough capsules. After intense communication with these companies and careful evaluation we at the
end signed a contract with NovaPack in Grenoble for a pre-serie of 25-30 chips. This happened on the
16th of February 2017.
 The chips from the pre-serie will be used for characterization before the production order. The total
sample is about 840 SALTRO ASIC:s.
 NovaPack has started the substrate design and this is expected to be completed by the end of week 15
(April 10-14).
 After that they will submit the CAD file for review by us.
 As we have approved the design the lead time will be about 6-7 weeks to get the substrates plus bout
2 weeks to deliver the initial packages.
In the middle of June we can expect the first parts.
Characterization of the chips can start

## Slide 6
 In the mean time we have started to prepare the test set-up.
 The MCM-board is a very dense board and not ideal for characterization of the chips.
 We have therefore produced a Development Board, in a size suitable for lab work.
 It is a stand alone board, containing only one SALTRO-chip in TQFP package and a CPLD for control of
the data handling but also the necessary voltage regulators.
 Several test points and connectors allow for connection to a logic analyzer.
 Data will be read out by serial readout over a 40 MHz DTC link connecting the MCM board
(alternatively the Development Board) and an SRU (Scalable Readout Unit). Data will be processed by a PC
 The Development Board has been used to test the FPGA firmware and to establish the communication
between the SALTRO16-chip in the TQFP package and the SRU.
 The firmware of the CPLD (on the Development Board) is taken from the ALICE experiment
and the firmware for the SRU is using the ALICE and TOTEM systems.
A master student is working on the modification and further development to meet our requirements.
 The PC DAQ is based on the EUDET TPC ALTRO DAQ, modified to handle Ethernet/UDP.

## Slide 7
The Development Board now has to be modified to enable test of the ASIC:s packaged in capsules
of size 12x9 mm2 and to adapt to the substrate design.
 Single Saltro16 readout in a BGA ZIF (zero insertion force) test-socket
 Same CPLD & connections as in final circuit
 Same power distribution
 Test-points on all signals
 8-layer board 210x145mm
 No impedance control, no signal timing/length matching
 HDI design because of BGA footprint (buried 0.1mm vias)
 3 signal layers, 5 power layers

## Slide 8
Saltro16 BGA test
socket
CPLD
(Lattice MachXO2)
Local power regulators
Power input
Readout
Testpulse
Signal input
Debug port

## Slide 9: BGA package

## Slide 10: Saltro16 ballpad connections
 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 
A | PasaTestP | Shutdown | PasaTestN | AdcTestP | RefN | Vcm | CmOut | chipadd [1] | chipadd [3] | adc_add0 | l2yb | ack_en | tstout | A
B | TestMode | PreampMode | BiasDecay | AdcTestN | VddAdcA | RefP | BiasGate | chipadd [0] | chipadd [2] | adc_add1 | trgb | ackb | bd [00] | B
C | InCDM15 | VddPasaA | VddPasaA | VddAdcA | VddAdcA | GndAdcDig | VddAdcDig | VDD | VDD | tsm | bd [01] | bd [02] | bd [03] | C
D | InCDM14 | VddPasaA | VddPasaA | VddAdcA | GndAdcA | GndAdcDig | VddAdcDig | VDD | GND | DVDD | bd [04] | bd [05] | bd [06] | D
E | InCDM13 | VddPasaA | GndPasaA | VddAdcA | GndAdcA | GndAdcDig | VddAdcDig | GND | DVSS | DVDD | bd [07] | bd [08] | bd [09] | E
F | InCDM12 | VddPasaA | GndPasaA | VddAdcA | GndAdcA | GndAdcDig | VddAdcDig | GND | DVSS | DVDD | bd [10] | bd [11] | bd [12] | F
G | InCDM11 | VddPasaA | GndPasaA | GndPasaA | GndAdcA | GndAdcDig | VddAdcDig | DVSS | DVSS | DVDD | scan ena | bd [13] | bd [14] | G
H | InCDM10 | VddPasaA | GndPasaA | GndPasaA | GndAdcA | GndAdcDig | VddAdcDig | DVSS | DVSS | DVDD | bd [15] | bd [16] | bd [17] | H
J | InCDM9 | VddPasaA | GndPasaA | GndPasaA | GndChipGuardRing | GndAdcDig | VddAdcDig | DVSS | DVSS | DVDD | bd [18] | bd [19] | rdoclk | J
K | InCDM8 | VddPasaA | GndPasaA | GndPasaA | GndChipGuardRing | GndAdcDig | VddAdcDig | DVSS | DVSS | DVDD | bd [20] | bd [21] | bd [22] | K
L | InCDM7 | VddPasaA | GndPasaA | GndPasaA | GndAdcA | GndAdcDig | VddAdcDig | DVSS | DVSS | DVDD | bd [23] | bd [24] | bd [25] | L
M | InCDM6 | VddPasaA | GndPasaA | GndPasaA | GndAdcA | GndAdcDig | VddAdcDig | DVSS | DVSS | DVDD | ScanMode | bd [26] | bd [27] | M
N | InCDM5 | VddPasaA | GndPasaA | GndPasaA | GndAdcA | GndAdcDig | VddAdcDig | GND | DVSS | DVDD | bd [28] | bd [29] | bd [30] | N
P | InCDM4 | VddPasaA | GndPasaA | VddAdcA | GndAdcA | GndAdcDig | VddAdcDig | GND | DVSS | DVDD | bd [31] | bd [32] | bd [33] | P
R | InCDM3 | VddPasaA | VddPasaA | VddAdcA | GndAdcA | GndAdcDig | VddAdcDig | VDD | GND | DVDD | bd [34] | bd [35] | bd [36] | R
T | InCDM2 | VddPasaA | VddPasaA | VddAdcA | VddAdcA | GndAdcDig | ClkSelect | VDD | VDD | SoftRst | bd [37] | bd [38] | bd [39] | T
U | InCDM1 | Gain1 | ShapingTime1 | ShapingTime3 | Vcm | CmOut | ClkAux | chipadd [4] | chipadd [5] | cstbb | grstb | writeb | dolo_en | U
V | InCDM0 | Gain2 | ShapingTime2 | Polarity | RefN | RefP | sclk | chipadd [6] | chipadd [7] | trsf_en | trsfb | errorb | dstb | V
 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 

## Slide 11
Status of the DAQ system:
The CPLD: The firmware is written and has been tested.
The SRU: For the firmware a lot is already done.
 we can read/write SALTRO and CPLD registers.
 we can send software tiggers to the SALTRO.
 we can read partial event data.
Still to do: The memory management on the SRU to read full events still has to be optimized.
A hardware trigger to the SRU:s has to be implemented.
The detector control between the MCM and SRU is ready and has been tested.
The readout to the PC is partly implemented.
Still to do is:  define data format
 test the full readout of the Development Board, the max data rate and the trigger rate
The PC DAQ: The modification of the ALTRO DAQ is in progress.
The detector control between the SRU and PC is mostly done and partly tested.
Still to do:  Some intermittent problems on the SRU still have to be solved
 Everything has to be tested

## Slide 12
Readout of the Development Board:
 The simple stand alone data reciever is ready.
 The implementation into the DAQ system has to be done.
Trigger handling:
 The trigger handling has to be implemented.
Monitor and run control:
 This is the same as in the ALTRO DAQ.
 The slightly different data format has to be implemented.

## Slide 13
SRU
Development Board
SALTRO
CPLD
The development set-up

## Slide 14
Event with two channels fully read (0 and 1) from the Development Board, the peak is induced by a
generated pulse from the SRU.
If reading all channels only 180 samples can currently be read, one must rearrange the memory
management on the SRU for full readout (to be done).
The standalone data receiver was used (to be implemented in PC DAQ).
Number of samples
ADC-value

## Slide 15
Towards the final system
The CPLD has to be tested with 8 chips on the MCM board.
For the SRU the following things have to be done:
 Implement handling of up to 40 MCM boards.
 Implement a memory manager for reading full events with up to 1000 samples.
 Implement resending of lost Ethernet/UDP data packets.
 Implement power pulsing.
 Implement clock synchronization between multiple SRU:s.
 Implement the trigger system.
For the PC DAQ the following things have to be done:
 Implement handling of multiple SRU:s.
 Implement handling of 40 MCM boards per SRU.
 Implement an interface to a common DAQ.
 Implement the trigger handling.

## Slide 16
Summary
 The packaging of the SALTRO16 ASIC:s is underway. Delivery of the preseries is expected mid June.
 The design of the MCM Development Board is ready and will be scrutinized before the PCB is ordered.
 The Test Socket is ordered and delivery is expected on April 25.
A first version of the DAQ system enables readout of data but further development will be done to
improve it