## Slide 1: Safety regulation for pressure equipment and crab cavity project in the framework of High Luminosity LHC Project
Luca Dassa on behalf of the engineering team
International Review of the Crab Cavity Performance for HiLumi on 3-5 April 2017 (CERN)

## Slide 2
SPS approach (with details)
Proposal for LHC installation

## Slide 3
SPS approach (with details)

## Slide 4: The cryomodule (1)
SPS cryomodule
Pressure components:
Dressed cavities
HOMs
Cryogenic lines
Thermal shield
Courtesy of: http://www.eucertification.com/ce-certification-2/ped/pressure-vessel-calculator
CERN:
SR-M
GSI-M-2
GSI-M-4
https://espace.cern.ch/safety-rules-regulations/en/rules/
PED 97/23/EC -> EC declaration of conformity
PS = 0.8 barg
V= 200 l
Gas group 2
Courtesy of: TUV

## Slide 5: The cryomodule (2)
Crab SPS prototype cryomodule
According to PED Annex 2, the cryomodule belongs to risk category I
equipment liable to have major Safety implications (GSI-M-4)
exempted from EC-marking
the equipment shall meet the Essential Safety Requirements (ESRs) stated PED 97/23/EC.
EU harmonized standards used whenever possible
If not possible, ASME Section VIII Div. 2 + compensatory measures in view of compliance with the ESRs of the PED.
Hydrostatic proof test will be replaced by alternative methods
Review and discussion with HSE unit (EDMS 1494776 + EDMS 1541969)
Everything discussed and approved by 14/09/2015
Valid only for SPS cryomodule
designers
engineers
HSE unit
management
manufacturing experts
materials not considered in harmonised standards
proof test with high risk to impact on RF performances

## Slide 7: Pressure bearing components (1)
Vacuum vessel
Thermal shield
Cryogenic lines
Standard approach for these components (not treated in following slides)
See later for proof test
Connection between titanium vessel and stainless steel cryogenic line using transition obtained by explosion bonding

## Slide 8: Pressure bearing components (2)
FPC
HOM
HOM
Leak tightness: welded joint (not pressure bearing joints!)
Edge welded-bellows (5x)
Welded pressure vessels: actively cooled HOMs
Bolted pressure vessel: dressed cavity

## Slide 9: Material tests
Materials:
Ti Gr. 2, Ti Gr. 5, high purity Nb and 55Ti-45Nb not considered in the Harmonised Standards
conformity with the ESRs required
compliance with PED would require PMA
  For Ti -> PMA (DESY) + material certificate + internal tests (EDMS 1538192)
  For Nb and NbTi -> PMA on similar material (DESY XFEL) -> material certificate + internal tests (EDMS 1722302 for Nb and EDMS 1493400 for NbTi )

## Slide 10: Design
Loads:
  PS = 1.8 bar abs
  T= 300 K
  Static pre-tuning = 0.2 mm
  Loads not considered:
  Dynamic fine tuning (only at cold)
Other remarks
  Fatigue life: not applicable (cycles < 500)
  Pressure test: derogation
Cavity, bolts, weld assessment -> EDMS 1549819
HOM assessment -> EDMS 1433086
Vessel assessment -> EDMS 1712011
Biphase line assessment -> EDMS 1727787

## Slide 11: Preliminary tests on components
Bolted He vessel = unusual pressure vessel
Courtesy of C. Zanoni, 11/11/2015 - SPS Cryo-module Engineering Review
Report EDMS 1705731
Report EDMS 1721359

## Slide 12: Manufacturing & Inspection (1)
Vacuum brazing (not for pressure equipment), electron-beam welding, TIG (GTAW) welding
Nb, NbTi, Ti
complex joints / not standard (not only for pressure)
exotic materials
Ti welded joints are not bearing pressure loads!
Back to the EDMS 1494776

## Slide 13: Manufacturing & Inspection (2)
qualification of each weld is performed in-house: for Nb according to harmonized standard (EN15614-11)
Guidelines from other projects: i.e. HG (former SPL) cavities for Nb welded joints and from other labs
Specification (EDMS 1389669)
No standard available for Nb welded joints
Acceptance criteria on imperfections: part of standards related to Al alloy have been retained and modified
Additional requirements for acceptance criteria due to RF restrictions
Qualifications and the acceptance criteria according to harmonized standards (in compliance with the Essential Safety Requirements from PED) whenever possible
Nb
Nb
Ti
Not pressure bearing joints
Ti
welding qualification according with the harmonized standard EN15613, not wrt the mechanical properties, but wrt to the configuration in the helium tank
welder and the operator are already qualified by the welding qualification itself
Specification (EDMS 1389669)
List of applicable EU harmonized standard standards with equivalent American standards
Ti

## Slide 14: Manufacturing & Inspection (3)
Bare cavity
EDMS 1581039 CERN documentation for welding qualification and tests
EDMS 1685099 Welding Book Crab Cavity DQW. Extremities and Final assembly. : ASSETS HCACFCA004‐CR000001 & CR000002.
EDMS 1758810, WELDING DOCUMENT CRAB CAVITY DQW
Material traceability:
EDMS 1549318 Summary of materials for Crab Cavity and He tank manufacturing at CERN
Ti vessel:
welding book in progress

## Slide 15: Proof test
Results of preliminary analysis shows that there could be risk of plasticization in some critical area during proof test.
There are NO RISKS for people safety!
De-tuning of the cavities at ambient temperature must be avoided
The mechanical properties of the material used for manufacture of the bare cavities is being verified by dedicated mechanical tests.
Alternative method proposed: pneumatic proof test at the maximum allowable pressure (0.8 barg) at ambient temperature.
Done in SM18 with service box
Accepted by HSE CERN unit
(see EDMS 1494776 + EDMS 1541969)
First cool-down and proof test preformed in SM18

## Slide 16: Documentation
cryomodule components
design
And similar for other components…
CERN-0000115183
CERN-0000115182
only for bare cavity

## Slide 17: SPS Crab Cavity Test Stand: Safety review
Approach accepted
All recommendations are in EDMS 1706929
See the talk :
“SPS Crab Cavity Test Stand: Safety review summary, Christelle Gaignant, International Review of the Crab Cavity Performance for HiLumi – 3rd/5th April 2017”

## Slide 18
Proposal for LHC installation

## Slide 19: LHC crab cavities
main reference: EDMS 1494776 “A proposal for compliance…” (for SPS cryomodule) + EDMS 1541969 “Safety Request Form - Safety compliance…”
risk due to pressure -> Category definition -> He volume on the LHC cryomodule -> we assume that we have the same volume (reasonable assumption: each cryomodule protected by a safety valve) as per SPS cryomodule -> Category 1
If so, the same safety requirements as per SPS cryomodule apply -> no need for Notified Body
Approach not yet discussed by HSE unit.
Clarification will be requested to HSE unit according to the statement of EDMS 1698982 and based on evaluation done in EDMS 1494776
We don’t see any reason to have for LHC series production a different approach from the SPS cavities

## Slide 20: Specifications for LHC crab cavities
Based on the hypothesis done in previous slide:
  Specification for dressed cavities (document EDMS 1389669) available and already presented to US LARP
  CERN take care of the global «certification» of the machine
  no CE marking is required, no notified body
  Possibility to use ASME or EU standard. Big effort has been done to state equivalence between standard (also using ITER experience) -> see example:
  US LARP has been updated about it

## Slide 21: Companies involved in LHC crab cavities
Example for bare cavity (from Specifications EDMS 1389669)
No exotic requirements
CERN has performed all the steps that will be required to companies
CERN is available to help and to clarify the spec, if needed
Notified body not required

## Slide 22: Thank you…

## Slide 23
Back-up slides

## Slide 24: Material tests (2)
Explosion bonding (Nickel interlayer) -> EDMS 1705993
Explosion bonding (Copper interlayer) -> EDMS 1724598
Explosion bonding (Tantalum interlayer) -> EDMS 1739621
Tests
Leak
Traction
Metallographic
Ultrasonic
+ thermal cycles on the real component + leak test
Chart: ε [%]

## Slide 25: Test on welded joints
EDMS 1458356 Technical Report: Titanium Welding Test for Crab Cavities project.
EDMS 1562825 WELDING HELIUM TANK PROTOTYPE CRAB CAVITY‐DQW
EDMS 1716204 Metallographic Weld Analyses for Crab Cavity Helium Tank Prototype

## Slide 26: Cryogenic circuits and safety protection
Flex lines
individual safety device for each cryogenic device (cold box, VBx, transfer line, SB)
exception: CCM (by integration reasons protection on SB)
Safety devices according to ISO-4126:
Helium guarded safety valve at 1.8 bara, sized to cover insulation vacuum break with 1.5 contingency factor
Dedicated rupture disc at 2.1 bara, sized to cover beam vacuum break with 1.5 contingency factor
Two DN150 discs (since common vacuum, one on SB and one on CCM) set at 1.05 bara (contingency factor of 2 considered)
Courtesy of K. Brodzinski, Crab cavities SPS Test Stand Safety Review – CERN – 9 November 2016

## Slide 27: Certification
Back to the EDMS 1494776

## Slide 28: Full diagram of SPS installation
Courtesy of K. Brodzinski, Crab cavities SPS Test Stand Safety Review – CERN – 9 November 2016
Flex lines

## Slide 29: MCI – CCM internal lines
Assumption: complete break of a beam pipe (D=90 mm)
Heat load due to N2 condensation inside the cavity = 40 kW/m2
Safety devices according to ISO-4126:
Helium guarded safety valve at 1.8 bara with diam.=25 mm, sized to cover insulation vacuum break with 1.5 contingency factor
Dedicated rupture disc at 2.1 bara with diam.=25 mm, sized to cover beam vacuum break with 1.5 contingency factor
Courtesy of K. Brodzinski, Crab cavities SPS Test Stand Safety Review – CERN – 9 November 2016

## Slide 30: MCI for the vacuum jacket
Vacuum jacket MCI: beak of a tuner bellow (He release section of ~3000 mm2)
Two DN150 discs (since common vacuum, one on SB and one on CCM) set at 1.05 bara will be installed (sizing includes contingency factor of 2 considered for first calculation, detailed analysis is still to be completed)
Courtesy of K. Brodzinski, Crab cavities SPS Test Stand Safety Review – CERN – 9 November 2016

## Slide 31
Courtesy of C. Zanoni, 11/11/2015 - SPS Cryo-module Engineering Review

## Slide 32
Courtesy of C. Zanoni, 11/11/2015 - SPS Cryo-module Engineering Review

## Slide 33
Courtesy of C. Zanoni, 11/11/2015 - SPS Cryo-module Engineering Review

## Slide 34
Courtesy of C. Zanoni, 11/11/2015 - SPS Cryo-module Engineering Review

## Slide 35: LHC crab cavities (1)
EDMS 1698982, S. Baird, head of the HSE Unit: Conformity approach for Pressure Equipment for the High Luminosity LHC Project (June 10th, 2016)
SR-M, GSI-M-2, GSI-M-4 -> PED (2014/68/EU) -> Category? -> CE marking -> general approach
Approach by exception -> per equipment basis -> unique in-house technical competency -> no Notified Body (it would not add any risk mitigation) -> HSE-SEE will act as de facto Notified body
Responsibility for design, manufacture, testing and quality insurance remains with HL-LHC Project team / technical justification shall be prepared by the equipment owner-work package leader
Formal derogation not required if the HSE-SEE classes the equipment as being “liable to have major safety implications”
Technical requirements
Organizational and quality assurance requirements
Example: see WP3 and EDMS 1753780 (07/02/2017)
Is the LHC crab cryomodule concerned???

## Slide 36
EDMS 1698982, S. Baird, head of the HSE Unit: Conformity approach for Pressure Equipment for the High Luminosity LHC Project (June 10th, 2016)