## Slide 1
THGEM and MICROMEGAS HV issues:
The experience gained with the novel
COMPASS photon detectors
C. Azevedo, S. Dalla Torre, S. Dasgupta, M. Gregori, G. Hamar,
S. Levorato, G. Menon, F. Tessarotto

## Slide 2: NOVEL DETECTOR HV ISSUES
quick reminder
THGEM and HV performance
  Reminder: during the 2016 run, 3 THGEM sectors (over 16) with lower gain
  Total surface: 1.5 m2
MM and HV performance
  Reminder: during the 2016 run, about 20 pads (over 19040) exhibiting shorts

## Slide 3
Quick remninder

## Slide 4: DETECTOR ARCHITECTURE
PCB
0.07 mm
fiberglass
Single pad
HV is applied here through a resistor (mesh @ ground)
Signal read-out from this pad
HV
Signals
Connections for groups of 48 pads
Resistor arrays
0.5 GW
60 x 60 cm2 detectors
formed by 30 x 60 cm2
active elements
CsI coating t

## Slide 5
THGEM and HV performance

## Slide 6: THGEM and HV performance
THGEM HV distribution
  1 chamber: 2 hybrids (detectors)
  1 hybrid: 2 sectors
  1 sector: 6 segments
   1 segment is 4% of a chamber
  HV is distributed sector by sector via
  independent HV channels for
  THGEM1_top_sector_ij
  THGEM1_bottom_sector_ij
  THGEM2_top_sector_ij
  THGEM2_bottom_sector_ij
  resulting in 16 THGEM-HV ch.s per detector, namely the segments are grouped 6 by 6
a relevant feature:
  when both THGEM1 and THGEM2 in a sector are on, the sparks always propagates to both
during standard operation it is not possible to disentangle the following
    Was the spark originated in THGEM1 or THGEM2 ?
    In which segment was the spark generated ?
Under laying principle of the fixing procedure
    Identify the feeble segments and isolate them providing separate HV ch.s
    GOAL: from 19% (= 3 sectors) to 3-5%(= 3-5 segments) low efficiency surface
chamber
sector (1 HV ch.)
hybrid
segment

## Slide 7: THGEM and HV performance
First exercise to identify the feeble channels
  only the THGEM1 (2) sectors on at high HV, with THGEM2(1) at lower HV
  THGEM1’s (CsI coated) are performing better !
  Top chambers are electrically better (already observed during the run)
  By product: an approximate reference to go on testing in Ar-CO2 (no methane) to be able to continue during the winter shut-down period
bottom chambers
top chambers

## Slide 8: THGEM and HV performance
Studying segment by segment
2 dedicated voltage distribution boxes built
  6 segments per box, independent HV supply
Dedicated software control tool
time
1 step every 10’
dot size ~ N (n. of sparks)
Voltage decreased when N > 4
Voltage
THE TEST LOGICS
TIME at the given HV
Number of sparks
Spark rate (1/min)
ANALYSIS OF THE RESULTS

## Slide 9: THGEM and HV performance
Parametrizing the test results via a single figure
Extract it from the spark rate vs voltage
  3 algorithm used
    When they give the same indication
    (LARGE MAJORITY OF CASES), use it
    When they are at variance, repeat the measurement
Results OK
REPEAT THE MEASUREMENT

## Slide 10
Chart: PD2
| 1 | 2 | 3 | 4 | 5 | 6
 | 1250 | 1240 | 1215 | 1210 | 1200 | 1170
 | 1135 | 1135 | 1185 | 1225 | 1200 | 1290
 | 1260 | 1240 | 1250 | 1260 | 1250 | 1280
 | 1265 | 1260 | 1249 | 1235 | 1210 | 1280
 | 500 | 1500
 | 500 | 1500
 | 500 | 1500
Chart: PD1
| 1 | 2 | 3 | 4 | 5 | 6
 | 1200 | 1150 | 1150 | 1150 | 1140 | 1150
1210 1200 1190 1170 1160 1190 | 1210 | 1200 | 1190 | 1170 | 1160 | 1190
 | 1190 | 1140 | 1130 | 1150 | 1160 | 1165
 | 1170 | 1135 | 1140 | 1150 | 1130 | 1130
13 14 15 16 17 18 | 1175 | 1125 | 1130 | 1135 | 1140 | 1150
 | 500 | 1500
 | 500 | 1500
 | 500 | 1500
Feeble sectors, study of the bottom chambers completed (top chambers on going)
THGEM and HV performance
Chamber PD1
Chamber PD2
Repeated measurements
THGEM 1
Bad segment that matches with what
observed during the 2016 run
THGEM 2
Sector 0,0
Sector 0,1
Sector 1,0
Sector 1,1
Sector 0,0
Sector 0,1
Sector 1,0
Sector 1,1

## Slide 11
Summary of segments prepared for separate HV supply
in colour, the feeble segments as detected during the 2016 run
THGEM and HV performance

## Slide 12: NEW: TOP CHMBERs COMPLETED
THGEM 1
THGEM 2
PD5

## Slide 13: NEW: TOP CHMBERs COMPLETED
THGEM 1
THGEM 2
PD6

## Slide 14: THGEM SPARK RATE DOES NOT DEPENED ON BEAM INTENSITY
BEAM INTENSITY from ppp on T6 (AVERAGE per h)
Spark rate (h-1)
The beam intensity is not an issue !
  Using slow control data
  Correlations between spark rate and other parameters being studied

## Slide 15
MICROMEGAS and HV performance

## Slide 16: MICROMEGAS and HV
REMINDER

## Slide 17: MICROMEGAS and HV
PD1, PD2, PD6 – detailed study of the pads with shorts
  To be performed for PD5
Pads with shorts identified, fixed by applying a reverse bias when possible, connection removed to the others
In PD1:
5 pads with shorts (over 4760)
  1 fixed
  4 isolated
+ 1 pad isolated due to connector
In PD2:
5 pads with shorts (over 4760)
  2 fixed
  3 isolated
In PD6:
7 pads with shorts (over 4760)
  1 fixed
  6 isolated
PD2
Pads off at 0.1 ‰ level

## Slide 18: MICROMEGAS and HV
More leak current when humidity condensed onto the resistor arrays that supply the pads
N2 flow to overcoming similar episodes (and reduce parasitic noise)