## Slide 1: LINAC4 – Instrumentation Status
F.Roncarolo BI-TB 9-Feb-2017
Lots of material (and work!) from
J.C. Alica Santamaria, S.Burger, T.Hoffmann, E.Barrios Diaz, G.Guidobboni, A.Guerrero, M.Bozzolan, L.Soby, S.Bart Pedersen, J.Tan, etc ….
We miss Uli!!!
LINAC4 Bear
… ah no: LINAC4 beam!

## Slide 2: LINAC4 Status and Plans
All LINAC RF structures and magnets installed up to the main dump
Transfer line (L4T) equipment installed apart from zone occupied by HST (including de-buncher + some diagnostics)
8-10 mA reached at 160 MeV, main limit in LEBT emittance
Schedule:
  HST till end of March
  April-May : remove HST, install de-buncher, temporary dump in place
  June  LS2: reliability run (with YETS ~overlapped with rest of the complex)

## Slide 3: LINAC4 Layout

## Slide 4: Last part of Transfer Line
Half Sector Test

## Slide 5: Half Sector Test (HST)

## Slide 6: Half Sector Test (HST)

## Slide 7: LINAC4 Diagnostics
Now HST
Missing

## Slide 8: Instrumentation Installation Status
All installed apart from
  2 WS inter-tanks PIMs
  2 WS L4T line
  Devices to be put at HST location
    2 SEM grids + 1 WS
    2 BCTs
    1 BPM
  2nd BSM to be moved to final location (now tested upstream of HST)
  Laser Emittance Meters

## Slide 9: Laser Emittance Meter

## Slide 10: Laser Emittance Meter
2 x (laser + diamond) stations

## Slide 11: Laser Emittance Meter (diamond read-out)

## Slide 12: Laser Emittance Meter (status)
Mechanics
  Design and intégration complete
  Vacuum chambres & supports in production
  Installation scheduled in May
Electronics
  Backplanes in production
  Pre-amplifiers in test phase
  Cables pulled in May
  Commissioning ~ June 2017
Laser
  Commissioning end of March at RHUL
Diamond Detectors
  Being currently assembled

## Slide 13
Issues concerning installed device

## Slide 14: BCT
1 L4Z (dump line) and 2 HST BCTs measure > 100 % transmission of H-
L4Z BCT measures < 90% transmission of protons after stripping (expected > 99%)
107%
110%
106%

## Slide 15: BCT

## Slide 16: BCTs
All HW and SW checked / swapped
  Problem still not understood
Next:
  Check BCT grounding
  Try to use BPM intensity signal (close to HST BCT) to measure transmission w.r.t. upstream BCT to see if intensity dependence ~ HST BCT
  Re-investigate beam intensity dependence

## Slide 17: SEM Grids
Case of wires giving very perturbed signals
Problem under investigation
OK
NOT OK

## Slide 18: SEM Grids
(OTHER GRID)
Case of wires giving very perturbed signals
Problem under investigation
NOT OK

## Slide 19: SEM grids
LEBT SEM grids at source test stand
  Removed, several broken wires
  Refurbish in the next few weeks
  New tank design (equipped with new generation of grid PCBs) in xxx months
    Will be then implemented also in tunnel LEBT (YETS?)
Source team very pushy to get priority
Need to review beam intensity-plse length limits to avoid other wire breakage

## Slide 20: WS
missing mechanical offsets for L4T scanners
(to have beam position on top of beam size)
New FESA + OP GUI featuring ppm acq. mode fully tested
To be noted:
Both grids and scanners not extensively used in 2017, other hiccups may show up
Need to plan a re-test of all devices before end of march to identify any need for maintenance during April-May

## Slide 21: BPM
New SW for trajectory operational
ToF was essential for Linac setup, BI expert mode
  SW development on-going
Use of intensity signal for HST under study
BSM
Two monitors validate in BI expert mode, essential to setup LINAC and TL
VME electronics + OP GUI to be made operational during reliability run

## Slide 22: BLMs
Installed
Extensively used for HST studies
Diamond BLM at HST suffering large disturbances every LINAC pulse, with no beam
Thresholds to be studied during reliability run

## Slide 23: Half Sector Test

## Slide 24: H0-H- beam current monitors

## Slide 25: Signal Monitoring Electronics

## Slide 26: Interlock Electronics

## Slide 27: H0-H- B.C.M., Signal disturbances
No beam
Analogue output of signals from VME card
Large disturbances every 1.2 sec
Interlock all the time
10 ms/div

## Slide 28: H0-H- BCM - calibration
Disturbances ~stable (change with BSW on-off) can be substracted (ref trace with no beam)
Calibration studies on-going
  No stripping foil, send full beam on single plate, compare to BCTs
New VME card
  Ready in few days

## Slide 29: HST - BTV
To observe beam and stripping foil
Operational (and essential)
Minor mechanical (screen slightly tilted at IN position) understood, will be ok for PSB

## Slide 30: Summary
Installation in April- May
  Missing grids and WS
  Laser EM
  (ESS BGI radiation tests system)
Issues with xxx SEM grid wires
Issues with LEBT grids
3 BCTs with too high current reading under investigation
BPM ToF to be made operational
H0-H- BCM to be calibrated (hopefully including interlock)

## Slide 31: The Plan
Up to March 31st:
  HST
  Re-assess (with beam if needed) overall instrumentation status)
  At the same time: review any need for SW consolidation/upgrade in visage of reliability run
April-May
  Installation and maintenance
May->YETS
  Reliability run, will also be BI reliability run