## Slide 1: SPS and LHC BCTDCs link to the SMP
Stephane Bart Pedersen, Stephane Gabourin, Lars Jensen, Patrick Odier, Ivan Romera Ramirez

## Slide 2
SPS BCTDCs and links to SMP
  2016
    Layout
  2017
    Status and layout
LHC BCTDCs and links to SMP
  2016
    Layout
  2017
    Status and layout
Commissioning
Outline

## Slide 3
Layout of the SPS BCTDCs link to the SMP in 2016
S.Gabourin

## Slide 4
Improved level of redundancy for high intensity: 2 separate systems
Status of the SPS BCTDCs link to the SMP in 2017
BA3
1 industrial BCTDC
High beam intensity measurement (protons); 14 bit; 100 S/s; rms noise ~6 E9 charges
1 copper link to SMP
BA4
1 industrial BCTDC
Low beam intensity measurement (ions, pilots); 14 bit; 100 S/s; rms noise ~6 E7 charges
2 optical links to SMP (2 CTDAB modules to replace 1 CTDCR and 1 CTDLT) NEW
BA5
2 CERN made BCTDCs
High beam intensity measurement (protons); 16 bit; 100 S/s; rms noise ~1.5 E9 charges
1 optical links to SMP NEW
monitors
electronics
acquisition crates
FESA class
HW modification to be tested in machine conditions for a potential reduction by a factor of ~ 4

## Slide 5
Layout of the SPS BCTDCs and link to the SMP in 2017
S.Gabourin

## Slide 6
Layout of the LHC BCTDC link to the SMP in 2016

## Slide 7
System A Beam 1 & 2
2 CERN made BCTDCs
16 bit ADC located in surface; 4 ranges + aux signals; 1 S/s; rms noise ~ 8 E8 charges
  MTT modules remain in place (links to SMP in 2016 still available)
24 bit ADC located next to the monitors; 1S/s; rms noise ~ 8 E8 charges
  2 optical links to SMP. NEW. Successfully tested in 2016 TS3
  no experience with ageing due to radiations
Status of the LHC BCTDC and link to the SMP in 2017
System B Beam 1 & 2
2 CERN made BCTDCs
16 bit ADC located in surface; 4 ranges + aux signals; 1 S/s; rms noise
  2 optical links to SMP; 10 S/s; SW filter on B2
24 bit ADC located next to the monitors; 1S/s; rms noise
Again good level of redundancy: 2 separate systems
~8.0 E8 charges (B1)
~1.6 E9 charges (B2)
monitors
electronics
acquisition crates
FESA class
~8.0 E8 charges (B1)
~1.6 E9 charges (B2)

## Slide 8
Layout of the LHC BCTDC link to the SMP in 2017

## Slide 9: Commissioning
Precise calibration of the BCTDCs
Reference current injected into the monitors, averaging over 60s for the offset and scaling factors measurement
  16 bit, 4 ranges
  24 bit 1 range
Check of the data transmission to the SMP
SPS and LHC: week 14 (3rd to 7th of April), to be discussed / confirmed