## Slide 1: Ground Vibration Monitoring at CERN as Part of the International Seismic Network
with contributions of many people from integration, civil engineering team, IT…
CERN – LabVIEW Developer Days Genève – 1st June 2017
EN/MME
Morgane Cabon
Michael Guinchard
EN/STI
Kevin Develle
Cedric Charrondiere

## Slide 2: Contents
Introduction
Technical proposal
Solution implemented
Stations performance
Conclusion

## Slide 3: Background knowledge

## Slide 4: LHC
The energy of a 200m long fast train at 155 km/hour
corresponds to
the energy of 360 MJ stored in one LHC beam
The Large Hadron Collider (LHC) : A 27km ring

## Slide 5: Motivations
  Continuous LHC monitoring
    Beam sizes around 1 mm during injection, 0,25 mm at top energy
    Stability of the magnetic field center (depending LHC operation):
      Safe conditions : < 5 μm
      Unstable conditions : 5 to 20 μm
    Beam abort : > 20 μm

## Slide 6: Motivations
HL-LHC Civil engineering activities
H2 (ω)
H1 (ω)
H0 (ω)
      Monitor our installation to anticipate some risks on the beam stability generated by civil engineering activities.
      Study ongoing to determinate the best excavation technique.

## Slide 7: Motivations
Geneva Program “Géothermie 2020”
      Evaluate effects of the micro-seismicity induced by the geo-thermal exploitation on CERN installations.
      Study by the engineering company Resonance mandated by SIG (Services Industriels de Genève)
Expected earthquakes | During the stimulation period | During the stationary phase | Expected cold mass movements
Magnitude 2 | Several per week | A few per month | ~1-10 μm
Magnitude 3 | A few per month | 2 to 3 per year | ~10-100 μm
      Need for a seismic network
        Measuring near source earthquakes to better know the region seismicity
        Integrated into the worldwide seismic network

## Slide 8: Requirements
Surface station | 
 | bdg. 1173 : CERNS
Underground station | 
 | CMS : CERN5
Underground station | 
 | ATLAS : CERN1

## Slide 9: Requirements
 | Needs for a common station | Specific needs for CERN
 | Geothermal exploitation, SED | HL-LHC excavation
Min amplitude | LHC ground motion | LHC ground motion
Max amplitude | 20 m/s2 | 20 m/s2
Min frequency | 1/30 Hz | 1 Hz
Max frequency | 20 Hz | 100 Hz
Timestamp precision | ~ 1 ms | ~ 10 ms
Data latency | 10 s | 1 min
Data format | MiniSEED | ASCII
Data transmitted to | SED servers | CERN servers, LHC database
Other needs | Available 24/7 | Independent systems with same software and hardware
Due to different needs, it was easier to develop a solution internally

## Slide 10: Hardware proposal
Model | Guralp 6T | Guralp 40T | Kinemetrics EpiSensor ES-T
Output | Velocity | Velocity | Acceleration
Triaxial | Yes | Yes | Yes
Frequency range | 30s to 100Hz | 60s to 100Hz | DC to 200Hz
Sensitivity | 2000 V/(m/s) | 800 V/(m/s) | 2,5 V/g
Analog/digital | both | both | Analog +-5V differential
Noise | 172 dB | 172 dB | 155 dB
Power supply [V] | 10 to 36V | 10 to 36V | 12V
LHC Ground motion level | Yes | Yes | No
Threshold level for earthquake | ≈ M3 @10km | ≈ M3,5 @10km | M 7,5 @10km
Sensors
Strong motion sensor
Broadband seismometer
With a 10V, 24 bits ADC

## Slide 11: Hardware proposal
      NI CompactRIO 9035 (1 GB RAM, 2 cores)
      with NI C-series 9239 (+/-10V, 24bits, fs = 1.6 to 50kHz, 70uV noise)
      Linux Real Time
      NTP timing source at CERN (<1ms accuracy)
      GPN connection
      Remote reboot: GUDE Expert power control 1103

## Slide 12: Software proposal
Power Spectral Density :
Ground motion monitoring
(frequency domain)
MiniSEED Files :
Standard for Exchange of Earthquake Data
(time domain)
Raw data
CERN Users
Mechanical Measurement Laboratory
CERN Experts
SED and worldwide seismic databases
Raw data (time domain)
Power Spectral Density (frequency domain)
Sum Level (time domain)

## Slide 13: Advantages
      CERN is a LabVIEW Center of excellence
      Huge experience with CompactRIO at CERN
      Timekeeper API
      Sounds and Vibration, HTTP, WebDAV palette from NI
      RADE palette from MTA Team to communicate with CERN infrastructure
      Distributed architecture
      System fulfilling the requirements
Challenges
      Manpower cost more important for the development
      Mseed converter to be develop
      Quality of the streaming to SED to be proven
      Non-standard seismic station
Technical proposal

## Slide 14: Software architecture
CERN Users
Mechanical Measurement Laboratory
CERN Experts
SED and worldwide seismic databases
1173 Station
UL16 Station
UP53 Station
Remote
connection
Settings
HTTPS
WEBDAV
CMW GPN
CMW TN
Program

## Slide 15: Software architecture
HTTPS
WEBDAV
CMW GPN
CMW TN

## Slide 16: Data Acquisition Module
      Read data from DMA
      Keep Timekeeper synchronized
      Down sample and filter from 2kHz to 250Hz
      Sinc filter with Blackman window

## Slide 17: PSD Module
Parameter [unit] | Value
Representation | Quadratic RMS [m2/Hz]
Window | Hanning
Block length | 64s
Overlap | 66.60%
Frequency resolution fr [Hz] | 0.015625
Max. frequency [Hz] | 97.7

## Slide 18: MiniSEED Converter & HTTP Module
      Build MiniSEED header (timestamp, sequence number, station name…)
      Send package to SED
      Handle SED server reboot, network disconnection
      Objective: 10s latency
6s
300ms
1~2s
data acquisition
mseed conv
mseed read
HTTP answer
CERN
SED
HTTP post
network
network

## Slide 19: System Monitoring Module
      Monitor CPU/RAM usage
        Maximum CPU peak at 50%
        Average CPU ~ 10%
        About 50% of the 1GB RAM available
      Monitor VI states
      Monitor number of elements in queue and RT fifo
      Send “alive” message to an other machine
      Monitor NTP synchronization
NTP server
seconds
Octal code
milliseconds

## Slide 20: The timing challenge
      Problem with SED :
      The 2kHz on-board clock is not perfect. The gap between the timestamp of sample 1999 of DMA 0 and sample 0 of DMA 1 had to be reduced below 1us to fulfil SED requirements
DMA 0
DMA 1
S1999
S0
S1
S0
S1
…
T0 from ntp
500
us
500
us
500
us
~505
us
T0 from ntp
theoretical time vector
With white rabbit, we measure the loop speed at ~500.003 us
1s
http://www.ohwr.org/projects/white-rabbit/wiki
S2
…
500
us

## Slide 21: The timing challenge
      Solution :
      The timekeeper API to timestamp every sample on the FPGA
DMA 0
DMA 1
S1999
S0
S1
S0
S1
…
No big gap anymore
1s
S2
…
500us
+~3ns
Time from timekeeper
500us
+~3ns
500us
+~3ns
500us
+~3ns

## Slide 22: Software Robustness
Stores experts files on USB drive
Stores everything on USB drive
Automatic restart
Problem
CompactRIO reaction
Sends warning and keeps streaming
Stores MiniSEED files on USB drive,
sends it when SED is back

## Slide 23: Installation

## Slide 24: Stations performance
Ground motion measurements versus position
Chart: Frequency [Hz]
Micro-seismic peak
=
average magnitude of FFT spectrum [rms] 𝑓𝑟: 𝐹𝑟𝑒𝑞𝑢𝑒𝑛𝑐𝑦 r𝑒𝑠𝑜𝑙𝑢𝑡𝑖𝑜𝑛

## Slide 25: Stations performance
Recorded earthquakes

## Slide 26: Conclusion
The seismic network already permits to:
  Monitor the LHC ground motion
  Track earthquakes on a wide range of amplitude
It satisfied successfully the needs of SED and CERN
      System compliant with LHC tunnel and CERN network
      Standard format files for seismologists
    SED is doing some regular checking on CERN stations as for their own stations, CERN is maintaining the stations.

## Slide 27: Conclusion
    “CERN stations fully integrated in network, expect to operate like any other SED station, and included permanent archive”
    “CERN stations now used in all earthquake related products”
    Dr. John Clinton, Director of Seismic Networks and Head of the Earthquake Monitoring section at the SED

## Slide 28: Future of the project
Further calculations will be implemented in order to set a warning trigger during the excavation.
A further collaboration will permit to get a warning in case of large earthquake.
The possibility of sending the data with a 2s latency could allow to add the stations into the fast triggering system and forecast the earthquakes arrival.
The CompactRIO is not using all of its CPU and RAM: The available slots can be used for future developments.
The price of a future station is estimated to 27 000 CHF.

## Slide 29: Thank you! Questions