## Slide 1: Vacuum qualification PSBBWSRA
Jose A. Ferreira
Tests: Alice Michet & Nicolas Thaus
Installation: Guillermo Merino & Paul Demarest

## Slide 2: Why vacuum qualification?
Mandate of TE-VSC is to guarantee the quality of the vacuum in CERN’s accelerators
Leak Tightness
Contamination
Outgassing rate (and time dependency)
Virtual Leaks
Leak detection
RGA
Pumpdown curve
Accumulation tests

## Slide 3: Which criteria?
New document to be circulated soon (15th of April)
Leak tightness does not change
Outgassing values will refer to 24h (instead of 10h), but equivalent values  PSB 5∙10-6 mbar∙l∙s-1 (24h) [4 m2 of SS]
Same RGA limits. Components with low outgassing  correction to request same partial pressure (more flexible)

## Slide 4: SPS experience
https://indico.cern.ch/event/393525/

## Slide 5: SPS experience
https://indico.cern.ch/event/393525/

## Slide 6: BWSRE-Materials
Item | SPSBWSRE0003 | Drawing | Material | To be checked
1 | Drum | SPSBWSRE0005 | 316L/316LN | 
2 | Optical disc | SPSBWSRE0009 | Steel, Glass, Inconel, Aluminium | 
3 | Optical sensor holder | SPSBWSRE0065 | Aluminium, stainless steel | Optical sensor
 |  |  |  | Optical sensor
4 | Shaft | SPSBWSRE0011 | 304L | 
5 | Shaft end | SPSBWSRE0013 | Alu EN AW6082 | 
6 | Washer | SPSBWSRE0007 | 304L | 
7 | Optical fibre cover | ? | ? | 
8 | Spring washer | SPSBWSRE0004 | Steel | 
9 | Round tube for optical fibre | ? | ? | 
10 | Stepped vacuum chamber assembly | SPSBWSRE0014 | 316LN | 
11 | RF screen | SPSBWSRE0012 | 304L, CuBe | 
12 | Fork | SPSBWSRE0028 | Ti | 
13 | Thin RF shield assembly | SPSBWSRE0027 | ? | X
14 | Bearing O32-O20-7 | ? | ? | XXX
15 | Bearing O47-O35-7 | ? | ? | XXX
16 | RF screen internal support | ? | ? | X
17 | RF screen hinge | ? | ? | XX
18 | Small RF screen support | SPSBWSRE0038 | 316L | 
19 | Small RF screen support | SPSBWSRE0039 | 316L | 
20 | Ring 10-13-1.35 | SPSBWSRE0051 | 304L | 
21 | Ring 25-28-2.4 | SPSBWSRE0052 | 304L | 
22 | Encoder stop ring | SPSBWSRE0053 | 304L | 
23 | Fork | SPSBWSRE0006 | Ti | 
24 | Optical feedthrough | ? | ? | XX
25 | Break | ? | ? | Outside vacuum
26 | Optical fibre | ? | ? | XX
27 | Encoder | ? | ADMOTEC: R03620 = 036 - 000(036-030A+036-020C) | Outside vacuum
28 | Brushless motor | ? | ? | Outside vacuum
29 | Wire | ? | ? | X
30-40 | nuts and bolts |  | steel | 
41 | Dome | SPSBWSRE0040 | Alu EN AW6082 | Outside vacuum
42 | Resolver and break support | SPSBWSRE0037 | Alu EN AW6082 | Outside vacuum
43 | Resolver stop | ? | Alu EN AW6082 | Outside vacuum
44 | Stator fitting tube | ? | Alu EN AW6082 | ?

## Slide 7: BWSRE-Materials
Item | SPSBWSRE0003 | Drawing | Material | To be checked
1 | Drum | SPSBWSRE0005 | 316L/316LN | 
2 | Optical disc | SPSBWSRE0009 | Steel, Glass, Inconel, Aluminium | 
3 | Optical sensor holder | SPSBWSRE0065 | Aluminium, stainless steel | Optical sensor
 |  |  |  | Optical sensor
4 | Shaft | SPSBWSRE0011 | 304L | 
5 | Shaft end | SPSBWSRE0013 | Alu EN AW6082 | 
6 | Washer | SPSBWSRE0007 | 304L | 
7 | Optical fibre cover | ? | ? | 
8 | Spring washer | SPSBWSRE0004 | Steel | 
9 | Round tube for optical fibre | ? | ? | 
10 | Stepped vacuum chamber assembly | SPSBWSRE0014 | 316LN | 
11 | RF screen | SPSBWSRE0012 | 304L, CuBe | 
12 | Fork | SPSBWSRE0028 | Ti | 
13 | Thin RF shield assembly | SPSBWSRE0027 | ? | X
14 | Bearing O32-O20-7 | ? | ? | XXX
15 | Bearing O47-O35-7 | ? | ? | XXX
16 | RF screen internal support | ? | ? | X
17 | RF screen hinge | ? | ? | XX
18 | Small RF screen support | SPSBWSRE0038 | 316L | 
19 | Small RF screen support | SPSBWSRE0039 | 316L | 
20 | Ring 10-13-1.35 | SPSBWSRE0051 | 304L | 
21 | Ring 25-28-2.4 | SPSBWSRE0052 | 304L | 
22 | Encoder stop ring | SPSBWSRE0053 | 304L | 
23 | Fork | SPSBWSRE0006 | Ti | 
24 | Optical feedthrough | ? | ? | XX
25 | Break | ? | ? | Outside vacuum
26 | Optical fibre | ? | ? | XX
27 | Encoder | ? | ADMOTEC: R03620 = 036 - 000(036-030A+036-020C) | Outside vacuum
28 | Brushless motor | ? | ? | Outside vacuum
29 | Wire | ? | ? | X
30-40 | nuts and bolts |  | steel | 
41 | Dome | SPSBWSRE0040 | Alu EN AW6082 | Outside vacuum
42 | Resolver and break support | SPSBWSRE0037 | Alu EN AW6082 | Outside vacuum
43 | Resolver stop | ? | Alu EN AW6082 | Outside vacuum
44 | Stator fitting tube | ? | Alu EN AW6082 | ?
Stainless steel, CuBe and Al series 6000

## Slide 8: BWSRE-Materials
Item | SPSBWSRE0003 | Drawing | Material | To be checked
1 | Drum | SPSBWSRE0005 | 316L/316LN | 
2 | Optical disc | SPSBWSRE0009 | Steel, Glass, Inconel, Aluminium | 
3 | Optical sensor holder | SPSBWSRE0065 | Aluminium, stainless steel | Optical sensor
 |  |  |  | Optical sensor
4 | Shaft | SPSBWSRE0011 | 304L | 
5 | Shaft end | SPSBWSRE0013 | Alu EN AW6082 | 
6 | Washer | SPSBWSRE0007 | 304L | 
7 | Optical fibre cover | ? | ? | 
8 | Spring washer | SPSBWSRE0004 | Steel | 
9 | Round tube for optical fibre | ? | ? | 
10 | Stepped vacuum chamber assembly | SPSBWSRE0014 | 316LN | 
11 | RF screen | SPSBWSRE0012 | 304L, CuBe | 
12 | Fork | SPSBWSRE0028 | Ti | 
13 | Thin RF shield assembly | SPSBWSRE0027 | ? | X
14 | Bearing O32-O20-7 | ? | ? | XXX
15 | Bearing O47-O35-7 | ? | ? | XXX
16 | RF screen internal support | ? | ? | X
17 | RF screen hinge | ? | ? | XX
18 | Small RF screen support | SPSBWSRE0038 | 316L | 
19 | Small RF screen support | SPSBWSRE0039 | 316L | 
20 | Ring 10-13-1.35 | SPSBWSRE0051 | 304L | 
21 | Ring 25-28-2.4 | SPSBWSRE0052 | 304L | 
22 | Encoder stop ring | SPSBWSRE0053 | 304L | 
23 | Fork | SPSBWSRE0006 | Ti | 
24 | Optical feedthrough | ? | ? | XX
25 | Break | ? | ? | Outside vacuum
26 | Optical fibre | ? | ? | XX
27 | Encoder | ? | ADMOTEC: R03620 = 036 - 000(036-030A+036-020C) | Outside vacuum
28 | Brushless motor | ? | ? | Outside vacuum
29 | Wire | ? | ? | X
30-40 | nuts and bolts |  | steel | 
41 | Dome | SPSBWSRE0040 | Alu EN AW6082 | Outside vacuum
42 | Resolver and break support | SPSBWSRE0037 | Alu EN AW6082 | Outside vacuum
43 | Resolver stop | ? | Alu EN AW6082 | Outside vacuum
44 | Stator fitting tube | ? | Alu EN AW6082 | ?
Stainless steel, CuBe and Al series 6000
Outside vacuum

## Slide 9: BWSRE potential OG sources
Item | SPSBWSRE0003 | Drawing | Material | To be checked
2 | Optical disc | SPSBWSRE0009 | Steel, Glass, Inconel, Aluminium | 
3 | Optical sensor holder | SPSBWSRE0065 | Aluminium, stainless steel | Optical sensor
 |  |  |  | Optical sensor
7 | Optical fibre cover | ? | ? | 
9 | Round tube for optical fibre | ? | ? | 
13 | Thin RF shield assembly | SPSBWSRE0027 | ? | X
14 | Bearing O32-O20-7 | ? | ? | XXX
15 | Bearing O47-O35-7 | ? | ? | XXX
16 | RF screen internal support | ? | ? | X
17 | RF screen hinge | ? | ? | XX
24 | Optical feedthrough | ? | ? | XX
26 | Optical fibre | ? | ? | XX
29 | Wire | ? | ? | X
44 | Stator fitting tube | ? | Alu EN AW6082 | ?
EDMS 1335217. Tests October 2013
Q10h< 1.7 10-7 mbar∙l∙s-1
Q10h< 4.9 10-7 mbar∙l∙s-1
Q10h< 4.9 10-7 mbar∙l∙s-1
Q10h< 8.3 10-7 mbar∙l∙s-1
Q10h< 1.1 10-6 mbar∙l∙s-1
Q10h< 1.7 10-7 mbar∙l∙s-1
Q10h< 6.5 10-7 mbar∙l∙s-1
Measurement of all the components
Q10h= 6.7 10-7 mbar∙l∙s-1
P∝1/t
No sign of contamination

## Slide 10: Acceptance test
  Q10h= 8.1∙10-5 mbar∙l∙s-1
P∝1/t0.5!
No contamination observed
Test December 2014
2 fibres B + 2 feedthroughs + 2 focusers (1st type) + 1 disk | 2.55∙10-6 mbar∙l∙s-1
2 fibres B + 2 feedthroughs + 2 focusers (2nd type) + 1 disk | 3.19∙10-6 mbar∙l∙s-1
2 fibres A + 2 feedthroughs + 2 focusers (1st type) + 1 disk | 1.28∙10-5 mbar∙l∙s-1
2 fibres A + 2 feedthroughs + 2 focusers (2nd type) + 1 disk | 1.34∙10-5 mbar∙l∙s-1
Chart: Time [h]
| 1 | 10 | 100 | 1000 | 10000
 | 1.2807e-05 | 4.05e-06 | 1.2807e-06 | 4.05e-07 | 1.2807e-07

## Slide 11: Residual Gas Analysis PSB FWS

## Slide 12: Residual Gas Analysis PSB FWS

## Slide 13: Residual Gas Analysis PSB FWS
Peak 32 linked to coaxial cable. Virtual leak

## Slide 14: Residual Gas Analysis PSB FWS
Issues after cleaning/assembly required to repeat the process to get here
Origin of contamination not fully understood
Cleaning and assembly of different components is critical

## Slide 15: Tests PSB wire scanner
Full BWSRA
4.7∙10-5 mbar∙l∙s-1 (24h)
1.1∙10-4 mbar∙l∙s-1 (10h)
Several tests without Kapton coaxial cable
From 2 to 7∙10-6 mbar∙l∙s-1 (24h)
PSB requirement 5∙10-6 mbar∙l∙s-1 (24h)

## Slide 16: Tests PSB wire scanner
Complete FWS
Without coaxial cable
Background
PSB limit

## Slide 17: Tests PSB wire scanner
Complete FWS
Without coaxial cable
Background
Without coaxial cable in acceptable values
Discrepancy of SPS results finally explained!
PSB limit

## Slide 18: Outgassing cable #2
1st Pumpdown
2nd Pumpdown after bakeout 150C@24h and 24h in air

## Slide 19: Outgassing cable #2
1st Pumpdown
2nd Pumpdown after bakeout 150C@24h and 24h in air
Approx. same outgassing as cable #1

## Slide 20: Outgassing cable #2
1st Pumpdown
2nd Pumpdown after bakeout and 24h in air
Some improvement (factor ~2) on the short term t<100h

## Slide 21: Outgassing cable #2
1st Pumpdown
2nd Pumpdown after bakeout 150C@24h and 24h in air
No improvement on the long term t<100h

## Slide 22: Outgassing cable #2
1st Pumpdown
2nd Pumpdown after bakeout 150C@24h and 24h in air
No improvement on the long term t<100h
Pre-degassing is not effective reducing outgassing!

## Slide 23: Outgassing cable #2
After 400h pumpdown in spec
PSB limit

## Slide 24: Interpretation
Diffusion of water in very thin Kapton® layer
Model validated against an extruded wire 0,2 mm (Axon HT3007SCA)
Only observable if <0,5 mm thickness
1 mm layer >1000h!

## Slide 25: Outgassing full wire scanner
During the final test  Observe same behavior?
If same behavior  FWS out of spec at 24h but in spec at >300h

## Slide 26: Outgassing fast wire scanner
FWS does not follow same trend! >10 times higher after 100 h
PSB limit

## Slide 27: Outgassing full wire scanner
During the final test  Observe same behavior?
If same behavior  FWS out of spec at 24h but in spec at >300h
Effective pumping speed 100 l/s. Ion pump 75 l/s added to the FWS  Gain factor of 2 but still out of spec
Vacuum recommendation  Installation but find a replacement for coax. cable in future units

## Slide 28: PS ring
PSB limit
PS limit
Ion run have stringent pressure requirement
Removing coaxial needed to reach the spec, but not enough
Highly recommended to add ion pump 200 – 150 l/s to keep it as an unbaked component
Simulation final location to refine this first estimation

## Slide 29: Conclusions
The coaxial wire is responsible of >90% of the gas load
Highly recommended to explore better alternatives
Retrofit the alternative to PSB and SPS wire scanners
To avoid in situ bake-out in PS Add ion pump. Detailed evaluation required. Position?
1 staff FTE∙month in acceptance tests only for this component until now. Compensation associated cost requested.

## Slide 30: Thank you for your attention