## Slide 1: Results and operational aspects of AD's CCC
Mini workshop on CCC with GSI, Jena and TU Darmstadt @ CERN
26/04/2017
CERN / Geneva
Main CERN contributors:
BE-BI: M. Fernandes, J. Tan, D. Louro, E. Oponowicz, M. Ludwig
TE-CRG: T. Koettig, A. Lees, A. Vacca, F. Carrasco
EN-MME: J.P. Brachet, D. Lombard, G. Foffano
BE-ICS: R. Speroni
My University PhD supervisor:
Carsten P. Welsch
miguel.fernandes@cern.ch

## Slide 2: AD cycle and beam parameters
AD beam parameters | 
Beta | 0.97 … 0.11
Frev [MHz] | 1.59 … 0.17
Cycle | ~114 s
N particles | (5 … 1) x 107
Current | (12 … 0.1) μA
Schottky Intensity Measurement
Coasting beam
Bunched beam
10
20
30
40
50
60
70
80
90
time [s]
0
1
2
3
4
N [x 107 pbar]
0.0
0.2
0.4
0.6
1.0
β [c]
0.8
Requirements for new monitor:
Current resolution: <10 nA
Intensity resolution: 5 x 105 charges
Bandwidth: DC - 1 kHz

## Slide 3: CCC in a nutshell
w
SQUID transfer function
SQUID slew-rate limit
Dynamics performance
Slew-rate of signal flux: < 5 𝝋0/𝜇𝑠
Bandwidth of FLL system: 20 𝑀𝐻𝑧
fc = 1 kHz
SC
SC
150 uF
1 Ohm

## Slide 4: AD slew-rate requirement
AD beam current slew-rate
Sib = = ~10 φ0/μA
Filtering | Slew-rate
None | 86 Gφ0/s
Lp permeability | 1.2 Gφ0/s
Lp + RC filter | 0.97 Mφ0/s

## Slide 5: CCC monitor acquisition
FAIR Nb shield
Di = 185 mm
Do = 280 mm
Nmeander = 20
Gapl = 0,5 mm
Gapt = 1 mm
Cartridge
Magnicon SQUID
Li = 400 nH
RC filter: fLP = 1 kHz
Sib = 10.4 φ0/μA
FLL Electronics
Magnicon (20MHz)
Settings:
  RFLL = 1 kOhm
  GBW = 1.04 GHz
  BWFLL = 1.04 GHz
Advantech
Serial Server
Calibration Source
Triax
VME Crate
CPU L866
CTRV timing
VD80 ADC
  16 channels
  16 bit
  +/- 10V
  800kOhm
  Isolated diff. inputs
RS-232
Ethernet
Calibration loop
NIM Crate
AD B-field cycle
Current/Nparticlenormalization
Acquisition / Controls
VME architecture
FESA SW
Acquisition of signals at end of cycle (1kHz)
Control of multiple parameters possible
HF filtre

## Slide 6: Installation in AD tunnel

## Slide 7: Cryostat Acquisition & Control

## Slide 8: CCC Acquisition and Control

## Slide 9: Software - FESA class
σtotal = 5.8 nA
mi = 1.2 nA
mf = -0.3 nA
BEAM CURRENT [uA]
NUMBER OF PARTICLES

## Slide 10: Performance in 2015
Monitor
Current resolution: 280 nA (BW=1kHz)
After filtering: 30 nA (BW=11Hz)
Offset jump at injection
RF-bypass in ceramic pipe responsible for excess noise (mostly at 50Hz)
Cryogenic
Cryostat & cryocooler clearly not performing as expected
Cryostat empty 3 days after a full refill
Heat in-leak reduced by:
  Better MLI shielding
  Thinner support rods
  Improved insulation vacum pumping
Raw signal (Ts = 0.4ms)
Filtered signal (Navg = 100)
σ(Ibeam) = 280 nA
σ(Ibeam) = 27 nA

## Slide 11: 2016: Performance first beams
Noise and perturbations are much more reduced after removal of RF-bypass
Noise figure depends on cryostat conditions:
  Pressure/temperature variations
  Reliquifier vibrations (when LHe level covers tip of transfer line)
  Worst case σ = 20 nA.
Jump at injection is still present
AD beam current
σ = 3.6 nA
σ = 4.2 nA
σ = 4.3 nA
Jump at injection
Raw signal fs=1kHz
No filtering
Beam ejection

## Slide 12: 2016: AD measurement
σ = 0.001x107
σ = 0.018x107
Measurement of number of particles

## Slide 13: Problem in CCC measurement
Very weird behaviour
It seemed like DC coupling is not working
Explanation (not fully confirmed):
  Accumulated flux in superconducting loops
  Due to jumps occuring at injection
  Made coupling SC loop non-SC
  Remnant magnetization in pickup core
But calibration pulses look fine

## Slide 14: 2017: AD intensity measurement
σ = 5.8 nA
Calibration signal
Number of particles
σ = 2.2x104

## Slide 15: Noise permormance
Measurement at lab in Jena
Measurement in AD-ring
Measurement in AD-ring performed with and without running vacum pumps connected to cryostat insulation vacum (primary and turbomolecular pump)

## Slide 16: Immunity to mechanical vibrations
Main vibration modes as captured by SQUID:
  65 Hz
  115 Hz
  78 Hz
  5 Hz
Preliminary !

## Slide 17: Noise performance improvements
σ = 5.5 nA
Perturbation from magnetic cycle
Perturbation from cryocooler
It should be possible to subtract both these perturbations in a robust way !

## Slide 18: First beams in 2017

## Slide 19: First beams in 2017

## Slide 20: First beams in 2017

## Slide 21: Conclusions and next steps
Measurement noise performance as initially specified 
Very good immunity to mechanical vibrations 
Stand-alone cryogenic system enables long term operations 
Flux jump at injection still present, and most likely perturbative sources have now been discarded… 
Develop expert GUI
Resolution can still be improved through signal processing
Solution for injection jump (additional EM filtering ?)
Assess long term stability over entire year run
Implement automatic SQUID setup in FESA class
Replace Magnicon “Connector box” and “Power supply” by improved electronics, to be installed in a single box
…

## Slide 22
Thank you!

## Slide 23: Baseline stability

## Slide 24: Limiting input signal slew-rate
AD inj current:
12muA
Injected bunch train
Laboratory measurement
Necessary to limit slew-rate of coupled signal to avoid flux-jumps
    AD injection: !!

## Slide 25: CCC Transfer function – RC series
Rf = 2.73kΩ
DC gain = 1.25 mV/nA
100nA = 768 mφ0RMS
10nA = 77 mφ0RMS
RC noise
Frequency response closely follows the theoretical curve
With 10 kHZ low-pass filter at SQUID output turned ON
R=1Ω; C=10uF

## Slide 26: CCC Transfer function – RC series
Noise from RC-shunt is a big contribution to total noise.
Rf = 1kΩ
Flux noise RMS (1Hz – 1MHz) [mφ] | Total | RC-shunt
Calibration @ Core | 34.4 | 14.5