## Slide 1
Collaboration Meeting Ultra-sensitive Beam Current Measurements for future Accelerator Facilities
26. + 27.04.2017, CERN
T.Sieber
Status of the FAIR CCC and Cryostat
26.+27.04.2017
CCC Collaboration Meeting
Thomas Sieber

## Slide 2
CCC Collaboration Meeting
Thomas Sieber
Contents
CCC status at FAIR
Design of the CCC cryostat
Outlook
26.+27.04.2017

## Slide 3
CCC Collaboration Meeting
Thomas Sieber
CCC application at FAIR / CRYRING
26.+27.04.2017
FAIR
Transmission / loss monitoring for slow extracted beams
Extraction efficiency of the synchrotrons
Analysis of the spill structure
Absolute calibration of BI devices like SEM
nA intensity measurements for experiments reaction rates
CR
CRYRING
Accurate tool for ring commissioning and ring operation
nA intensity measurements for experiments reaction rates

## Slide 4
CCC Collaboration Meeting
Thomas Sieber
GSI
FAIR
SIS100
HESR
Cryring
SIS18
UNILAC
Proton Linac
CR
CBM
Super FRS
NuSTAR
PANDA
CCC Position
CCC locations at FAIR
26.+27.04.2017

## Slide 5
CCC Collaboration Meeting
Thomas Sieber
CCC requirements at FAIR
26.+27.04.2017
Detector
- Current resolution as demonstrated in the 2014 runs (or better)
High bandwidth (10 kHz or better)
Improved system robustness
Radiation hardness; estimate 300 Gy/20a  tests at SIS18
Full integration in FAIR control system (FESA)
Cryostat
UHV / 150 mm beam tubes
Stand alone system  reliquefier (not in CRYRING)
Standard operation and testbench for CCC development
Vibration damping relativized but still an issue ( N. Marsic)
Full integration in FAIR control system (FESA)

## Slide 6
CCC Collaboration Meeting
Thomas Sieber
26.+27.04.2017
Radiation hardness test
Desired total dose: 100 Gy
Accumulated dose measured weekly by radiation protection
→ after 4 months only 1 Gy accumulated due to low intensities and improved extraction scheme
Supracon SQUIDs + electronics
Magnicon SQUIDS
Installation in SIS18 extraction-line in May 2016

## Slide 7
CCC Collaboration Meeting
Thomas Sieber
CCC requirements at FAIR
26.+27.04.2017
Detector
- Current resolution as demonstrated in the 2014 runs (or better)
High bandwidth (10 kHz or better)
Improved system robustness
Radiation hardness; estimate 300 Gy/20a  tests at SIS18
Full integration in FAIR control system (FESA)
Cryostat
UHV / 150 mm beam tubes
Stand alone system  reliquefier (not in CRYRING)
Standard operation and testbench for CCC development
Vibration damping relativized but still an issue ( N. Marcsic)
Full integration in FAIR control system (FESA)

## Slide 8
CCC Collaboration Meeting
Thomas Sieber
Status at FAIR
26.+27.04.2017
- Shielding and SQUID system of FAIR CCC successfully tested at Jena 
- Basic design of the cryostat done (with FSU), detailed spec in review 
- LOBIDS started to look into control application 
- Priority on project planning for FAIR during last months 
- Call for tender still blocked by budget restrictions 
Demonstration for Tech. Director + Letter to FAIR management
Alternative solutions ? (Cryring)
- First beam in Cryring achieved, bakeout going on
- ecool commissioning in May, first ‚competing‘ experiments in 2018

## Slide 9
CCC Collaboration Meeting
Thomas Sieber
Design of the cryostat
- GSI BI develops basic cryo + mechanical design and specifications
- FSU Jena provides 3D model and preliminary set of drawings
- GSI cryo department reviews design and spec
TU generates in parallel model for mechanical simulations
 call for tender (negotiated procedure)
- manufacturer provides detailed cryo concept and drawings
- TU verifies mechanical properties by adjusting the theoretical model
 production
Parameter | Value
Pump down time for isolation vacuum (2 x 500 l/s turbo) | <10 hours
Vacuum pressure in IVC (mbar) | 5*10-7  3*10-7
Temperature of cooled thermal shielding (without He) (K) | <50
Leakage rate isolation vacuum (mbar l/s) | <1*10-9
Leakage rate UHV beam tube (mbar l/s) | <1*10-11
Number of temperature sensors (cryo-physics) | 8
Standing time with 80 l He (hours) | >80
Maximum heat load to the He-container (shield at 50K) | <200 mW
Basic parameters fixed
In detailed spec
26.+27.04.2017

## Slide 10
CCC Collaboration Meeting
Thomas Sieber
CCC system developed during early 1990’s at GSI (A. Peters, H. Reeg, W. Vodel et al.)
Test device required for new improved sensor components  system was recommissioned 2013 / 2014
Beam
Titanium suspensions
40 K Radiation shield covered by MLI
30 liter helium container / bath cryostat
Low thermal conductivity feedthroughs
Vacuum vessel; isolation vacuum = beamline vacuum
Pickup unit (SC shield, pickup coil and ring core)
Warmhole for beam passage
GM cryocooler cold head / Enthalpie cooling
Liquid helium level sensor
CCC system developed during early 1990’s at GSI (A. Peters, H. Reeg, W. Vodel et al.)
Test device required for new improved sensor components  system was recommissioned 2013 / 2014
Rubber feet
SQUID
Ceramic gap
He Gas Outlet
The GSI prototype
26.+27.04.2017

## Slide 11
CCC Collaboration Meeting
Thomas Sieber
The GSI CCC prototype
26.+27.04.2017

## Slide 12
CCC Collaboration Meeting
Thomas Sieber
Design of the FAIR cryostat
26.+27.04.2017

## Slide 13
CCC Collaboration Meeting
Thomas Sieber
Mechanical layout
26.+27.04.2017

## Slide 14
CCC Collaboration Meeting
Thomas Sieber
Mechanical layout
Top plate, feedthroughs, He filling
Vacuum chamber
Ti Suspensions
He lines
Thermal shield, Cu
MLI
Liquid He container
Pickup unit/GFK spacers
He container inner tube
150mm beam tube
Thermal shield tube
26.+27.04.2017
yz -cut
xy -cut
Connection to refrigerator
(strips)
Turbo pump
CCC support

## Slide 15
CCC Collaboration Meeting
Thomas Sieber
Mechanical layout
26.+27.04.2017
Vacuum chamber
Ti Suspensions
Thermal shield, Cu
Liquid He container
Pickup unit/GFK spacers
He container inner tube
150mm beam tube
Thermal shield tube
CCC support

## Slide 16
CCC Collaboration Meeting
Thomas Sieber
Re-Liquefaction schemes (TransMIT)
26.+27.04.2017
Liquefaction rate with pulsed tube cooler (He at room temperature): 0.7 l/hour. Costs: ~ 50 k€
Displacement measured at the cooler: 5 mm, damping reduces motion by factor ~102
Inlet exhaust
gas from shield
gas cooling
Inlet exhaust
gas from shield
gas cooling
direct
direct with damping
separate shield cooling
separate shield cooling
with damping

## Slide 17
CCC Collaboration Meeting
Thomas Sieber
Summary / Outlook
CCC detector
Some lessons learned from 2016 (Nb cabling, thermal drift, radiation...)
- Investigation of temperature drift in Jena / develop standard measurement for noise figure
- Repeat radiation hardness test at GSI
Cryostat
Detailed spec: review process to be finished
→ Call for tender as soon as budget is clarified (in the meantime intensify contact with industry)
Provide input for mech. simulations at TU
New PhD student, David Haider, starting in Nov to support this work !!!
Cryring
vacuum commisioning ongoing
restart with beam in May 2017 for 4 weeks
 shutdown. Further planning depending on results
26.+27.04.2017

## Slide 18
CCC Collaboration Meeting
Thomas Sieber
Thank you for your attention !
26.+27.04.2017