## Slide 1: Introduction to RF measurements and instrumentation
Daniel Valuch, CERN BE/RF, daniel.valuch@cern.ch

## Slide 2: Content
RF power measurement
Spectrum analyzers
Vector network analyzers

## Slide 3: RF power measurement
  Diode envelope detection: Most common, large dynamic range, suitable for constant signals
  Thermal: Very accurate, limited dynamic range, ideal for complex modulated signals
  Electronic receivers: Most complex, suitable for fast, pulsed signals
Most popular methods to measure the RF power

## Slide 4: RF power measurement
How to read specifications of a power sensor
Frequency range:
Diode sensors ~10MHz
Thermal sensors DC
Pulsed sensors 50-100MHz
Power/dynamic range:
Minimum and maximum power for your application. Values change typ. factor 1000+ between room temperature and superconducting state
External trigger and interface:
Synchronized acquisition, USB interface, ethernet…

## Slide 5: RF power measurement
How to use a RF power meter
1. Set the frequency
2. Zero the offset
3. Connect sensor to the signal source

## Slide 6: RF power measurement
Chose the desired measurement unit
If measuring through couplers/ attenuators etc. set-up the offset for direct reading of the source power value, instead of measured power value

## Slide 7: RF power measurement
Some comments:
  The reading is noisier at the bottom of the dynamic range
  The power sensor really burns when overloaded (very costly to repair)
  Sensors have optimized dynamic range for type of measurements they do (very accurate, very fast, RMS…)
  Sensors and instruments need certain time to measure a data point, be careful with automation
  Decide what to buy based on how are you going to use the instrument: Table top, USB sensor, Ethernet sensor…

## Slide 8: Spectrum analyzer
Spectrum analyser is a device, which measures frequency content of a signal.

## Slide 9: Spectrum analyser user interface
Frequency
Amplitude

## Slide 10: Spectrum analyser user interface
Start
X axis, frequency span: Span
Center frequency: Center
Stop

## Slide 11: Spectrum analyser user interface
Y axis, 1 square (/division)
Y axis, full scale 10 square
Set by: Amplitude/scale

## Slide 12: Spectrum analyser user interface
Other important settings

## Slide 13: Spectrum analyser user interface

## Slide 14: Spectrum analyzer user interface

## Slide 15: Spectrum analyzer user interface
Screen + Soft keys
Frequency and amplitude
Markers and advanced
RF input

## Slide 16: Spectrum analyser user interface

## Slide 17: Spectrum analyser user interface
Screen + Soft keys
Frequency and amplitude
Markers and advanced
RF and IF inputs

## Slide 18
IF detector bandwidth – resolution bandwidth
Resolution bandwidth
Sweep time
Receiver signal bandwidth prior the power detector
IF/resolution bandwidth defines the measurement noise floor AND how close two different signals can be in order to still distinguish them
Different signals
Noise floor

## Slide 19
IF detector bandwidth – resolution bandwidth
Receiver signal bandwidth prior the power detector
IF/resolution bandwidth defines the measurement noise floor AND how close two different signals can be in order to still distinguish them

## Slide 20: Real time spectrum analyzer

## Slide 21: Signal demodulation by s.a.

## Slide 22: Modern spectrum analysers

## Slide 23: Vector network analyser
Vector network analyser is a device, which excites a RF network and measures a response from its ports. All signals are measured as phasors.
If we know RFL and FWD
we can calculate:
VSWR
S-parameters S11, S12
Reflection coefficient G
Impedance R+jX
Admitance Y+jB
Input matching
If we know TRN and FWD
we can calculate:
Gain, attenuation
S-parameters S21, S22
Transmission coefficient T
Group delay
Phase shift

## Slide 24: Vector network analyser

## Slide 25: Vector network analyser
How does the VNA measure?
  Send signal from port 1 and measure the response at port 1 and port 2

## Slide 26: Vector network analyser
How does the VNA measure?
  Send signal from port 1 and measure response at port 1 and port 2
S11=A-/R

## Slide 27: Vector network analyser
How does the VNA measure?
  Send signal from port 1 and measure response at port 1 and port 2
S21=B-/A+

## Slide 28: VNA user interface
X Axis, typ. Frequency
Start/stop
Center/span
Menu bar

## Slide 29: VNA user interface
Y axis, measured parameter

## Slide 30: VNA user interface
Instrument status, calibration, reference etc.

## Slide 31: VNA user interface
Calculation of a derived parameter using the measured trace

## Slide 32: Setting up process…
Super important for high-Q device measurements

## Slide 33: Few notes on IF bandwidth
VNA uses a super-heterodyne receivers to measure the RF signals
Same noise handling procedures as for the spectrum analyser apply

## Slide 34: IF (detector) bandwidth
Bandpass filter measurement, fc=1GHz
IF bandwidth 100kHz

## Slide 35: IF (detector) bandwidth
Bandpass filter measurement, fc=1GHz
IF bandwidth 10kHz

## Slide 36: IF (detector) bandwidth
Bandpass filter measurement, fc=1GHz
IF bandwidth 1kHz

## Slide 37: IF (detector) bandwidth
Bandpass filter measurement, fc=1GHz
IF bandwidth 100Hz

## Slide 38: IF (detector) bandwidth
Bandpass filter measurement, fc=1GHz
IF bandwidth 10Hz
Noise floor for IFBW=100kHz
Noise floor for IFBW=10Hz

## Slide 39: IF (detector) bandwidth
Bandpass filter measurement, fc=1GHz
IF bandwidth 10Hz
Plus averaging

## Slide 40: Measurement errors and calibration
The instrument always shows some curves…
…but in 99% cases this is not what you want to measure
verydemotivational.com

## Slide 41: Measurement errors and calibration
“Hello Daniel, this is Nikolai. We are trying to measure the 1.3GHz superconducting cavity but the instrument shows something strange…”

## Slide 42: Measurement errors and calibration
We work in the RF domain
The instrument measures our DUT
…but also everything around (cables, connectors, adapters, spurious reflections etc.)
…and the instrument is sensitive to temporal alignment of the signals as well
…finally the instrument also measures himself

## Slide 43: Measurement errors and calibration
All networks which connect DUT to the instrument introduce static systematic errors
We can measure them and mathematically de-embed them from the measurement
This process is called Calibration

## Slide 44: Measurement errors and calibration
All networks which connect DUT to the instrument introduce static systematic errors
We can measure them and mathematically de-embed them from the measurement
This process is called Calibration

## Slide 45: Measurement errors and calibration
Directional coupler measurement without and with full calibration
Without calibration
Full calibration

## Slide 46: Calibration process
3 standards: Open, Short, Load + Thru

## Slide 47: Thank you for your attention
Thank you for your attention
After the break: hands on part