## Slide 1
Fast timing measurement for CMS RPC Phase II upgrade
Christophe COMBARET, IPNL – CNRS IN2P3
On behalf of the CMS Muon Group
XIV Workshop on Resisitive Plate Chambers and Related Detectors

## Slide 2
CMS muon upgrade at high eta
Proposed iRPC
72 chambers to equip area with 1.8 < |η| < 2.4
Add track hits in muon reconstruction
Search for Heavy Stable Charge Particle

## Slide 3
Proposed iRPC and front end electronics for RE3/1 and RE4/1
Strip board :
96 strips per chamber (1cm strip width at η = 2.8) between double gap RPC detectors
Front-end :
Readout on both sides of the strip
Relative timing between both channels of same strip (σ<100ps)
Position along the strip (σ = 1.8 cm )
Absolute timing (σ << 1ns)
 HSCP search
To
BE
Preampl.
+cmp.
Strip
TDC
GBT

## Slide 4
Global prototype measurement chain
iRPC top
iRPC bottom
Front-end board
TDC mezzanine
Petiroc
FPGA
TDC
W5300
Tcp/Ip
TCP-IP
ZDAQ acquisition software
cf. Laurent Mirabito’s talk
DAQ PC
Clock and Control Board
Synchro. FPGA
Clk
Sync. cmds
Busy
HV
Gas system
Z matching network
32
Requirements :
Fast preamp. : Petiroc ASIC
Precise TDC : Wave union TDC

## Slide 5
Petiroc ASIC
Parameter | Value
Number of Channels | 32
Signal Polarity | positive or negative
Sensitivity | Voltage input amplifier, 200 Ohm matching
Timing Resolution | ~ 18 ps RMS on trigger output (4 photoelectrons injected)
Dynamic Range | 160 fC up to 400pC
Packaging & Dimension | TQFP 208 (28x28x1.4 mm) TFBGA 353 (12x12x1.2mm)
Power Consumption | 6 mW/channel
Inputs | 32 analogue inputs, No external component required Inputs DC level adjustable
Outputs | 32-channel trigger outputs ASIC level general trigger (OR of all channel) ASIC level second level general trigger (OR of all channels) Charge measurement (10 bits) Time measurement (10 bits)
Internal Programmable Features | Common trigger threshold adjustment and 6bit-DAC/channel Shaping time & gain of the charge shaper 32 x 8bit-input DAC over 1V span

## Slide 6
Wave union TDC
Implementation of carry chain TDC in FPGA :
Width of bins are different (Vcc and temperature). Typ. 60ps for Altera Cyclone II
Some ultra-wide bins (LAB boundary crossing). Typ. 165ps for Altera Cyclone II
Wave union TDC solves this problem by splitting the input signal in 2 edges and ensuring that in any case, at least one edge is not in the ultra-wide bin.
Details in : QI Ji, DENG Zhi, Liu Yi-nong. Nuclear Electronics &Detection Technology, 2011, 31(4): 378-381.
IN
T0
T1
T2
T3
ΔT
LUT
ΔT
ΔT
ΔT
Principle :
Input signal is propagated into a chain of identical delays (bins)
All delays output are latched on the system clock
Fine delay to clock edge is measured by position of last crossed delay element.

## Slide 7
Small size Electronics prototype
50cm
DAQ PC
TCP/IP
DCC board
Clk and sync
Trigger
PMs
Top iRPC
Bottom iRPC
Pcb with buried strips, readout on both sides
- 32 x 50cm strips (3,5mm pitch ) and 32 off detector return lines (1mm wide)
- 2 iRPC chambers : 1.4/1.4 mm and 1.6/1.6 mm
- 2x Petiroc ASIC and 2x Wave union TDC mezzanine
Time resolution with 5pC injection

## Slide 8
May 2017 beamtest at CERN (H4) : setup
H4 beam area
- Muon beam (150GeV)
- Tiny scintillators (1cm width)
- Scan studies were performed using moving tables (< 1mm resolution)

## Slide 9
May 2017 beamtest at CERN (H4): Results
 σ
Signal Propagation Speed : V ~2*1 cm/0.11 ns ~18 cm/ns
σ(T2-T1) ~100 ps

## Slide 10
Real size strip PCB
Each strip PCB covers half of a cassette with 48 strips (1 cm each)
The strip, both gaps and cassette behave as a stripline where the cassette is the ground planes
Impedance of the strip is defined by :
              Where :
              Rs : Resistance
              Ls : Inductance
              Gp : Conductance
              Cp : Capacitance
48 strips
Side | Cp (pF) | Gp (S) | Ls (nH) | Rs (m | Zc (Ohm)
Wide | 244 | 934 | 482 | 467 | 43,5
Narrow | 244 | 934 | 487 | 461 | 44
Side | Cp (pF) | Gp (S) | Ls (nH) | Rs (m | Zc (Ohm)
Wide | 244 | 934 | 482 | 467 | 43,5
Narrow | 244 | 934 | 487 | 461 | 44
1,6m
GND area
GND area

## Slide 11
Real size PCB, Impedances and EMC
To minimize signal reflections, the stripline impedance must be controlled up to the asic.
3 methods were used to measure strip impedance :
Direct measurement of line parameters with a RLC meter (at 2MHz)
Direct measurement with potentiometric line adaptation
            Zc= 46 Ω
Reflection Method
Side | Cp (pF) | Gp (µS) | Ls (nH) | Rs (mΩ) | Zc (Ω)
Wide | 244 | 934 | 482 | 467 | 43,5
Narrow | 244 | 934 | 487 | 461 | 44
Zc=41Ω
E
Rg
Strip
Rvar
E
Rg
Strip

## Slide 12
Calibrations (2) : injection on the PCB - principle
Change is injected on one side (point A)
Charge is divided in two equal halves, each propagates in one direction of the strip
Cross talk is measured on adjacent strip
B
A
A
A’
Injected charges measured at both ends of the strip
Cross talk on two adjacent strips
A : 2mV/div
A’ : 2mV/div
Timebase : 2ns/div
A : 2mV/div
A’ : 2mV/div
Timebase : 2ns/div
oscilloscope
Pulse generator
Agilent 81134A
50Ω
1pF
A
B
50Ω coax.
50Ω coax.
A’
50Ω load
50Ω load

## Slide 13
FE Calibrations : noise response, no injection
FE Calibration method :
With only one channel activated at once in the Petiroc ASIC :
  Threshold of comparator lowered step by step with short acquisition window
Alignment of all channels pedestal to the same value (individually saved in configuration database) . Erf fit function is used to compute pedestal value
Pedestal = 480 DAC

## Slide 14
FE Calibrations : noise response, no injection
FE Calibration method :
With all Petiroc channels activated together :
    Threshold of comparator lowered step by step with short acquisition window (alignment check)
Channels alignment = 480
Turn On = 481

## Slide 15
FE Calibrations : noise response, no injection
FE Calibration method :
With all channels activated together and strips connected to the FE board :
    Threshold of comparator lowered step by step with short acquisition window (alignment check and noise sources hunt)
Threshold =500
Alignment = 480
Turn On= 486

## Slide 16
Calibrations (3) : injection on the PCB – results (for FE#2)
For this FE Board :
1pF injection capacitance used on high radius
For high radius :
Pedestal from fit = 476 DAC
Threshold set to 500 DAC
Minimal charge that can be seen is
  = 117 fC for the strip
 59 fC per channel on injection side
  On Low radius side, = 206 fC for the strip
   103 fC per channel on other side

## Slide 17
Calibrations (4) : injection on the PCB – results (for FE#2)
 : injection side
 : opposite side
Full efficiency on injection side : 170 fC (85 fC seen on each side)
Full efficiency on both sides : 230 fC (115 fC seen on each side), less than 200 ps resolution (3,6 cm)
170fC
230fC
230fC

## Slide 18
Conclusions and prospects
Longitudinal position on the strip measured by difference of signals time arrival on both sides achieved with good resolution
Long Strip PCB behavior and impedance understood and under control
Wave union TDC used with Petiroc ASIC
Next steps and prospects :
FE board version 2 : 64 channels embedded in cyclone V GT FPGA
Petiroc3 ASIC (lower threshold and improved noise immunity)
Time over Threshold (ToT) to measure charge and correct time-walk effects
Integration in CMS central DAQ (GBT data link )