## Slide 1
TTC-PON test (CTP/LTU-Board)
CTP – PRR - Meeting
Luis Alberto Pérez Moreno (BUAP)
On behalf of the ALICE-CTP group
A Large Ion Collider Experiment
July – 11th - 2017.
Benemérita Universidad Autónoma de Puebla

## Slide 2
1
TTC-PON test (CTP/LTU-Board)
Outline
Introduction.
Evolution of the TTC_PON in ALICE CTP.
LTU_Firmware (Structure)
Experimental Setup 1 (CTP-LTU-Ev.Board)
Downstream latency measurements
BERT
A Large Ion Collider Experiment
Experimental Setup 2 (CTP-LTU-CRU)
Downstream latency measurements
BERT
July – 11th - 2017.
CTP – PRR - Meeting

## Slide 3
2
Introduction
A Large Ion Collider Experiment
July – 11th - 2017.
CTP – PRR - Meeting
Ref : The 10G TTC-PON :Challenges, solutions and performance (TTC.-PON team) , TWEPP 2016

## Slide 4
3
A Large Ion Collider Experiment
July – 11th - 2017.
CTP – PRR - Meeting
Ref : The 10G TTC-PON :Challenges, solutions and performance (TTC.-PON team) , TWEPP 2016
Introduction
  9.6Gbps, 6.4Gbps User Bandwidth maximum
    User Field: 20 bytes (160 bits)

## Slide 5
4
A Large Ion Collider Experiment
Ref : The 10G TTC-PON :Challenges, solutions and performance (TTC.-PON team) , TWEPP 2016
July – 11th - 2017.
CTP – PRR - Meeting
Introduction

## Slide 6
5
A Large Ion Collider Experiment
Ref : The 10G TTC-PON :Challenges, solutions and performance (TTC.-PON team) , TWEPP 2016
July – 11th - 2017.
CTP – PRR - Meeting
Introduction

## Slide 7
A Large Ion Collider Experiment
TTC_PON in ALICE CTP
6
LTU Logic
July – 11th - 2017.
CTP – PRR - Meeting
TTC-PON ver_0_0 in KC705
1 OLT. Ctrl/Cfg: Microblaze, SDK
1 ONU. Ctrl/Cfg: Microblaze, SDK
Testability: Counter, PRBS7, PRBS23
Downstream
TTC-PON ver_0_1 in KC705
1 OLT. Ctrl/Cfg: IPbus
1 ONU. Ctrl/Cfg: Microblaze, SDK
Testability: CTP_EMU, Counter, PRBS7, PRBS23
Downstream.
TTC-PON ver_0_1
1 OLT (KCU105). Ctrl/Cfg: IPbus
1 ONU (KC705). Ctrl/Cfg: Microblaze, SDK
Testability: CTP_EMU, Counter, PRBS7, PRBS23
Downstream & Upstream.
TTC-PON ver_0_2
1 OLT (KCU105) Ctrl/Cfg: IPbus
2 ONUs (KC705). Ctrl/Cfg: Microblaze, SDK
Testability: CTP_EMU, Counter, PRBS7, PRBS23
Downstream/Upstream
Include calibration process
TTC-PON ver_0_3
1 OLT (KCU105) Ctrl/Cfg: IPbus
use the FMC_TTC_PON.
2 ONUs (KC705). Ctrl/Cfg: Microblaze, SDK
Testability: CTP_EMU, Counter, PRBS7, PRBS23
Downstream/Upstream
Include calibration process
TTC-PON ver_0_4
9 OLT (KCU105) Ctrl/Cfg: IPbus
1 ONU (KCU105) Ctrl/Cfg: Ipbus
2 FMC_TTC_PON (OLT/ONU).
2 ONUs (KC705). Ctrl/Cfg: Microblaze, SDK
Testability: CTP_EMU, Counter, PRBS7, PRBS23
Downstream/Upstream
Include calibration process

## Slide 8
A Large Ion Collider Experiment
CTP_TTC-PON (Block-Diagram - Firmware)
Firmware V1
25 IPbus Slaves
8 OLTs
1 ONU
2 SPI ctrl
1 I2C ctrl
1 CTP_ Emulator
8 ILAs (Logic Analyzer )
7
July – 11th - 2017.
CTP – PRR - Meeting
“Thanks TTC-PON team for support”

## Slide 9
A Large Ion Collider Experiment
CTP_TTC-PON (Features of the TTC-PON system)
Firmware V1
25 IPbus Slaves
8 OLTs
1 ONU
2 SPI ctrl
1 I2C ctrl
1 CTP_ Emulator
8 ILAs (Logic Analyzer )
8
Ctrl & cfg. of all OLTs via IPbus
Ctrl & cfg. of the ONU via IPbus
Full Calibration.
Light Calibration.
Network Init.
Network Health.
Network Presence.
BERT Config. (Down/Up)
BERT Read. (Down/Up)
Downstream Latency
Status ONU
Status OLT
Toggling (TX & RX)
OLT Reset loop
TTC_PON test Menu (IPbus)
ONU_Ref_clock via IPbus (IPbus to I2C)
July – 11th - 2017.
CTP – PRR - Meeting

## Slide 10
A Large Ion Collider Experiment
Experimental Setup 1 (CTP to LTU to 2 Ev.boards as CRU)
9
July – 11th - 2017.
CTP – PRR - Meeting
Experimental Setup 1
CTP
LTU
KC705
KCU105

## Slide 11
A Large Ion Collider Experiment
Experimental Setup (CTP to LTU to 2 Ev.boards as CRU)
10
CTP/LTU board as CTP
LHC Clock
( LHC Machine Interface )
IPbus software
CTP/LTU board as LTU
IPbus software
Microblaze software
KCU105
KC705
Optical Link
Electrical Link
OLT
OLT & ONU
OLT & ONU
ONU
In
Out
In
Out
Out
TTC-PON: Time & Trigger Control – Passive Optical Network
CTP: Central Trigger Processor
LTU: Local Trigger Unit
OLT: Optical Line Terminal
ONU: Optical Network Unit
CRU: Common Readout Unit
July – 11th - 2017.
CTP – PRR - Meeting

## Slide 12
A Large Ion Collider Experiment
Experimental Setup (Calibration process)
11
Calibration Process (LTU to 2 Ev. Boards)
Result ONU Config:
OLT_RX (ILA) After calibration.
Result OLT Config.
July – 11th - 2017.
CTP – PRR - Meeting

## Slide 13
A Large Ion Collider Experiment
TTC_PON Downstream
12
July – 11th - 2017.
CTP – PRR - Meeting
Data_Gen
Regbank
OLT Core
bbert
OLT User
Regbank
Data_check
bbert
ONU Core
ONU User
OLT Match Flag
ONU Match Flag
FPGA
FPGA
Start
End
Downstream Latency
*
*
*
Similar structure for OLT/ONU
The downstream latency must be constant, after reset process
..01
OLT in counter mode
..02
..f8
..f9
..00
..01
Match Flag
..01
..02
..f8
..f9
..00
..01
Match Flag

## Slide 14
A Large Ion Collider Experiment
Experimental Setup ( Downstream Latency)
13
CTP/LTU as CTP
LHC Clock
( LHC Machine Interface )
IPbus software
CTP/LTU as LTU
IPbus software
KCU105
1- Start in OLT
2- Match Flag in ONU
Configuration Send data that is received from the ONU
“in the first PON layer”
OLT= Counter Mode
(CTP to LTU) ~ 11 m.
(LTU to KCU) ~ 11 m.
|
1
2
3
1-2 ~ 157 ns (- 55ns of cables) ~ 102 ns
1-3 ~ 318 ns (-110ns of cables) ~ 208 ns
------------------------------------------------
2-3 ~ 161 ns (-55ns of cables) ~ 106 ns
(CTP to Scope) = 1.5 m.
(LTU to Scope) = 1.5 m.
3 - Match Flag in ONU
(KCU to Scope) = 1.5 m.
In
In
Out
Out
CTP to LTU
CTP to KCU
LTU to KCU
Downstream Latency .
July – 11th - 2017.
CTP – PRR - Meeting

## Slide 15
A Large Ion Collider Experiment
Experimental Setup ( Downstream Latency after several resets)
14
CTP/LTU board as CTP
LHC Clock
( LHC Machine Interface )
IPbus software
CTP/LTU board as LTU
IPbus software
KCU105
1- Start in OLT
2- Match Flag in ONU
OLT= Counter Mode
(CTP to LTU) = 11 m.
(LTU to KCU) = 11 m.
|
1
2
3
(LTU to Scope) = 1.5 m.
3 - Match Flag in ONU
(KCU to Scope) = 1.5 m.
In
In
Out
Out
1-2 ~ 157 ns (- 55ns of cables) ~ 102 ns
1-3 ~ 318 ns (-110ns of cables) ~ 208 ns
------------------------------------------------
After several resets in CTP_OLT, LTU_OLT
Stability of the latency.
Configuration Send data that is received from the ONU
“in the first PON layer”
July – 11th - 2017.
CTP – PRR - Meeting

## Slide 16
A Large Ion Collider Experiment
Experimental Setup ( BER Test after calibration process)
15
Downstream 9.6 Gbps
Upstream 2.4 Gbps
CTP
LTU
KCU & KC7
LTU
BERT Downstream (CTP to LTU). No Errors ~ 2hrs
BERT Upstream (LTU to CTP). No Errors ~ 2hrs
BERT Downstream (LTU to 2Ev.Board). No Errors ~ 2:30hrs
BERT Upstream (KCU to LTU). No Errors ~ 2:30hrs
ONU_RX (ILA) PRBS lock and No errors.
BER_Downstream ~ 10 -14
BER_Upstream ~ 10 -13
July – 11th - 2017.
CTP – PRR - Meeting

## Slide 17
A Large Ion Collider Experiment
Experimental Setup (LTU standalone )
16
IPbus software
Microblaze software
KCU105
KC705
OLT & ONU
OLT & ONU
ONU
In
Out
Out
Downstream 9.6 Gbps
Upstream 2.4 Gbps
BERT Downstream (LTU to 2 ONUs). No Errors
BERT Upstream (ONUs to LTU). No Errors
After Calibration process
NO LHC CLOCK
July – 11th - 2017.
CTP – PRR - Meeting
Internal Clock
( Local Oscillator)
CTP/LTU board as LTU
BER_Downstream ~ 10 -14
BER_Upstream ~ 10 -13

## Slide 18
A Large Ion Collider Experiment
Experimental Setup 2 (CTP to LTU to 2 Ev.boards as CRU)
17
July – 11th - 2017.
CTP – PRR - Meeting
Experimental Setup 2
CTP
LTU
KC705
CRU 1
KCU105
CRU 2

## Slide 19
A Large Ion Collider Experiment
Experimental Setup (CTP-CRU)
18
CTP/LTU board as CTP
LHC Clock
( LHC Machine Interface )
IPbus software
IPbus software
Microblaze software
KCU105
Optical Link
Electrical Link
OLT
OLT & ONU
OLT & ONU
ONU
In
Out
In
Out
Out
KC705
Out
Out
ONU
CRU
CRU
ONU
CTP/LTU board as LTU
VLDB (GBTx)
(CTP to LTU) = 1 m.
(CTP to KCU) = 1 m.
(LTU to CRUs) = 11 m.
LTU
to
KC705
=
1 m.
(CRU to VLDBs) = 11 m.
TTC-PON: Time & Trigger Control – Passive Optical Network
CTP: Central Trigger Processor
LTU: Local Trigger Unit
OLT: Optical Line Terminal
ONU: Optical Network Unit
CRU: Common Readout Unit
July – 11th - 2017.
CTP – PRR - Meeting

## Slide 20
A Large Ion Collider Experiment
Experimental Setup (Calibration process)
19
LTU(OLT_RX )(ILA) After calibration.
Calibration Process (LTU to 2 CRUs and 1 KC705 )
Result ONU Config:
Result OLT Config:
KC705
CRUs
July – 11th - 2017.
CTP – PRR - Meeting

## Slide 21
A Large Ion Collider Experiment
Experimental Setup (BERT Results )
20
Downstream 9.6 Gbps
Upstream 2.4 Gbps
CRUs
LTU
CTP
LTU & KCU
BERT Downstream (LTU & 2CRUs & 1kc705 ). No Errors > 22hrs
CRU1
CRU2
BERT Upstream (CRUs to LTU). No Errors >22 hrs
CRU1
CRU2
BER_Downstream ~ 10 -15
BER_Upstream ~ 10 -14
CTP – PRR - Meeting
July – 11th - 2017.

## Slide 22
A Large Ion Collider Experiment
Experimental Setup (Downstream Latency After Several Resets)
21
CTP/LTU board as CTP
LHC Clock
( LHC Machine Interface )
IPbus software
IPbus software
Microblaze software
Optical Link
Electrical Link
OLT
OLT & ONU
ONU
In
Out
In
Out
KC705
Out
Out
CRU
ONU
CTP/LTU board as LTU
VLDB (GBTx)
(CTP to LTU) ~ 1 m.
(LTU to CRUs) ~ 11 m.
LTU
to
KC705
~
1 m.
(CRU to VLDBs) ~ 11 m.
1- Start in OLT
2- ONU Match flag
3- GBTx Match flag
1-2 ~ 105 ns (- 3 ns of cables) ~ 102 ns
1-3 ~ 590 ns (-120 ns of cables) ~ 365 ns (CRU + VLDB)
-----------------------------------------------------------------
After several resets in CTP_OLT, LTU_OLT
1
2
3
OLT= Counter Mode
CTP to LTU
CTP to eLink (VLDB)
Stability of the latency.
July – 11th - 2017.
CTP – PRR - Meeting

## Slide 23
A Large Ion Collider Experiment
End
22
July – 11th - 2017.
CTP – PRR - Meeting
Summary
We tested the TTC-PON system on the CTP/LTU board.
First version of the LTU firmware well advanced
The consistence of the downstream latency between CTP-LTU-CRU-VLDB tested.
Downstream Latency CTP-LTU ~ 102 ns
Downstream Latency of the full chain ~ 590 ns
BER between LTU-CRU tested.
  BER downstream ~ 10-15
  BER upstream ~ 10-14
Next Steps:
Include the GBT in LTU firmware
Long BER test
Thank you !!

## Slide 24
ALICE Central Tigger Processor (Upgrade)
22
Thank you !!
A Large Ion Collider Experiment
July – 11th - 2017.
CTP – PRR - Meeting

## Slide 25
ALICE Central Tigger Processor (Upgrade)
22
Back Up!!
A Large Ion Collider Experiment
June – 26th - 2017.
CTP - Upgrade - Meeting