## Slide 2: High-frequency impedance measurements of coated waveguides
A.Passarelli, A. Andreone, H. Bartosik, O. Boine-Frankenheim
M.R. Masullo, G.Papari, G. Rumolo, V.G. Vaccaro

## Slide 3: Goal
Study the electromagnetic properties of Amorphous Carbon coating in Sub-THz range.
DUT

## Slide 4: CLIC Damping Rings
Parameters | CLIC @3 TeV
Energy [GeV] | 2.86
Circumference [m] | 427.5
RF voltage [MV] | 5.1
Damping time x/s [ms] | 2/1
Bunch population [109] | 4.1
Hor. norm. emit. [nm.rad] | 500
Ver. norm. emit. [nm.rad] | 5
Bunch length [mm] | 1.8
Number of trains | 2
Bunches per train | 156
Bunch spacing [ns] | 1
The short rms σz=1.8 mm translates into
hundreds of GHz in the frequency domain.
Two DRs, one for
each species

## Slide 5: Electron Cloud
When electrons hit the pipe wall, they do not just disappear…..
  High energy electrons easily survive and actually multiply through secondary electron emission
  Low energy electrons tend to survive long because they are likely to be elastically reflected.
Secondary electron emission is governed by the curve below
Ep
q
secondaries
elastically reflected
The electroectromagnetic (EM) characterization of a-C properties up to is required for the impedance modelling of the CLIC DR components.
G. Rumolo, H. Bartosik and K. Li
USPAS Lectures

## Slide 6: Outline
Analytical evaluations
Measurement setup
Measurement results
Coatings
Conclusions
Next steps

## Slide 7: Waveguide
Square waveguide
copper conductivity: σcu = 6∙107 S/m
waveguide length: 7cm
waveguide side: 1.27mm

## Slide 8: Analytical evaluations
Waveguide side a [mm] | Frequency range [GHz]
1.27 | 120-170
a
The frequency range goes from fTE1,0= fTE1,0 to fTE1,1

## Slide 9: Analytical evaluations
Power attenuation per length unit α is defined by:
The attenuation α due to the losses in the walls is:
Zs : surf. char.imp.
Z : mode char.imp.
Htan : tang. mag. field
Ht : trans. mag. field
N.Markuvitz – Waveguide Handbook

## Slide 10: Analytical evaluations
Along the waveguide walls of length :
The attenuation for TE1,0 mode:

## Slide 11: Analytical evaluations
Cu
aC
d

## Slide 12: Analytical results
Good agreement between TD CST simulations and analytical evaluations.
copper conductivity: σcu = 6∙107 S/m aC conductivity:
waveguide length: 7cm
waveguide side: 1.27mm
What happens with two modes?
σac = 104 S/m

## Slide 13: Analytical evaluations
Along the waveguide walls of length :
The attenuation for both TE1,0 and TE 0,1 mode:
The attenuation is the same as the single mode case.
The losses on the walls for the two modes are just double of those of one mode alone as well as the flowing power.

## Slide 14: Analytical results
Bad agreement between TD CST simulations and analytical evaluations.
copper conductivity: σcu = 6∙107 S/m aC conductivity: σac = 104 S/m
waveguide length: 7cm
waveguide side: 1.27mm
If we rotate the structure and polarization?

## Slide 15: Analytical results
Good agreement between TD CST simulations and analytical evaluations
if we coherently rotate the structure and the polarization.
copper conductivity: σcu = 6∙107 S/m aC conductivity: σac = 104 S/m
waveguide length: 7cm
waveguide side: 1.27mm

## Slide 16: It is time to contact CST support

## Slide 17: Analytical results
copper conductivity: σcu = 6∙107 S/m aC conductivity: σac = 104 S/m
waveguide length: 7cm
waveguide side: 1.27mm
Good agreement between FD CST simulations and analytical evaluations.
If we coat a central foil?

## Slide 18: Analytical evaluations
Along the foil of length l:
The attenuation for both TE1,0 and TE 0,1 mode:

## Slide 19: Analytical results
copper conductivity: σcu = 6∙107 S/m aC conductivity: σac = 104 S/m
waveguide length: 7cm
waveguide side: 1.27mm
Good agreement between FD CST simulations and analytical evaluations.
We need two pyramidal transitions…

## Slide 20: Analytical evaluations
Along the waveguide walls of length :
The attenuation for both TE1,0 and TE 0,1 mode:
Pyramidal transitions to enhance the signal collection.

## Slide 21: Analytical results
copper conductivity: σcu = 6∙107 S/m aC conductivity: σac = 104 S/m
transition length: 3.7cm
transition side: 5mm → 1.27mm
Good agreement between FD CST simulations and analytical evaluations.
If we coat a central foil?

## Slide 22: Analytical evaluations
Along the foil of length l:
The attenuation for both TE1,0 and TE 0,1 mode:

## Slide 23: Analytical results
copper conductivity: σcu = 6∙107 S/m aC conductivity: σac = 104 S/m
transition length: 3.7cm
transition side: 5mm → 1.27mm
Good agreement between FD CST simulations and analytical evaluations.

## Slide 24: Analytical results
The signal attenuation along a central coated slab is equal to the one on coated walls.
Comparison between attenuation on walls and on foil for different thickness
The advantages of the setup with the central coated slab are:
  intrinsic simplification of the manipulation of all the setup.
  possibility to have a uniform deposition on the slab.

## Slide 25: Outline
Analytical evaluations
Measurement setup
Measurement results
Coatings
Conclusions
Next steps

## Slide 26: Measurement system
0
1
100
T-Light/
FS-Laser
ODU
PortA
PortB
Input fiber port
Output fiber port
Tx
Tera15-SL25-FC
Rx
Tera15-DP25-FC
Attenuator
Lens1
Lens2
Lens3
Lens4
2F
DDS Out1
PA310
Supply
Out
±15V
Vin
System specifications:
Spectral Range: > 3 THz
Dynamic Range > 70 dB
Scanning Range ~400ps
Spectral Resolution < 3GHz
N

## Slide 27: Outline
Analytical evaluations
Measurement setup
Measurement results
Coatings
Conclusions
Next steps

## Slide 28: Two different configurations
120mm
Lens1
Lens2
Lens3
Lens4
152mm
spot diameter 2.4 mm
< 5 mm aperture!
120mm
Lens1
Lens4
5 mm aperture!
8 mm
Adapted “confocal” geometry
Collimated geometry

## Slide 29: Comparison
Two geometries show similar results.
The collimated geometry is preferred for the planar wave similarity.

## Slide 30: Outline
Analytical evaluations
Measurement setup
Measurement results
Coatings (in collaboration with TE-VSC-SCC)
Conclusions
Next steps

## Slide 31: Coatings
DC magnetron sputtering system
C. Yin Vallgren
CERN-THESIS-2011-063

## Slide 32: Coatings
One side coating
Good thickness 5μm No Peel off
Sample under stress.
Two sides coating
Good thickness ~5μm
Peel off
No stress on foil

## Slide 33: Backup solution
Use the NEG Coatings as Eirini K.P. did in her thesis to validate our measurement method
E. Koukovini Platia
CERN-THESIS-2015-152

## Slide 34: Backup solution
We need from 4 to 5 μm of NEG coating thickness on copper foil

## Slide 35: Outline
Analytical evaluations
Measurement setup
Measurement results
Coatings
Conclusions
Next steps

## Slide 36: Conclusions
Good agreement between Analytical evaluation and CST FD solver.
The rhomboidal waveguides is used to have a smoother transition with pyramidal horns.
The collimated geometry is preferred for the planar wave similarity.
We will use a foil as long as the entire length of the guide + transitions.
Problems in the a-C coating process.
Evaluate NEG conductivity to validate measurement method.

## Slide 37: Next steps
Have 5 μm of NEG coatings on copper foil
Evaluate NEG conductivity

## Slide 38
Thank you!

## Slide 39
Backup slides

## Slide 40: CST support reply
Simulation considering the coating thickness of 1um.
The attenuation is similar for all the simulation in frequency domain and in the time domain with the waveguide aligned to the axis.
The simulation with the “rotated” waveguide show different results also increasing the number of meshes.

## Slide 41: Skin depth

## Slide 42: Analytical results
6/7 μm of aC coating thickness is needed to appreciate a signal attenuation.
copper conductivity: σcu = 6∙107 S/m aC conductivity: σac = 104 S/m
waveguide length: 7cm
waveguide radius: 0.9mm

## Slide 43: Analytical results
copper conductivity: σcu = 6∙107 S/m aC conductivity: σac = 104 S/m
waveguide length: 7cm
waveguide radius: 0.9mm
6/7 μm of aC coating thickness is needed to appreciate a signal attenuation.