## Slide 1: Destructive detectors for low energy ion beams
CERN BE-BI-EA
Sergey Sadovich
AVA School on Low Energy Antimatter Physics
25-29 June 2018
CERN

## Slide 2: Outline
1. Interaction of particles with matter
2. Faraday Cup
3. Scicnillators/screen
4. Silicon detector

## Slide 3: Interaction of particles with matter
Ionization
  Creation of electrons/ions pairs
    Secondary electron emission (low energy electrons)
    Emission of photons (decay of excited states)
Elastic and Inelastic scattering
  Dislocations
  Production of secondary particles (high energy particles)
Radiation
  Cherenkov radiation
  Bremsstrahlung
  Optical transition radiation

## Slide 4: Energy deposition

## Slide 5: Scintillation
E.Bravin - Transverse profile measurements
CAS Beam Instrumentation 2018

## Slide 6: Faraday Cup
Collector
Repeller
<1 pA detection

## Slide 7: FC: Example 1
Aperture 60x100 mm
Large aperture – up to 1kV on repeller
Grounded shielding around repeller

## Slide 8: FC: Example 2

## Slide 9: FC: Example 3
Aperture 140mm
High current/density/energy:
V shape of the FC collector
Coating of the FC collector with refractory metals
Water cooling

## Slide 10: FC: Signal extraction
Microphonic/triboelectric effect /ground loops – proper cables/cabling/connectors are required
Preamplifier (low noise, changeable gain) + ADC

## Slide 11: FC: application
Current
Profile/position
Emittance

## Slide 12: FC and Scanning slit (profile measurement)
detector
beam

## Slide 13: FC and Scanning slit (profile measurement)

## Slide 14: FC and 2 Scanning slits (transverse emittance)

## Slide 15: FC: pin-hole / array
1 mm aperture in molybdenum cup
0.2 mm aperture in tungsten foil
Faraday Cup collector

## Slide 16: FC: summary
Cheap and robust device for current measurement (< 1 pA), for lower values magic with improving signal to noise ratio is required
Updated functionality with slit (profile and position measurement)
Beam stopper – energy absorption / dissipation

## Slide 17: Silicon detector
Preamplifier (Operating as a charge to voltage converter) (0.45V per pC -> 20mV per MeV)
500 MeV for 0…10V digitizer
Single particle detection
Fragile/degradation

## Slide 18: Silicon detector

## Slide 19: Silicon detector
DPP-PHA
Digital Pulse Processing – Pulse Height Analysis

## Slide 20: Silicon telescope
K.Schmidt and O.Wyszynski
An evolutionary strategy for ∆E-E identification

## Slide 21: Silicon detector: TOF
7.76 m
time [ns]
N / NMAX
Bunch time structure for beams with different energies measured at SD1
time [ns]
N / NMAX
Bunch time structure for beams with different energies measured at SD2
The TOF system uses the time information provided by two silicon detectors separated 7.76 m from each other
Energy changes smaller than 0.5 % could be easily resolved
Bunch structure in second Si detector has partially degraded
Annular silicon detector

## Slide 22: Degradation of SD
Typical leakage current: 2-5 nA
Leakage current increases during opration
Degradation (not critical for relative measurement, can affect final measurements)

## Slide 23: Energy measurement: Pulse Height Defect
# Counts
Energy bin #
Energy spectrum for a beam with A/q = 4.0 and REX energy
132Xe33+
40Ar10+
20Ne5+
16O4+
4He1+
12C3+
Mass Number (A)
Ebin# / A
Deviation from linearity as a function of mass number
XT00.1000 deviation from linearity
Mass Number (A)
Ebin# / A
XT00.1900 deviation from linearity
Mass Number (A)
Ebin# / A
Cocktail of beams (EBIS and Xe from GPS target) was sent to silicon detector
PHD is visible for heavy ions – time of flight technique should be used for energy measurement

## Slide 24: Silicon detectors: summary
Absolute energy measurement of the beam (0.5%): roughly measurement using energy output and fine correction using timing output
Single particle detection
Fragile – can be destroyed in one direct pulse
Degradation during operation (corrections for data analysis are required / calibration using alpha source)
Corrections for PHD should be taken into account for data analysis

## Slide 25: Scintillators
Hao Zhang - Beam Diagnostics I (Destructive monitors)
F. Miyahara et all
Response of scintillating screens to high charge density electron beam
Widely used for beam profile measurements
Decay time
Radiation hardness
For chemical compound
Saturation with very dense beam (pC/mm2) is possible

## Slide 26: Scintillator: summary
Phosphors have very high light yields, but can only be used as thin coating on a rigid support and get damaged very quickly (often used in MCP)
Normally used for very low intensity beams
CHROMOX (Al2O3:Cr, Aluminum Oxide) is a very common choice because it is a very robust ceramic
ZZZ-doped YAG is also a very frequent choice (fast)
E.Bravin - Transverse profile measurements
CAS Beam Instrumentation 2018

## Slide 27: Micro Channel Plate
Saturation of the MCP (channels)
www.hamamatsu.com

## Slide 28: PepperPot
www.dreebit-ibt.com
Hao Zhang - Beam Diagnostics I (Destructive monitors)

## Slide 29: PepperPot: scanning plate

## Slide 30: PepperPot: image processing

## Slide 31: How to analyze images

## Slide 32: MCP summary
Sensitive tool for particles detection
MCP-PMT with a single photon detection
Fragile
Linearity and working point (to avoid saturation)
Image processing can be non-trivial

## Slide 33: The End