## Slide 1
Update on surface TCAD Radiation Damage Modeling in Silicon Detectors
F. Moscatelli(1,2), D. Passeri(3,1), A. Morozzi(1), G.-F. Dalla Betta(4), G.M. Bilei(1)
(1) INFN Perugia, Perugia, Italy
(2) CNR-IOM Perugia, Perugia, Italy
(3) Department of Engineering, University of Perugia, Perugia, Italy
(4) Department of Industrial Engineering-University of Trento and TIFPA,, Trento, Italy

## Slide 2: Outline
Motivations
Experimental Measurements (X-rays irradiation)
  Different test-structures (Gated Diodes, MOS capacitors, MOSFETs)
  Different providers (FBK, HPK, IFX) and substrates (p-type, n-type)
  Simulation Results and Comparison with experimental data
  Surface damage
    Update of the TCAD Si/SiO2 interface damage model

## Slide 3: AIDA2020 - Perugia
INFN and University of Perugia are involved in WP7.
7.2 TCAD model radiation damage:
- development of a TCAD model for the study of radiation damage effect in silicon detectors valid up to the particle fluences expected at the end of HL-LHC operations in the inner layers of the pixel systems (ATLAS/CMS).
Main goal synopsis:
  extension of the past “Perugia” TCAD model to simulate the effects of radiation damage on solid-state silicon devices;
  development of a comprehensive bulk damage as well as surface trap states build up modeling for a particle fluence > 2×1016 1MeV neq/cm2.
Participants: G.M. Bilei (INFN PG staff), D. Passeri (UniPG staff), F. Moscatelli (CNR-IOM staff), E. Fiandrini (UniPG staff), A. Morozzi (INFN post-doc).

## Slide 4: Work Plan
Measurements on dedicated test structures e.g. gated diodes, MOS capacitors and MOSFETs on p-stop and p-spray/different substrates.
Different technologies/providers (e.g. FBK, IMM, HPK, IFX …).
High-Frequency and Quasi-Stationary C, MOSFET and I-V characteristics, …
Irradiation campaign with gammas, X-rays and neutrons/(protons).
Measurements after irradiation → trap parameter extraction, TCAD model validation.
Predictive application of the model → sensor design and optimization
Added
To do.
Completed
On going
Bulk radiation damage modelling:
- extension of the three-level UniPG modelling (capture cross sections, charge multiplication, avalanche effects).
Surface radiation damage modelling:
- oxide fixed charge and interface trap state @dose;
- systematic study of acceptor/donor states at different energies and doses.
Comparison with literature data/dedicated measurements in terms of static behaviour (R, C) and charge collection properties.
Comprehensive modelling (bulk + interface, 2D/3D).
Measurements
Simulations

## Slide 5: Motivations
Simulate the effects of radiation damage on silicon devices at very high fluences (HL-LHC operation > 2×1016 1MeV neq/cm2).
Development of a combined bulk damage effect and interface trap states build-up numerical TCAD model:
    Physically-grounded approach,
    Extend the predictive capability to high fluences,
    Keep low the number of traps (e.g. avoiding fitting parameters / minimize the number of traps),
    Compatibility with already developed bulk damage model
    No over-specific modeling (e. g. device and technology independent)
Extraction from simple test structures of relevant parameters to be included within the model
Validation of the new modeling scheme through comparison with measurements of different test structures before and after irradiation.
Modeling the radiation damage effects.
Predictive insight of the behaviour of detectors, aiming at their performance optimization.
Goal

## Slide 6: New “University of Perugia” model
TEST STRUCTURE MEASUREMENTS
DETECTOR OPTIMIZATION
MODEL
MODEL VALIDATION
Setting-up the measurements:
Measurement Campaign: X-ray irradiation
  carried out in Padova (IT)
  doses range: 50 - 20 Mrad (SiO2)
Different devices, different providers, different substrates (p-on-n and n-on-p)
Measurements after irradiation / annealing 80°C 10 min.
PARAMETERS EXTRAPOLATION

## Slide 7: Test Structures
    Test Structures: MOS capacitors, Gated Diodes, MOSFETs, Strips (Rint meas.)
    Substrates: n and p
    Strip Isolation: p-stop and p-spray (for p-type substrate)
Al-gate MOS
capacitor
Gated-diode
MOSFET
Al-gate MOS
capacitor
Gated-diode
MOSFET
p-on-n
n-on-p
Gated-diode
MOS
capacitor
Cap-TS
for Rint
MOS
capacitor
Gated-diode
FBK
HPK
IFX

## Slide 8: FBK - MOS Capacitors: measurements
n-type substrate
HF measurements at 100 kHz with a small signal amplitude of 15 mV.
The QS characteristics were measured with delay times of 0.7 s using a voltage step of 100 mV.
Effective oxide charge density NEFF obtained from VFB measurements.
Unbiased devices during the irradiation steps.
Pre X-ray
Post X-ray

## Slide 9: FBK - MOS Capacitors: parameters extraction
The interface trap density (DIT) was estimated by using the C-V High-Low method.
Relatively low interface trap state density region
Non-negligible interface trap state density
Acceptor interface trap states evaluation
from -type substrates
Donor interface trap states evaluation
from -type substrates

## Slide 10: FBK – Gated Diodes
Unbiased devices during the irradiation steps.
From I-V measurements the surface velocity s0 was evaluated as a function of the dose.
p-type substrate
n-type substrate

## Slide 11: FBK - MOSFETs
x
x
x
+
Vth = -0.8 ÷ 0.1 V (unirradiated)
Unbiased devices during the irradiation steps
Radiation  interface traps (NIT) + trapped-oxide (NOX)  Vth shift (ΔVth).
ΔVth is separated into a contribution due to NIT and due to NOX, from IDS-VGS of MOSFET (method proposed in McWorther Applied Physics Letters 48, 133 (1986))
pMOSFET
nMOSFET

## Slide 12: Summary of measurements – p-type substrate
Integrated interface trap density (NIT)
Describe the Donor trap state characterisics
Effective oxide charge density (NEFF)
as INPUT PARAMETERS to the TCAD tool
+

## Slide 13: Summary of measurements – n-type substrate
Integrated interface trap density (NIT)
Effective oxide charge density (NEFF)
Describe the Acceptor trap state characterisics
as INPUT PARAMETERS to the TCAD tool
+

## Slide 14: IFX: MOS C-V Measurements after X-rays
VFB  -10 V at 50 krad
VFB  -17 V at 100 krad
VFB  -30 V at 500 krad
VFB-42 V at 1 Mrad
VFB-50 V at 10-20 Mrad
p-type substrate

## Slide 15: HPK and IFX: p-type Summary of measurements
NEFF and NIT :
  Extrapolated parameters with the same methodologies of FBK
  Similar values between the two vendors
p-type substrate

## Slide 16: IFX and HPK: p-type GCD after X-ray irradiation
Annealing 80°C 10 min
Surface velocity s0 evaluated as a function of the dose
Area 11.71 mm2
Area 6.14 mm2
IFX
HPK

## Slide 17: New “University of Perugia” model
TEST STRUCTURE MEASUREMENTS
DETECTOR OPTIMIZATION
MODEL
MODEL VALIDATION
PARAMETERS EXTRAPOLATION

## Slide 18: Surface Damage Model: Levels
EC - 0.40
EC
EV
EC - 0.60
EV + 0.70
Interface trap state energy modelling
Type | Peak Energy (eV) | Density (cm-2)
Acceptor | EC - 0.40 | 40% of acceptor NIT [1] (NIT=M·NOX)
Acceptor | EC - 0.60 | 60% of acceptor NIT [1] (NIT=M·NOX)
Donor | EV + 0.70 | 100% of donor NIT (NIT=M·NOX)
F. Moscatelli et al., Combined Bulk and Surface Radiation Damage Effects at Very High Fluences in Silicon Detectors: Measurements and TCAD Simulations, IEEE Transactions on Nuclear Science, 2016, Vol. 63, Issue: 5, 2716 - 2723.
Conduction Band
Valence Band

## Slide 19: Surface Damage Model: Gaussian
EC - 0.40
EC
EV
EC - 0.60
EV + 0.70
Interface trap state energy modelling
Type | Peak Energy (eV) | Density (cm-2) |  (eV)
Acceptor | EC - 0.40 | 40% of acceptor NIT [1] (NIT=M·NOX) | 0.07
Acceptor | EC - 0.60 | 60% of acceptor NIT [1] (NIT=M·NOX) | 0.07
Donor | EV + 0.70 | 100% of donor NIT (NIT=M·NOX) | 0.07
F. Moscatelli et al., Effects of Interface Donor Trap States on Isolation Properties of Detectors Operating at High-Luminosity LHC, IEEE Transactions on Nuclear Science, 2017, Vol. 64, Issue: 8, 2259 - 2267
Conduction Band
Valence Band

## Slide 20: Surface Damage Model: Uniform Bands
EC
EV
EV + 0.6 eV
EC - 0.56 eV
Interface trap state energy modelling
Acceptor Band
Donor Band
EV + 0.3 eV
Conduction Band
Valence Band
Type | Energy (eV) | Band width (eV) | Concentration (cm-2)
Acceptor | EC - 0.30 | 0.56 | Dit = Dit(Φ)
Donor | EV + 0.45 | 0.30 | Dit = Dit(Φ)

## Slide 21: FBK MOS Capacitors – Model validation
@ 50 krad
@ 100 krad
@ 10 Mrad
@ 20 Mrad
n-type substrate
Irradiated structures FBK p-type.
C-V measurements compared to simulations at different doses from 0.05 to 20 Mrad(SiO2).
  Good agreement!

## Slide 22: HPK MOS capacitors: simulations
50 krad
1 Mrad
20 Mrad
p-type substrate
Irradiated structures HPK p-type.
C-V measurements compared to simulations at different doses from 0.05 to 20 Mrad(SiO2).
  Good agreement!
  same TCAD surface model of FBK, but with different values (i.e. extrapolated from meas)

## Slide 23: IFX MOS capacitors: simulations
p-type substrate
Irradiated structures IFX p-type.
C-V measurements compared to simulations at different doses, from 0.05 to 20 Mrad(SiO2).
  Good agreement for IFX devices!
  same TCAD surface model of FBK, but with different values (i.e. extrapolated from meas)
50 krad
1 Mrad
20 Mrad

## Slide 24: Conclusions and future plans
  Extensive measurements campaign:
    X-ray with doses 0.05-20 Mrad(SiO2).
    Dedicated FBK, HPK and IFX p-on-n and n-on-p test structures before and after irradiation with X-rays.
  Surface radiation damage effects deeply investigated aiming at the extrapolation of the most relevant parameters:
    cross-check of NEFF, NIT, NOX evaluated by different methodologies from different test structures.
  Development of the radiation damage modelling scheme (bulk + surface), suitable for commercial TCAD tools (e.g. Synopsys Sentaurus).
    No over-specific modeling.
    Predictive capabilities of the “University of Perugia” TCAD model extended to HL-LHC radiation damage levels (e.g. Φ > 2.0×1016 1 MeV neq/cm2).
  Application to the analysis and optimization of different classes of (pixel) detectors to be used in the future HEP experiments (3D detectors, 2D planar detectors, timing …).