Data Acquisition Hardware
(Hands-on Approach)
        ISOTDAQ2010, February 1st
S. Ballestrero University of Johannesburg & CERN - W.Vandelli CERN/PH-ATD

---
Outline
- Introduction
- Measure energy deposition
  - Scintillator setup
  - Photomultiplier
  - Analog-to-Digital conversion
  - Charge-to-Digital conversion
  - QDC in real life
- Measure position
  - Wire chamber setup
  - Time-to-Digital conversion
  - TDC in real life
- Corollary

---
Introduction
- This wants to be a hands-on approach to the basic DAQ hardware
- We will discuss two different experiments, requiring different techniques and components
- We also have some good real data to discuss
  - You we will see, we are talking about real life here
- Let's get started!
Material and ideas have been taken from CERN Summer Student lectures of P.Farthouat, C.Joram and O.Ullaland

---
1st experiment

---
Energy measurement
- Measure the energy deposited by a particle traversing a (special) medium
- The (detector) medium is a scintillator → The molecules, excited by the passing particle, relax emitting light
  - The amount of light is proportional to the deposited energy
- The light is then collected, using dedicated optical means (light guide), and fed into a photo-detector → photomultiplier
Scintillator
Light
guide
Photomultiplier
HV
Signal

---
Photomultiplier
- Photo cathode: photon to electron conversion via photo-electric effect
  - Typical quantum efficiency ≈1-10% (depends on material and wavelength)
- Dynodes: electrodes that amplify the number of electrons thanks to secondary emission
  - Typical overall gain ≈106
- Dark current: current flowing in the PMT without light → noise
Anode
Photo
Cathode
   -HV
Dynodes
i(t)
-(4/5)HV
-(3/5)HV
-(4/5)HV
-(2/5)HV
-(1/5)HV

---
Getting started
b
i(t)
V(t)
h
R=50
20ns
100mV

---
Good old scope
- Approximate Q measurement using oscilloscope
  - Linear approximation of a exponential decay
b
i(t)
V(t)
h
R=50
20ns
100mV

---
Getting started
Done! (In less than 5 slides)
Approximate measurement
Deadtime ~3000%/Hz (if you are good)
Encode the data into some sort of electronic format by hand
Wouldn't be much more convenient to have a direct electronic measurement? It could save the data in some digital format and fill a histogram on-line. Wouldn't be cool?
The scope method is still fundamental: allows for the validation of your DAQ (and, yes, you should never thrust it a priori!)
Approximate Q measurement using oscilloscope
b
i(t)
V(t)
h
R=50
20ns
100mV

---
Analog to digital conversion: introduction
Digitization → Encode a analog value into a binary representation
Compare entity with a ruler
Entity to be measured
Ruler unit

---
Analog to digital conversion: Flash ADC
Digitization → Encode a analog value into a binary representation
Compare entity with a ruler
Entity to be measured
Ruler unit
Flash ADC
- Flash ADC simplest and fastest implementation
- Performs M comparisons in parallel
  - Input voltage is compared with M fractions of a reference voltage: Vref//(2M) → (M-1/2)Vref/M
- Result is encoded into a compact binary form of N bits
Encoder
start
R/2
#M
Vref
#1
Differential comparator
N=Log2 (M-1)
R
R
R
Vin
busy
R/2

---
Analog to digital conversion: Flash ADC
Digitization → Encode a analog value into a binary representation
Compare entity with a ruler
Entity to be measured
Ruler unit
Flash ADC
Example M=3 → N=2
Vin/Vref	Comparison / results	Encoded / form
<1/6	000	00
1/6≤ <3/6	001	01
3/6≤ <5/6	011	10
5/6≤	111	11
Encoder
start
R/2
#M
Vref
#1
Differential comparator
N=Log2 (M-1)
R
R
R
Vin
busy
R/2

---
ADC Characteristics
Digitization → Encode a analog value into a binary representation
Compare entity with a ruler
Entity to be measured
Ruler unit
- Resolution (LSB), the ruler unit: Vmax/2N
  - 8bit, 1V → LSB=3.9mV
- Quantization error, because of finite size of the ruler unit: LSB/2
- Dynamic range: Vmax/LSB
  - N for linear ADC
  - >N for non-linear ADC
    - Constant relative resolution on the valid input range

---
ADC phase-space
Many different ADC technique exists, mostly because of the trade-off between speed and resolution
Speed (sampling rate)
Flash
Sub-Ranging
Pipeline
Successive Approximation
Ramp
Sigma-Delta
GHz
Hz
6
22
bipolar
CMOS
Discrete
Power
>W
<mW
#Bits

---
ADC (In)Accuracies
ADC Transfer function
Integral & Differential
Non-Linearity
Extreme Differential
Non-Linearity

---
Charge to Digital → QDC
- ADC converts a voltage into a digital representation. However, in our experiment, we have a current and we are interested in the total charge
- QDC → Charge to Digital Converter
  - Essentially an integration step followed by an ADC
- Integration requires limits → gate

---
QDC and experiment 1
- ADC converts a voltage into a digital representation. However, in our experiment, we have a current and we are interested in the total charge
- QDC → Charge to Digital Converter
  - Essentially an integration step followed by an ADC
- Integration requires limits → gate
Gate
gen.
QDC
0xa3, 0x15, 0x8d, …
delay
We are done: we have a digital representation of:
Question: what should follow the QDC? With which aim?

---
QDC and experiment 1
- ADC converts a voltage into a digital representation. However, in our experiment, we have a current and we are interested in the total charge
- QDC → Charge to Digital Converter
  - Essentially an integration step followed by an ADC
- Integration requires limits → gate
dN/dq
Gate
gen.
QDC
0xa3, 0x15, 0x8d, …
delay
QDC count
We are done: we have a digital representation of:
Question: what should follow the QDC? With which aim?

---
QDC: Timing
Gate
gen.
QDC
delay
- Relative timing between signal and gate is important
  - Delay tuning
- Gate should be large enough to contain the full pulse and to accommodate for the jitter
  - Fluctuations are always with us!
- Gate should not be too large → increases the noise level
  - By the way, which is the noise contribution to our charge measurement?
Q: the tail is exponential,
how large is large enough?
f(x)
a
b

---
QDC: pedestal subtraction
clock
Gate
gen.
QDC
delay
Calibration data from a test beam @CERN
dN/dQ
- With an out-of-phase trigger we can measure the baseline (pedestal)
  - PMT dark current, thermal noise, …
  - Out-of-time particles
- The same noise enters our physics measurements and contributes with an offset to the distribution
- The result of a pedestal measurement has to be subtracted from our charge measurements
QDC count

---
A word on front-end electronics
Advanced FE
Complex FE
Impedance match
Noise suppression
Tail integration
Wide range (calorimetry)
Event tags, data blocks
Amplifier
Filter
Shaper
Range compression
clock
Sampling
Digital filter
Zero suppression
Buffer
Format & Readout
Buffer
Feature extraction
Detector
Pre-amplifier
Shaping
Peak finder
ADC
Format & Readout
Detector
Zero suppression

---
Real QDCs at work
- Real data from a beam test @CERN
- PbWO4 (scintillating) crystal equipped with two PMTs and exposed to e, and  beams

---
Real QDCs at work
- Real data from a beam test @CERN
- PbWO4 (scintillating) crystal equipped with two PMTs and exposed to e, and  beams
Nice pion-beam charge-distribution for one PMT
But, what are all those little peaks? Just statistical fluctuations? Let's zoom in!
QDC count
dN/dQ

---
Real QDCs at work
Bin with N entries shall fluctuate with
=√N
√360~18 → (540-360)/18=10
Spikes are regularly distributed
Some systematic effect must be taking place
Zoom in a bit more!
QDC count
dN/dQ

---
Real QDCs at work
415 & 416 → 0x19f & 0x1a0
431 & 432 → 0x1af & 0x1b0
447 & 448 → 0x1bf & 0x1c0
Can you see the effect?
The QDC prefers output configurations of type xxx0 in respect of those like xxxf
QDC count
dN/dQ

---
Real QDCs at work
415 & 416 → 0x19f & 0x1a0
431 & 432 → 0x1af & 0x1b0
447 & 448 → 0x1bf & 0x1c0
Can you see the effect?
The QDC prefers output configurations of type xxx0 in respect of those like xxxf
Homework: which is the simplest way to fix this problem in the data? At which cost? Can you understand the module name?
Module: 4c6543726f79204c31313832
QDC count
dN/dQ

---
2nd experiment

---
Position measurement
- This time we want to measure the position of particle with a wire chamber
- The ionization electrons created by the passage of the particle will take a time t to reach the anode wire
- Transit time is normally negligible with respect to t
- If we consider a constant drift speed vD (50m/ns), then position is:
x
Set
Q
Clear
clock → f
N bit
counter
reset
delay

---
Triggering
- The wire chamber alone is not sufficient however
- We need a triggering system
- Therefore we will measure a relative time t+t0 =t*
  - t0 accounts for the time delays, offsets, … between the wire chamber and the triggering system
x
Set
Q
Clear
clock → f
N bit
counter
reset
delay

---
Time measurement
- We can measure the time offset between the two signals using a N-bit digital counter driven by a clock of frequency f
- The wire signal acts as a start signal
- The trigger provides the stop signal
x
Set
Q
Clear
clock → f
N bit
counter
reset
delay

---
Time measurement → TDC
TDC
- Time-to-Digital Converter
- Resolution: 1/f
- Dynamic range: N
- Single hit TDC
  - e.g. a noise spikes comes just before the signal → measure is lost
x
Set
Q
Clear
clock → f
N bit
counter
reset
delay

---
Multi-hit TDC
clock → f
N bit
counter
N bit
counter
FIFO
gate
reset
hit
- Gate resets and starts the counter. It also provides the measurement period. It must be smaller than 2N/f
- Each “hit” (i.e. signal) forces the FIFO to load the current value of the counter, that is the delay after the gate start
  - In order to distinguish between hits belonging to different gates, some additional logic is need to tag the data
- Common-start configuration

---
Actual TDCs
Common-stop TDC
- Real TDCs provide advanced functionalities for fine-tuning the hit-trigger matching
  - Internal programmable delays
  - Internal generation of programmable gates
  - Programmable rejection frames
Can you imagine/sketch a common stop setup?

---
Real life wire chamber & TDC
- XDWC: delay wire chambers used on the SPS extracted lines to measure beam profiles
- Two cathode planes provide X and Y positions
- Measurement is based on the delay gained along a delay line
SL-Note-98-023

---
Principle
- XDWC: delay wire chambers used on the SPS extracted lines to measure beam profiles
- Two cathode planes provide X and Y positions
- Measurement is based on the delay gained along a delay line
trigger
top
t
bottom

---
Raw time data
BOTTOM
TOP
TDC count
TDC count
Individual channel distributions
LEFT
RIGHT
TDC count
TDC count

---
(Uncalibrated) Beam Profile
TOP-BOTTOM
RIGHT-LEFT
TDC count
- Beam sizes are still in TDC counts
  - Not very useful, though
- How do we convert this into a known scale (e.g. cm)?
TDC count

---
Corollary: calibration

---
Calibration
- Both the experiments we discussed provide relative measurements. The values obtained via our system are in some (known) relation with the interesting quantity
  - Scintillator
  - XDWC
- Our instruments need to be calibrated in order to give us the answer we are looking for
  - We have to determine the parameters that transform the raw data into a physics quantity
  - The parameters normally depend on the experimental setup (e.g. cable length, delay settings, HV settings, … )
- In the design of our detector and DAQ we have to foreseen calibration mechanisms/procedures

---
Ge crystal for isotope identification
Crystal HPGe
Trigger & front-end
Readout (ADC)

---
Crystal calibration
- 152Europium reference source
  - Known  emission lines
  - Peak find & fit
- Allow for definition of the parameters describing functional relation between ADC count and E
- Reality is not perfectly linear
  - Polynomial fits
  - Physical models

---
Isotope identification
Calibrate crystal setup can be used to identify isotopes generated in -irradiated samples

---
XDWC Calibration
- XDWC chamber have 3 calibration inputs that allow for independent calibrations of X and Y axis with only 3 different sets of data
- The calibration input simulate signals from particles respectively hitting
  - Right-top corner (X=Y=30mm)
  - Center (X=Y=0mm)
  - Left-bottom corner (X=Y=-30mm)
- The calibration data sets shall be taken with final setup and TDC
- Interpolating the three points in the t-x space, the parameters of the calibration equation can be measured

---
Calibrated XDWC
Beam profile
TOP-BOTTOM
Y (cm)
X (cm)
X (cm)
RIGHT-LEFT
Y (cm)

---
Wrap-up
- Digitization techniques produce data directly manageable by digital systems (e.g. a computer)
  - Greatly simplifies the down-stream data-handling
  - Root of every modern DAQ system
  - Available on a variety of platforms: VME, PCI, USB, …
- Open the “black box” and see where the numbers come from
  - Real electronics does not behave as the ideal one
- Trade-offs between speed/precision/cost exist
  - You have to choose the solution that best suits you
- Physics quantities are derived from the raw data via calibration
  - Calibration procedures have to be foreseen for your detector and DAQ

---
Example – NA59 Drift Chambers
- More complex, precise detectors need more sophisticated calibration
- Calibrate both “detector behavior” and “DAQ behavior” (eg TDC), sometimes at once.