## Slide 1: Fiber imaging system state of the work
Damiano Celeste, Stephane Burger, Enrico Bravin, Federico Roncarolo

## Slide 2: Project framework
Beam profile imaging nowadays: BTV – Vidicon
  Advantages:
    Most Rad Hard so far…
    No active components inside.
  Drawbacks:
    Limited life.
    Limited sensitivity (observation instruments).
    Issues of procurement (no production since > 5years, still ≈100 pieces in stock).
Beam profile imaging upgrade: imaging fiber system
  Principle: Transport an image away from a high radiation
  level area up to the camera located in a ‘safer’ place.
  Procurement of 2x2m and 2x10m long/bunch of 10k fibers.
  Advantages: Increased life time, improved image resolution(?).
Camera
*BTV principle, ack. Stephane Burger BE-BI-PM

## Slide 3: Index
Set-up characterization
  Target design and LED screen
  Comparison between 2 different lens orientation
  Light loss and attenuation:
    1900 CCD camera with 25mm lens @500mm distance from screen
    Two 25mm lenses and CCD camera
    Same setup but tested with 3 different optical fibers between the two 25mm lenses:
      Old 10m opt. fib. By (UNKNOWN)
      New 2m opt. fib. by Fujikura
      New 10m opt. fib. By Fujikura
    Same setup, changing the 1st lens: one 9mm lens and one 25mm lens tested with the same optical fibers.
  Reproducible measurements: comparison between two 2m fibers and two 10m fibers by Fujikura
  Resolution loss evaluation
  Comparison with eye-piece (f=24mm) in terms of light collection and spatial resolution
Fiber radiation hardness test: ongoing measurements and perspectives

## Slide 4: Target design and LED screen

## Slide 5: Target design and LED screen

## Slide 6: Picture of the setup: example with two 25mm lens
S1=517mm
S2=38mm
S3=42mm

## Slide 7: Study of the effect of 180° rotation of the 2nd lens
1st configuration
2nd configuration
Int light: 15858842 a.u.
Int light: 15993512 a.u.
<1%
Reference image:
Target @500mm
CCD camera
One 25mm lens

## Slide 8: Study of the effect of 180° rotation of the 2nd lens
Reference image:
Target @500mm
CCD camera
One 25mm lens
Chart: Pixels
From 38%
To 52%
Light profile and spatial resolution
Central square

## Slide 9: Beam profile imaging with optical fibers
Lens 1
Opt fiber in
Opt fiber out
Lens 2
CCD camera
Set up description:
LED screen at 500mm from the 1st lens (25mm or 9mm)
The optical fiber interface is placed at a distance d=f to catch the focalized image of the screen, created by the 1st lens.
A 2nd lens is placed at the fiber output to obtain a certain magnification factor of the image on the CCD camera.

## Slide 10: Acquired images: same MAG on camera
Reference image:
Cam + Lens 25mm
Cam + 2 Lens 25mm
Cam + 2 Lens 25mm
Old 10m opt. fiber
New 2m opt. fiber
New 10m opt. fiber
Cam + 1 Lens 9mm + 1 Lens 25mm
Old 10m
opt. fiber
New 2m
opt. fiber
New 10m
opt. fiber

## Slide 11: X-integrated light profile (a.u.)
#Pixel
Reference image one 25mm lens
Image with two 25mm lens

## Slide 12: X-integrated light profile (a.u.)
Reference image one 25mm lens
Image with two 25mm lens
a.u.
Old 10
New 2
New 10
#Pixel

## Slide 13: X-integrated light profile (a.u.)
Reference image one 25mm lens
Image with two 25mm lens
a.u.
Old 10m fib.
New 2m fib.
New 10m fib.
#Pixel
#pixel
a.u.
Old 10
New 2
New 10

## Slide 14: Effect of the presence of the optical fiber
Integral value:
1217222
319084
43197
35316
24443
1
2
3
4
5
Normalized integral value:
1
0.2621
0.0355
0.0290
0.0201
Central square
#Pixel
a.u.
Old 10
New 2
New 10
Reference
Two lens

## Slide 15: Light attenuation in fibers
#pixel
a.u.
New 2m
New 10m
#pixel
a.u.
-12.6%
-15.5%
Light loss ~ 14% in 8m (1.75%/m if linear -> 0.68 dB/km)
New 2m
New 10m
9mm lens
25mm lens

## Slide 16: Reproducible measurements (camera texp = 20 ms)
1) 6820955 a.u. 1) 5743142 a.u.
2) 7005164 a.u. 2) 5921904 a.u.
Differences < 3%
Comparison between two new 2m opt. fibers by Fujikura
Comparison between two new 10m opt. fibers by Fujikura
Differences < 3%

## Slide 17: Resolution loss: analysis of the central square
Cam + Lens 25mm
Cam + 2xLens 25mm
1 Lens 9mm + OptFib+ 1 Lens 25mm
2m New fiber
10m Old fiber
10m New fiber
1 Lens 25mm +
OptFib +
1 Lens 25mm
2m New fiber
10m Old fiber
10m New fiber
Conclusions:
The first main factor worsening the resolution is the presence of the 2nd lens: with respect to the mean value, we go from 90% to 34% approximately (graph A).
The presence of the optical fibers does not drive further resolution losses in the case with two 25mm lens (graph B).
In the case of the 9mm lens before the optical fiber, a further decrease of the spatial resolution is noticed: 16%
A
B
C

## Slide 18: Comparison with Eyepiece
Main difference:
The objective lens collects light and
brings it to focus creating an image.
The eyepiece is placed near the focal point
of the objective to magnify this image.
Main goals:
  Increase light collection
  Improve spatial resolution
Choice of the eye piece: f=24mm
Comparison with the 25mm lens

## Slide 19: Comparison with eye-piece
Set-up and measurements:
The 9mm lens is used as 1st optical input in the fiber.
The 2nd camera lens (f=25mm) is substituted with an eye-piece (f=24mm) and the distance with the camera is adjusted to obtain approximately the same MAG.
Two comparative measurements were performed with both the 2m and 10m new optical fibers, with texp=20ms.
Eye-piece 24mm
Camera lens 25mm
2m Opt fib
10m Opt fib
X-Profiles
Integral values [a.u.]
EyeP_2m: 9066809
Lens_2m: 7005164
EyeP_10m: 7515769
Lens_10m: 5921904
+ 29.4%
+ 26.9 %

## Slide 20: Comparison with eye-piece
Eye-piece 24mm
Camera lens 25mm
10m Opt fib
X-Profiles
Spatial resolution value:
From 20% to 30%
(in 2.5mm gaps)
Chart: Pixels
From 10% to 15%
(in 1mm gaps)

## Slide 21: Best images
Target @500mm, texp=20ms, 1st lens 9mm, 2nd lens 25mm
Target @200mm, texp=40ms, 1st lens 9mm, 2nd lens 25mm

## Slide 22: Conclusions optical study
Analysis | Results
Rotation 180° 2nd lens: light integration | No remarkable difference (<1%)
Rotation 180° 2nd lens: spatial resolution | Best with 2 lens with the same orientation (38%->52%)
Light loss due to 2nd lens | From 100% to 26.2%
Light loss due to Fujikura fiber interfaces | From 26% to 2.9%
Light attenuation in fibers | 0.65 dB/km
Reproducibility: light integration | No remarkable difference (<3%)
Spatial resolution loss due to two 25mm lens | From 90% to 34%
Spatial resolution loss due to fibers with 25mm lens | No remarkable difference from 34%
Spatial resolution loss due to fibers with 9mm lens | From 34% to 16%
Eye piece: light integration gain | Approximately +28%
Eye piece: spatial resolution gain | Approximately +50%

## Slide 23: Fiber radiation hardness test: on going
Simulating 10 years of operation
Dose integration of approximately 500kGy
On line results, taking into account light attenuation and spatial resolution
Choice of gamma irradiation
First tests in Co60 facility at CERN Bldg. 772
**Radiation-Resitant Single-Mode Optical Fibers, Zauhiko Aikawa et Al, Fujikura TR, 2008

## Slide 24: Fiber radiation hardness: mapping test set up
60Co
irradiator
window
source
Optical fiber
Radiation cone
d
Where d is the distance chosen to maximize the dose rate
Ion chamber
Ion chamber
35 cm
40 cm

## Slide 25: Choice of the distance
Distance [cm] | Dose rate [Gy/h]
72 | 4.16
68 | 4.29
66 | 4.73
65 | 4.54
63 | 3.12
60 | 2.30
55 | 0.98
Chart: 
| 72 | 68 | 66 | 65 | 63 | 60 | 55
Dr [Gy/h] | 4.16 | 4.29 | 4.73 | 4.54 | 3.12 | 2.3 | 0.98
Radiation
cone limit
Dose rate
≃ 1/r2
Region out of
radiation field
The distance of 66 cm between the fiber support and the source is chosen
Gy/h
cm

## Slide 26: Uniformity check
Position [deg] |  | Gy/10s | Gy/s | mean | std%
0 | 1 | 14.39 | 5.1804 | 5.27652 | 0.766482
 | 2 | 14.84 | 5.3424 |  | 
 | 3 | 14.64 | 5.2704 |  | 
 | 4 | 14.64 | 5.2704 |  | 
 | 5 | 14.6 | 5.256 |  | 
 | 6 | 14.69 | 5.2884 |  | 
 | 7 | 14.76 | 5.3136 |  | 
 | 8 | 14.61 | 5.2596 |  | 
 | 9 | 14.7 | 5.292 |  | 
 | 10 | 14.7 | 5.292 |  | 
For each point, 10 dose measurements integrated
over 10s were taken and then the avg and the std dev is calculated, as follows:
angle = 0° Dose Rate = 5.43 Gy/h
angle = 180° Dose Rate = 5.03 Gy/h
angle = 90° Dose Rate = 4.40 Gy/h
angle = 270° Dose Rate = 4.89 Gy/h
angle = 45° Dose Rate = 5.20 Gy/h
angle = 315° Dose Rate = 5.23 Gy/h
angle = 225° Dose Rate = 4.94 Gy/h
angle = 135° Dose Rate = 4.94 Gy/h
The obtained map is presented below, where the average dose rate is approximately 5 Gy/h over the all fiber. In this situation at the limit of the radiation cone and very close to the source, even 1-2 cm of displacement can drive a big difference in the dose rate.
The best obtained result, looking for a compromise between
the maximum achievable dose rate and the irradiation uniformity,
is ≃ 5Gy/h
angle = 0° Dose Rate = 5.27 Gy/h
angle = 180° Dose Rate = 5.05 Gy/h
angle = 90° Dose Rate = 4.56 Gy/h
angle = 270° Dose Rate = 4.71 Gy/h
angle = 45° Dose Rate = 5.14 Gy/h
angle = 315° Dose Rate = 5.07 Gy/h
angle = 225° Dose Rate = 4.98 Gy/h
angle = 135° Dose Rate = 4.99 Gy/h

## Slide 27: Starting irradiation test
Testing the new 10m fiber by Fujikura
10 TBq active Co60 source
Integrated dose over 3 weeks: 2 kGy
Reference image taken at the beginning
Acquired image every hour
Light-tight box with all optical components and supports
9mm lens
25mm lens
LED screen
Opt fib OUT
Opt fib IN
CCD cam
Opt fib
bundle

## Slide 28: Conclusions and next steps
Radiation hardness tests in Co60 lab:
  Mapping tests revealed the maximum uniform possible dose rate: 5 Gy/h
  These results showed us that 500 kGy cannot be achieved
  These tests are important to fix the set up for next measurements
New higher intensity facility
  LABRA laboratory in Saclay (up to 7 kGy/h) -> 47’000 euros for 96h
  …looking for other option with the EN-STI group
Final set up design and prototype