## Slide 1: Pushing the limits of hybrid pixel detectors in photon science with MÖNCH
Roberto Dinapoli :: Photon Science Detectors Group :: Paul Scherrer Institut
FEE 2018
Notes: MÖNCH is a 25 μm pitch charge integrating detector aimed at exploring the limits of current hybrid silicon detector technology.
It opens new perspectives for many synchrotron applications where currently the detector is the limiting factor e.g. inelastic x-ray scattering (IXS), Laue diffraction and soft X-ray or high resolution imaging.
The reduced input capacitance allows to obtain a low noise of less than 35 electrons RMS and the high segmentation provides high resolution images.
However, the small pixels present many challenges in terms of bump bond yield, electronics design and data throughput.
At this pitch, the charge produced by single X-rays is almost always shared between neighboring pixels.
This allows to interpolate between channels obtaining ~ micron spatial resolution, but on the other hand deteriorates the energy resolving power of the detector.
The possibility of exploiting the charge sharing to improve the spatial resolution requires detailed modeling and characterization, which can also help to retrieve the energy information.

## Slide 2: Overview
Hybrid detectors at PSI
MÖNCH introduction
Detector characterization
Applications
MÖNCH 1.0 for Athos
Conclusions & Outlook

## Slide 3: Hybrid Detectors @ PSI
Pixels
Single
Photon
Counting
Charge
Integrating
172 µm
AGIPD
PILATUS
200 µm
75 µm
EIGER
JUNGFRAU

## Slide 4: Hybrid Detectors @ PSI
Pixels
Single
Photon
Counting
Charge
Integrating
172 µm
AGIPD
PILATUS
200 µm
EIGER
JUNGFRAU
MÖNCH

25 µm
75 µm

## Slide 5: Challenges and applications
High resolution
Low noise
Large dynamic range
Bump bonding yield
100% charge sharing
Sensitive to leakage current
Many pixels per unit area
  Limited resources per pixel
  Data throughput
  Data storage
  Data treatment
25x25 µm2
Pixel

## Slide 6: Challenges and applications
High resolution
Low noise
Large dynamic range
HI-res imaging
Laue diffraction
(R)IXS
XRF
Color imaging
Soft X-rays
Electrons
Sensor studies

## Slide 7: Prototypes
UMC 110nm AE (Al only, shrunk technology, great quality/price)
Pixel 25x25 µm2
Area 4x4 mm2, 160x160 = 25.6kpixels
Five different pixel architectures
4 analog outputs
2 kHz frame rate (1Gb port)
MÖNCH 02
MÖNCH 01
JUNGFRAU01
Multipitch strip sensor
MÖNCH 01
“Guest” of Jungfrau01 MPW
Test structures
Seven single pixel channels
Wirebonded to a strip detector for calibration

## Slide 8: MÖNCH 03
Area 10x10 mm2, 400x400 = 160k pixels
Architecture optimized for very low noise (single photon)
32 analog outputs
Currently 2 kHz frame rate (10 Gb port)
Up to 6 kHz by design (960 Mpixel/s)
Region of interest possible (frame rate > 6kHz)

## Slide 9: Basic pixel architecture
CDS
Stage
Cfb
Detector
Test
Calibrate
CS
Preamp
Off-pixel
buffer
Column bus
Column
receiver
PIXEL
CDS
reset
Preamp
reset
Storage
Pads
Single ended
to differential
output driver
(1x/25cols)
MÖNCH 02:
Five different sub-blocks
Column selection
MUX

## Slide 10: Add-ons
CDS
Stage
Cfb
Detector
Test
Calibrate
CS
Preamp
Off-pixel
buffer
Column bus
Column
receiver
PIXEL
CDS
reset
Preamp
reset
Storage
Caps
Cf2s
Cfb
Low
Gain
(static)
Single ended
to differential
output driver
(1x/25cols)
Pads
Static gain selection
Column selection
MUX

## Slide 11: Add-ons
CDS
Stage
Cfb
Detector
Test
Calibrate
CS
Preamp
Off-pixel
buffer
Column bus
Column
receiver
PIXEL
CDS
reset
Preamp
reset
Storage
Caps
Pads
Single ended
to differential
output driver
(1x/25cols)
Selectable CDS Gain
Two storage caps
Column selection
MUX

## Slide 12: Add-ons
CDS
Stage
Cfb
Detector
Test
Calibrate
CS
Preamp
Off-pixel
buffer
Column bus
Column
receiver
PIXEL
CDS
reset
Preamp
reset
Storage
Caps
Pads
Single ended
to differential
output driver
(1x/25cols)
Selectable CDS Gain
Two storage caps
CDS can be bypassed
Column selection
MUX

## Slide 13: Bump bonding at PSI
In house:
Under-Bump-Metallization
Indium deposition
Bump bonding
Hit map from Ge Kα fluorescence
Hits per pixel distribution
Yield > 99.95%
Bad bump bonds
Neighbour pixels
Pilatus 172 μm
Eiger 75 μm
x5.3
Mönch 25 μm
x47
MÖNCH 0.3
1x1 cm2, 160 kpixel
25 µm pitch

## Slide 14: Gain and Noise
Distribution of RMS noise across the sensor
Water cooled, 25 µs integration time
S/N=1 at 130 eV
260 eV for a 2x2 cluster
Acquisition at different energies
  (e.g. Kα fluorescence)
Gain from linear fit
Noise from pedestal RMS
Mo
Ge
Cu
  Dynamic range up to ≈1800 ∙ 1keV γ, ≈150 ∙ 12keV γ

## Slide 15: Position interpolation
High resolution digital image
Low resolution analog image
Charge sharing
Division of charge between pixels is strongly position dependent
Can improve the spatial resolution
  Pitch ≈ charge sharing distance
    320 µm thick Si, 120V ≈ 20 µm
    Depends on the drift time:
      Applied bias voltage
      Absorption depth
        Sensor thickness
        X-ray energy
  High SNR
    Low noise
  Isolated photons
    Fast frame rate
    Low flux
A. Bergamaschi, ICXOM24
  Max. 100 photons/mm2/frame to separate single photons
  @10 kHz frame rate => 106 photons/mm2/s
Clustering+2D ETA algorithm

## Slide 16: Does it work?
100 um
100 um
Gold on Silicon sample by M. Lebugle, Acquired at MS beamline, 8keV, with MÖNCH
Eiger Mönch Jungfrau

## Slide 17: High resolution imaging with MÖNCH
25 µm pixels
Interpolated
SOLEIL, Nanoscopium
11 keV
Siemens star
2 µm bins
Synchrotron
W-tube
55kV
EIGER HDI
5 µm bins
X-ray tube
SLS, PHOENIX
2 keV
61µm pitch metal grid
1 µm bins
Low energy
100µm

## Slide 18: Spatial resolution
The spatial resolution is higher in the corner between pixels than in the center of the pixels due to limited charge sharing
Energy dependent
  Higher energies, higher SNR
  Lower energies, larger charge sharing
Possible improvements
  Smaller pitch
  Thicker sensors
  Lower bias voltage
6 keV
Pencil beam scan
experiment
@ cSAXS beamline, SLS
Horizontal resolution
Vertical resolution
3.3
1.2
3.3
1.4
µm
µm

## Slide 19: Laue diffraction
Photon counts
# of photons
Energy
Energy (keV)
White beam at
microXAS beamline, SLS
Stressed Si wafer
1 ph.
2 ph.
3 ph.
Simultaneous position and energy measurements
http://pd.chem.ucl.ac.uk

## Slide 20: Color Imaging
K-edge subtraction
Fluorescence detection and rejection
Removal of beam hardening
Gold structures on silicon wafer fabricated at PSI with silicon microspheres
Measured spectrum of
W-tube
2x2 clusters
Ca. 30 s

## Slide 21: G2-less grating interferometry with MÖNCH
No stepping
  Lower dose
  Shorter exposure
  Simpler mechanics
Avoid G2 manufacturing issues
6 µm pitch gratings
16.7 keV beam @ TOMCAT
S. Cartier et al., JOURNAL OF SYNCHROTRON RADIATION 23, 1462-1473 (2016).
Absorption
Differential Phase Shift
Dark field
1 mm
Notes: The fringes are integrated in the direction of the gratings to improve statistics

## Slide 22: Towards MÖNCH 1.0
SPECIFICATIONS (as good as physicists can give them to you)
Energy range: 250eV-2keV
Single photon resolution
Dynamic range: 10000 x 700eV photons (Qmax=1.95Me-), data quality limited by poissonian fluctuations of the source
>50% quantum efficiency down to 500 eV
Total area required: 2 x (4x3cm2)
Frame rate 100 Hz (God bless the Swiss)
Baseline: the low energy ATHOS beamline at the SwissFEL needs a detector.
  ASIC rad-tolerance not a problem
  Frame rate and data throughput less demanding
  Hi spatial resolution not needed

## Slide 23: Possible ASIC and detector arrangements
= Beam position
Large area MÖNCH (Low noise, high dynamic range)
  ASIC 2 x 3 cm2 (768 x 1024 channels ~800k pixel)
  Modules up to 4 x 3 cm2 (~ 1.6M pixel)
  Almost 1G pixel after interpolation (not needed for ATHOS)

## Slide 24: Soft X-ray diffraction
Direct
beam
Bragg
peaks
Magnetic
peaks
Soft X-rays
  Single photon resolution down to about 800 eV
Large dynamic range
  Simultaneous imaging of direct beam, Bragg peaks and low intensity signal

## Slide 25: Soft X-ray Ptychography at PolLux
MÖNCH 03 at PolLux beam line at PSI
Out of focus beam
zero order suppressed
The target is scanned with overlapping regions
Shape is reconstructed from the diffraction patterns created (low flux)
Sample: 40 nm pitch grating
1 keV photon energy
2 ms exposure time/frame
1000 frames per point
Total photons per pixel
# of photons
Soft X-rays
Fast frame rate
Large dynamic range

## Slide 26: Noise
Improve control/readout system
LGAD sensor (collaboration with FBK)
Optimize preamp gain/CDS gain
Multisampling: Noise= Noise/ sqrt(Nsamples)
  Off chip-Off pixel (bottom of the column)-On pixel
CDS
Stage
Pixel
buffer
CDS
reset
Storage
Caps, N=4

## Slide 27: Dynamic gain switching
Single ended
to differential
output driver
(1x/25cols)
CDS
Stage
Cfb
Detector
Test
Calibrate
CS
Preamp
Off-pixel
buffer
Column bus
Column
receiver
PIXEL
CDS
reset
Preamp
reset
Storage
Caps
Pads
Vth
S
R
DGS
Q
reset
Cfb2
Cfb
Column selection
MUX

## Slide 28: Dynamic gain switching
Vth
S
R
DGS
Q
reset
Cfb2
Cfb
preamp
Already implemented in Gotthard,
Jungfau and AGIPD. However:
Very low power and area budget
Not implemented in Moench 0.3
Only 1 switching feasible: ratio of capacitors crucial
Additional pixel output (digital)
Channel without DGS
Channel with DGS
ADC#
Test pulse
DGS activates
Pixel and readout chain connected in “Test pulse mode”

## Slide 29: DGS at work with an IR laser
DGS mechanism for synchrotron and FEL applications
  single photon resolution
  high dynamic range
JUNGFRAU
AGIPD GOTTHARD
One half-supercolumn of MÖNCH 02 features a preliminary DGS design
Characterization is ongoing (FEL beam now available)
DGS supercolumn under IR laser
Pixel 92,16

## Slide 30: Droop
The chip frame rate will be lower, more droop. Biggest contributor: gate leakage current of off pixel buffer.
  Input transistor of Off-pixel buffer: p-type, smaller
  Additional switch
Off-pixel
buffer
Storage
Caps
CDS
Stage
CDS
reset
STO 1
MMC
STO 2
NCAP25
Bypass 1: switch kept closed
Bypass 2: switch operated

## Slide 31: Droop: pedestal shift
 Pedestal = Pedestalt – Pedestalref
Comparing the pedestal of each single pixel at different waiting times for the 6 cases
Results: droop is halved operating the switch
    2 caps help
    NCAP leaks more
To be done:
Same measurements but with signal on caps
MMC+NCAP25
Switch operated
NCAP25 only
Switch not operated

## Slide 32: Thin entrance window
Quantum Efficiency
Sensor backplane illuminated
Entrance window reduces quantum efficiency
Sensor not optimized for low energies
1.5 µm thick entrance window
entrance
window
photon
pixels
P implant
P diffusion
As implant (low energy)
As implant
Collaboration with FBK
Tests ongoing on different types of backside (20-200nm thickness; LGAD)
Irradiation at 1400 eV (bias 90V)

## Slide 33: Vacuum compatibility
We have a vacuum compatible system to test MOENCH 0.3 for electrons detection
System for ATHOS will be very different
Vacuum from 10-5 mbar downward
Heat dissipation of the electronics

## Slide 34: Conclusions and outlook
A 25 µm pixel pitch hybrid detector is both feasible…
  Good bump-bond yield, Low noise, Large dynamic range
  CHALLENGING (still work in progress)
… and useful
Many applications already tested:
  HI-res imaging; Laue diffraction; (R)IXS; XRF; Color imaging; Soft X-rays
  More to come (Electrons, Sensor studies)
We are working on a detector system for Swissfel (low energy)
    Specifications defined, down to 250eV
    Large area ASIC
    Hi dynamic range with DGS
    Thin backplane sensors (FBK), 20-200nm entrance window
    Vacuum compatibility
    Capability to handle big data amount (for use at synchrotrons)

## Slide 35: The PSI Photon Science Detector Group
M. Ruat, B. Schmitt, S. Redford, A. Mozzanica, E. Fröjdh
J. Zhang, C. Lopez-Cuenca, M. Andrä, R. Barten, M. Brückner, C. Ruder, D. Greiffenberg, S. Vetter
X. Shi, D. Thattil, G. Tinti, A. Bergamaschi, M. Ramilli, R. Dinapoli, D. Mezza, S. Chiriotti-Alvarez

## Slide 36: Thank you!

## Slide 37: Thank you

## Slide 38: Clustering
With charge integrating readout in single photon regime, we can retrieve the full charge by clustering
  Photon counting –like performance
  Preserves spectral information
    Suppresses charge sharing
    Noise is increased depending on the size of the cluster
MÖNCH 25 µm
JUNGFRAU 75 µm
MÖNCH 3x3 25 µm
Raw data
Clusters

## Slide 39: Iterative correction method
The bins of the η-distribution histogram are assigned to position bin within the pixel
  η flat distribution gives flat flatfield image
Iterative algorithm defines η bins map to obtain flat distribution
Algorithm converges once the flat field image is flat
η bin map
Flat field
1
10
20
30
40

## Slide 40: Nanoscope imaging
Magnification: ~80x
No Zernike phase dots
Si111 and 40 μm Cu
Siemens star test images
MÖNCH02
25μm pixel
interpolation
5 x stitching
not cooled
Photonic Science CCD
4μm pixel
2 x stitching
cooled
80μm

## Slide 41: Charge sharing
The charge diffuses while drifting
  The charge can be collected by several pixels
  Charge sharing ≈ 20 µm in 320 µm thick Si, 120V bias
    Depends on the drift time:
      Bias voltage,
      Sensor thickness,
      X-ray energy
X-ray
Pixel
Photon counting detectors lose counts in the pixel corners
Photon counting does not make sense without charge summation for pitches below 40 µm
Artefacts when structures smaller than pixels size
  Corner effect in protein crystallography
20keV pencil beam scan
Photon counting
Charge integrating

## Slide 42: Linear charge collection model
d
p
Full charge collection
Charge sharing
Sides
Corners
Noise convolution
α : dchargeSharing /pitch
Q0 : total charge
Q : collected charge
σ : electronic noise ENCrms
A : number of photons

## Slide 43: Dynamic range for MÖNCH 03
Highest Gain Setting
Highest Gain Setting
Dynamic range:
600 3 keV photons at lowest gain setting
18 3 keV photons at highest gain setting
  non-linearity due output buffer saturation
  optimize chip settings
IR laser
ND filters
Lowest Gain Setting
3 keV photons
Kapton thickness

## Slide 44: Energy resolution
Compton
Edge
1.45 keV
Main beam
20 keV
Si escape
18.2 keV
In
3.3 keV
Ti
4.5 keV
Cu
8.0 keV
Au
9.7 keV
Rb?
13.3 keV
No spectral information from single pixel with no cuts
Clustering increases the noise
Cuts to charge sharing as “offline collimation”
  99% rejection rate
Use larger pixels
MÖNCH02 20keV monochromatic data.
The fluorescence lines come from the detector itself.
Ni
7.4 keV
Br
11.9 keV

## Slide 45: Subpixel interpolation
T/(T+B)
R/(R+L)
The signal is fully contained in a 2x2 pixels cluster
We analyze the ratio between neighbors
  =
  =
Custom algorithms must be developed to correct for non linear charge sharing
Notes: If we look more in detail, the charge sharing is usually in a 2x2 pixels cluster. To interpolate, we analyze the ratio of the right pixels compared to the total for the horizontal coordinate, and of the top to the total for the vertical, which we call eta x and y respectively.
What happens is that if a photon come in the center of the bottom right pixel, the eta will be located in the bottom right position, if it come between two pixels in vertical of horizontal the eta will be about 0.5 in that coordinate and if the photon comes in the corner between the pixels, the eta will be in the center of the distribution plot.
If charge sharing would be linear and uniform, the eta would represent the relative position within the cluster. However some more advanced algorithm is required to correct for the non uniform charge sharing, in particular for the regions in the center of the pixels where there is no charge sharing, as you can see from the peaks in the eta distribution.

## Slide 46: Energy calibration
Developed linear model to deal with 100% charge sharing
Gain used to convert pedestal distribution into noise in e-
Width of the photon peak given by
ENC and Fano Noise
Dynamic range and linearity still to be fully characterized

## Slide 47: Grating-based phase contrast imaging
Phase shift
Creates periodic
intensity
modulations
G2 Analyzer grating
Makes fringes
visible on large
detector pixels
Darkfield
Absorption
No sample
Sample
A detector with sufficient spatial resolution and SNR allows to remove G2
Lower dose
Shorter exposure
No stepping
Larger area
Higher energies
~ 5 um
G1 Phase grating

## Slide 48: Omnidirectional grating interferometry
M. Kagias, et al. (2016) Phys. Rev. Lett. 116, 093902
Circular gratings can be detected using MÖNCH
  25 µm pitch, 5 µm period
Tests with a sample to detect DPC not yet performed
100 µm
A. Bergamaschi, XNPIG 2017
Notes: X-ray grating interferometry (GI) can provide simultaneously three complementary contrasts: absorption, differential phase, and small-angle scattering [1–3] also known as dark-field. The phase signal is highly sensitive to the electron density variations in the sample and can reveal differences between materials with similar absorption properties. The dark-field signal is able to access unresolved structural variations of the sample in the (sub)micrometer scale, which is beyond the resolution capability of the imaging modality. It has been demonstrated that both differential phase and dark-field signals can provide valuable information additional to the conventional absorption contrast in medical imaging [4–6], material science, and nondestructive testing [7]. Especially, the dark-field signal has drawn great attention due to its success in providing quantitative or inaccessible structural information in radiographic applications [8–12]. In general, the dark-field signal exhibits highly directional behavior if the underlying microstructure of the sample is anisotropic. However, up to now GI has mainly been composed of linear gratings and the dark-field sensitivity is therefore only perpendicular to the grating lines. To obtain multiple-direction dark-field sensitivity, either the sample or the interferometer needs to be rotated [13–14], which is a time-consuming procedure. The usage of 2D gratings can mitigate the issue but can only provide dark-field sensitivity in up to four directions. Moreover, the noise performance is not the same in all directions due to the different modulation orders of the phase stepping curves [15]. An alternative speckle scanning technique [16–17] is proposed to sense the dark-field signal by scanning a membrane in the direction of interest; however, only two directions can be obtained. To cope with samples containing structures with an unknown orientation, it would be favorable to achieve an omnidirectional dark-field sensitivity with a straightforward imaging setup, without requiring motion of the sample or the optical elements. This would ensure that all microstructures, highly oriented or not, can be detected. In this work, we propose a single-shot imaging method with 2D-ominidirectional dark-field sensitivity using a grating design consisting of a 2D repetition of a small pitch circular grating. To achieve single-shot imaging, the self-image of the grating is recorded with sufficient spatial resolution. Local fringe analysis is then performed to obtain the dark-field signal of the sample in all directions of the imaging plane. In addition, the differential phase contrast (DPC) in the vertical and horizontal directions can also be retrieved.

## Slide 49: Laue-Indexing of complex samples
Photon counts
# of photons
Energy
Energy (keV)
White beam at
microXAS beamline, SLS
Stressed TiNi alloy
Sample is translated to illuminate different grains
No rotation required
http://pd.chem.ucl.ac.uk
Ni
Ti

## Slide 50
twares XMAS and LaueTools
Am = T R A0
Elastic distortion matrix
Angular rotation matrix
Undistorted orientation matrix
Ge
Ice, G. E., & Pang, J. W. L. (2009). Tutorial on x-ray micro-Laue diffraction. Materials Characterization, 60(11), 1191–1201.
White-beam Laue diffraction
Diffraction pattern identified by:
Peak position
Peak energy
Notes: And, well, just a few words about The Laue Method, to understand why and how Laue can give us the information on the crystilline nature, deformation and orientation of materials, here goes a very very short explanation:
Basically described by the Bragg’s law;

## Slide 51: Pilot test with Resonant Soft X-ray Scattering
Circular +
Charge/frame [ADCu]
Scattering at 708 eV (Fe LIII-edge) – used to probe spin and angular momenta
Bragg peaks (high intensity)
Other structures due to resonant magnetic scattering (low intensity: ~ 1ph/s/pixel)
Integral(C+)/Integral(C-) = 1.07
can detect a ~5% variation in the Bragg peaks intensity
C+/C-
CCDs are commonly used but:
Bragg peaks need to be blocked to avoid damage to the detector
Masks specific for each experiment must be built and carefully placed
Circular +
3σ noise cut