## Slide 1: Towards a new generation of Medipix and Timepix ASICs
  X. Llopart,
  On behalf of the Medipix Collaborations
22nd May 2018
Front-End Electronics 2018

## Slide 2: Medipix3 Collaboration
University of Canterbury, Christchurch, New Zealand
CEA, Paris, France
CERN, Geneva, Switzerland,
DESY-Hamburg, Germany
Albert-Ludwigs-Universität Freiburg, Germany
University of Glasgow, Scotland, UK
Leiden University, The Netherlands
NIKHEF, Amsterdam, The Netherlands
Mid Sweden University, Sundsvall, Sweden
IEAP, Czech Technical University, Prague, Czech Republic
ESRF, Grenoble, France
Universität Erlangen-Nurnberg, Erlangen, Germany
University of California, Berkeley, USA
VTT, Information Technology, Espoo, Finland
KIT/ANKA, Forschungszentrum Karlsruhe, Germany
University of Houston, USA
Diamond Light Source, Oxfordshire, England, UK
Universidad de los Andes, Bogota, Colombia
University of Bonn, Germany
AMOLF, Amsterdan, The Netherlands
Technical University of Munich, Germany
Brazilian Light Source, Campinas, Brazil

## Slide 3: Medipix3RX (2011)
14100 µm
15880 µm
Sensitive Area (14080 µm)
Application | Photon counting Colour imaging
Technology | IBM 130nm DM 4-1
Pixel size | 55 x 55 / 110 x 110 µm2
Pixel arrangement | 256 x 256
Acquisition modes | Single Pixel Mode (55/110) Charge Summing Mode (55/110)
Readout Type | Frame-based (up to 24-bit) Continuous RW (up to 12-bit)
Minimum threshold | ~ 500 e- (SPM) ~ 1000 e- (CSM)
Power consumption | <1W @1.5 V
Number of Thresholds | Frame-based 2/4/8 Continuous RW 1/4
Floorplan | 3 sides buttable and minimum periphery
TSVs possibility | YES. With 0.8mm periphery
Max Count Rate [E.Frojdh JINST/2014] | 614 x 106 hits/mm2/s (SPM 55) 114 x 106 hits/mm2/s (CSM 55) 130 x 106 hits/mm2/s (SPM 110) 29 x 106 hits/mm2/s (CSM 110)
Output bandwidth | 1 to 8 LVDS @ 320MHz each

## Slide 4: Measurement (60keV, 110µm pitch, 2mm CdTe)
T. Koenig
~4.4KeV FWHM
Fluorescence photons are included in charge sum if their deposition takes place within the volume of the pixels neighbouring the initial deposition
No offline correction, all chip hits added up
Energy resolution limited by the pixel-to-pixel gain mismatch and linearity

## Slide 5: Count rate measurements 110x110 µm pixel
Energy resolution degraded by count rate
110µm pitch, 2mm CdTe
Measurements: T. Koenig
Notes: The idea behind this slide is that in order to have a “reasonable” spectrum the incoming flux should be at least 5 times lower than the dead time.

## Slide 6: Timepix3 (2013)
14100 µm
16210 µm
Sensitive Area (14080 µm)
Application | General Purpose particle tracking
Technology | IBM 130nm DM 4-1
Pixel size | 55 x 55 µm2
Pixel arrangement | 256 x 256 (2x4 superpixels)
Acquisition modes | Time (TOA) AND Charge (TOT) Time (TOA) PC & integral charge (iTOT)
Readout Type | Data driven (Shutter-less) Frame-based (Shutter)
Thresholds | 1
Minimum threshold | > 500 e-
Time resolution (TOA) | 1.562 ns
Energy Resolution (TOT) | ~2 keVFWHM
Power consumption | <1.5W @1.5 V
Floorplan | 3 sides buttable and minimum periphery
TSVs possibility | YES. With 1.2mm periphery
Count Rate | Data-Driven: ~0.43 x 106 hits/mm2/s Frame-based: 826 x 106 hits/mm2/s
Output bandwidth | 1 to 8 SLVS DDR @640Mbps each

## Slide 7: Trigger-less event-by-event data driven and zero-supressed readout
Achievable count rate:
  uniformly distributed events → ~40 Mhits/s/cm2 @5.12Gbps
Full matrix readout: ~800 µs @5.12Gbps
Acquisition time
Address[16-bit]
0xB
Data[28-bits]
Data Packet (48 bits)
48bit
48bit
48bit
48bit
48bit
48bit
End of Command (48 bits)
Shutter
Qin
DataOut
ChipID [32b]
0x71
0xB0

## Slide 8: Pixel Operation in TOA & TOT
Tpeak < 25ns
TOT (10 bits) =4
Preamp Out
Disc Out
Clk (40MHz)
FTOA (4 bits)=7
VCO Clk (640MHz)
TOT Clk (40MHz)
TOA (14-bit)
16383
X
TOA (14 bits)=16383
16384
0
2
3
4
16383
16382
1
Pixel Readout Starts (475ns→ 19 clock cycles)
Global TOA (14-bit)

## Slide 9: Am241 Spectrum with a Si 300µm
Chart: [Ke-]
Chart: [Ke-]
Threshold ~= 1.8 KeV
Threshold ~= 9 KeV

## Slide 10: Velopix (2016)
14140 µm
16600 µm
Sensitive Area (14080 µm)
Application | LHCb Vertex Locator upgrade
Technology | CMOS 130nm
Pixel size | 55 x 55 µm2
Pixel arrangement | 256 x 256 (2x4 superpixels)
Acquisition modes | Time (TOA) PC or charge (TOT) 6-bits
Readout Type | Data driven (Shutter-less) Frame-based (Shutter)
Thresholds | 1
Minimum threshold | > 500 e-
Time resolution (TOA) | 25 ns
Power consumption | <1.5W @1.2 V
Floorplan | 3 sides buttable and minimum periphery
TSVs possibility | YES. Multi-dicing scheme as Medipix3
Count Rate | Data-Driven: ~5 x 106 hits/mm2/s
Output bandwidth | 4 serializers @ 5.12 Gbps each

## Slide 11: 5.12 Gbps Velopix High Speed link
Full link 3-chip module with all links at 5.12 Gbps over a 1m copper line

## Slide 12: ClicPix2 (2017)
Application | CLIC Vertex detector prototype
Technology | CMOS 65nm
Pixel size | 25 x 25 µm2
Pixel arrangement | 128 x 128 (2x8 superpixels)
Acquisition modes | Time (TOA) 8-bits and Charge (TOT) 5-bits PC or charge (iTOT) 13-bits
Readout Type | Full Frame Zero compression (pixel, super-pixel and column skipping)
Thresholds | 1
Minimum threshold | > 500 e-
Time resolution (TOA) | 10 ns
Power consumption | < 80mW @1.2 V (power-pulsing 156ns/20ms)
Floorplan | 3 sides buttable and minimum periphery
TSVs possibility | NO
Readout | < 800 µs
Output bandwidth | 1 CML @ 640 Mbps
3340 µm
3700 µm
Sensitive Area (3200 µm)

## Slide 13: Pixel detectors ASICs developed @CERN
Chart: CMOS process [nm]
| 600 | 250 | 250 | 130 | 130 | 65 | 130 | 130
 | 0.013841 | 0.17521 | 0.18182 | 0.56198 | 0.89256 | 1.7578 | 0.6281 | 0.45455
Medipix3RX (2011)
Medipix1 (1998)
Timepix (2006)
Timepix3 (2013)
ClicPix (2013)
ClickPix2 (2016)
VeloPix (2016)
Dosepix (2010)
Medipix2 (2004)

## Slide 14: Through Silicon Via (TSV)
The goal is to develop a reliable "Via-last" TSV (Though Silicon Via) process for the Medipix Hybrid Pixel Detectors
The main benefit is to limit the need of classical wire bonding to sensor biasing only  Large area tiling with high percentage of active area
Medipix3RX, Timepix3 and Velopix are TSV compatible chips
TSV process development by CEA-LETI
Sensor fabrication and flip-chip assembly ADVACAM
CERN Medipix TSV project supported by AIDA, CERN-LCD and Medipix3 collaboration
“TSV last for hybrid pixel detectors: Application to particle physics and imaging experiments”, D. Henry, et al, (ECTC), 2013 IEEE 63rd

## Slide 15: TSV Process
MEDIPIX3 pixel side native thickness
TSV processed chip
“BGA” bottom distribution
SEM cross section of TSVs (CEA-LETI)
First 2x2 TSV Medipix3RX (ESRF)

## Slide 16: Medipix4 Collaboration
Medipix4 Collaboration is set to provide the next generation of Medipix4 family chips (Medipix4 and Timepix4)
Main characteristics:
  Use of a commercial 65nm CMOS technology
  4-side buttable architecture  periphery integrated inside the pixel matrix
  Performance improvements from previous generation:
    Medipix4: Count rate, gain linearity
    Timepix4: Count rate, TOA and TOT resolution
  Larger ASICs
  Digital-On-Top design methodology
Agreement signed on May 2016
Today already 14 groups

## Slide 17: Medipix4 Collaboration
CEA, Paris, France
CERN, Geneva, Switzerland,
DESY-Hamburg, Germany
Diamond Light Source, Oxfordshire, England, UK
IEAP, Czech Technical University, Prague, Czech Republic
JINR, Dubna, Russian Federation
NIKHEF, Amsterdam, The Netherlands
University of California, Berkeley, USA
University of Houston, USA
University of Maastricht, The Netherlands
University of Canterbury, New Zealand
University of Oxford, England, UK
University of Geneva, Switzerland
IFAE, Barcelona, Spain (In signature process)

## Slide 18: Pixel detectors ASICs developed @CERN
Chart: CMOS process [nm]
| 600 | 250 | 130 | 130 | 65 | 130 | 130 | 65
 | 0.013841 | 0.18182 | 0.56198 | 0.89256 | 1.7578 | 0.6281 | 0.45455 | 1.7578
Medipix3RX (2011)
Medipix1 (1998)
Timepix (2006)
Timepix3 (2013)
ClicPix (2013)
ClickPix2 (2016)
VeloPix (2016)
Dosepix (2010)
Medipix4 Timepix4

## Slide 19: Timepix3  Timepix4
 |  |  | Timepix3 (2013) | Timepix4 (2018/19)
Technology |  |  | CMOS 130nm – 8 metal | CMOS 65nm – 10 metal
Pixel Size |  |  | 55 x 55 µm | 55 x 55 µm
Pixel arrangement |  |  | 3-side buttable 256 x 256 | 4-side buttable 512 x 448
Sensitive area |  |  | 1.98 cm2 | 6.94 cm2
Readout Modes | Data driven (Tracking) | Mode | TOT and TOA | 
 |  | Event Packet | 48-bit | 64-bit
 |  | Max rate | < 43 Mhits/cm2/s | 178.8 Mhits/cm2/s
 | Frame based (Imaging) | Mode | PC (10-bit) and iTOT (14-bit) | CRW: PC (8 or 16-bit)
 |  | Frame | Zero-suppressed (with pixel addr) | Full Frame (without pixel addr) CRW (8-bit / 16-bit) Up to 44 KHz frame @8b
 |  | Max count rate | 82 Ghits/cm2/s | ~800 Ghits/cm2/s
TOT energy resolution |  |  | < 2KeV | < 1Kev
Time resolution |  |  | 1.56ns | ~200ps
Readout bandwidth |  |  | ≤5.12Gb (8x SLVS@640 Mbps) | ≤81.92 Gbps (16x @5.12 Gbps)
Target global minimum threshold |  |  | <500 e- | <500 e-
3.5x
4x
2x
8x
10x
33%
Timepix4: A 4-side tillable large single threshold particle detector chip with improved energy and time resolution and with high-rate imaging capabilities

## Slide 20: 4-side buttable: Motivation
Sensor
ASIC
PCB
Build large area detectors by combining smaller modules

## Slide 21: 4-side buttable: Challenge
Sensor
ASIC Pixel Matrix
ASIC Periphery
FC
Wire Bonding
PCB
FC
BGA
ASIC
PCB
Sensor
TSVs
The through-silicon vias (TSVs) are the key technology for this paradigm shift

## Slide 22: 3-side vs 4-side
Advantages:
  Build large area detectors without dead area
  Better power distribution on chip (TSV)  Larger single ASICs
  Mechanically more robust (no wire-bonds)
Disadvantages:
  Periphery hidden or integrated in the pixels  RDL between sensor and pixel input
    RDL  Increase input capacitance and crosstalk (~30-50fF)
  Top metal layers “blocked”  No MiMCaps, Inductors…
  Devices are only available after TSV processing (no wire-bonds)

## Slide 23: Timepix4 Floorplan
Chip size 28.16mm x 24.64 mm (no Wirebonds)
512 x 448  229376 pixels
Pixel size 55µm x 55µm
Analog Periphery (< 800 µm):
  BandGap + Temperature sensor
  Biasing DACs
  Monitoring ADC
  Analog supply
  Digital supply
2 x Digital Periphery ( < 400 µm):
  8 x 5.12Gbps serializers (configurable)
  PLL(s)
  Analog supply
  Digital supply
2 x Pixel matrix (13.28 mm x 24.64mm):
  256 x 448 pixels 55 µm x 51.875 µm
  5.68% smaller than 55 µm x 55 µm
  RDL to compensate up to 400 µm

## Slide 24: On-chip pad to pixel RDL
Use of 10 metal option (1p9m + RDL)
Use of internal metal layers for vertical redistribution:
  Equalized Cin for all pixels  ~40 fF increase in Cin for a 400 µm periphery
  Shielding of RDL layers to minimize/eliminate cross-coupling
RDL routing over analog pixel circuitry and periphery  M6 used as shield
Chart: Periphery Height [um]
| 640 | 512 | 409.6 | 384 | 256
average_b(worst) | 61.2932 | 48.9508 | 39.4025 | 36.9546 | 24.6663
average_b(best) | 67.3216 | 53.7899 | 43.2739 | 40.5797 | 27.0889
average_b(typ) | 64.3105 | 51.3507 | 41.3451 | 38.7735 | 25.8736
Chart: row
~0.1fF/µm [0.1/1.2/GNDA]
384 µm periphery  55 µm/52 µm pixel

## Slide 25: Timepix4 Pixel Schematic
Global threshold
Front-end
Leakage
Current
compensation
Preamp
Input
pad
1 pixel
VCO
@640MHz
Super pixel
Control
voltage
8 pixels
640MHz
TpA
TpB
TestBit
MaskBit
Front-end
Counters
&
Latches
Time stamp
14-bits
Synchronizer
&
Clock gating
OP Mode
5-bit Local
Threshold
~50mV/ke-
2.5fF
3fF
Sync
Readout
Data out
to EOC
8-bits
clock
(40MHz)
8-bits
4
8-bits
ADB
32 pixels
8
Clock
local
 | bits
ADDR | 9
ToA | 16
ufToA_r | 4
ufToA_f | 4
fToA_r | 5
fToA_f | 5
ToT | 12
Pileup | 1
SPG

## Slide 26: Pixel Operation in TOA & TOT [DD]
65534
X
TOA (16 bits)=65534
Tpeak < 25ns
TOT (10 bits) =4
Preamp Out
Disc Out
Clk (40MHz)
VCO Clk (640MHz)
TOT Clk (40MHz)
TOA (16-bit)
65535
0
2
3
4
65534
65533
1
Global TOA (16-bit)
FTOA_R (4 bits)=7
FTOA_F (4 bits)=11
0
X
UFTOA_Start (4-bits)
8
X
UFTOA_Stop (4-bits)

## Slide 27: On-pixel < 200ps time resolution
FE + discriminator jitter < 50 psrms
Column dDLL to distribute the 40 MHz
  Controlled skew on the stop signal (<100ps)
  Minimizes noise from pixel matrix clock distribution
Share an on-pixel 640 MHz VCO among 8-pixels:
  Oscillation frequency locked (Vcntrl) with periphery PLL for PVT control
  1.56 ns resolution (as in Timepix3)
  195 ps obtained latching the internal VCO phases
Chart: input jitter [ns]
| 0 | 0 | 10 | 21 | 29 | 40 | 47 | 50 | 52 | 60 | 62 | 71 | 72 | 78 | 89 | 95 | 102 | 110 | 121 | 147 | 198 | 237 | 301 | 354
Tpix3 measured TOA resolution | 433 | 433 | 433 | 433 | 436 | 434 | 434 | 435 | 436 | 437 | 438 | 437 | 436 | 439 | 442 | 441 | 442 | 445 | 449 | 451 | 481 | 486 | 510 | 556
Tpix3 measured TOA resolution | 65 | 65 | 64 | 68 | 70 | 74 | 79 | 80 | 79 | 88 | 87 | 94 | 93 | 101 | 110 | 115 | 121 | 126 | 139 | 159 | 208 | 245 | 308 | 360

## Slide 28: Timepix4 FE
The Timepix4 chip front-end has to be versatile in order to be used:
  With different types of sensors:
    Bipolar detection and <10nA leakage currents (high-Z materials)
  In tracking and as well in imaging:
    Imaging at high rates (e.g. synchrotrons) imposes a minimum pulse width. In order to be at the state of the art in terms of rate capability, the dead time of the front-end should be smaller than ~200ns.
    Tracking applications impose an input jitter in the front-end below 50psrms (time bin 200ps)
FE architecture follows the “classic” Medipix approach (Krummenacher FE)
  Gain adjustment (log gain in hole collection)
  Improved for low jitter and fast count-rate when required  power
Qin=2.2ke-
IKRUM=20nA

## Slide 29: Timepix4 FE summary
 | e- collection | h+ collection | h+ collection (log gain)
Gain | ~50 mV/ke- | ~50 mV/ke- | ~25 mV/ke-
ENC (@Cin=50fF) | ~60 e-rms | ~60 e-rms | ~65 mV/ke-
Minimum threshold | < 400 e- | < 400 e- | < 450 e-
TOA Jitter | <40 ps rms Qin > 10Ke- |  | 
TOT linearity | < 250 ke- | < 200 ke- | < 800 ke-
Pixel analog power | <7.5uA (@1.2V, 9 µW) |  | 
ENC vs Cin slope ~0.3 e-/fF
ENC vs Ileak slope ~4 e-/nA

## Slide 30: dDLL for the reference pixel clock distribution
Design of a digital delay-locked loop (dDLL) to control the clock phases (40 MHz) that arrive in every SPG (Super Pixel Group, 2x16 pixels)  Robust to PVT variations
1 complete dDLL every double column: 32 Delay stations over ~2.7cm (!)  Target station delay 781.25 ps
Digital design:
  Custom designed cells characterized using Liberate using 4 corners  Full double column verification possible (functional, dynamic power)
Chart: Delay code
BCspectre | 0.27757 | 0.33289 | 0.3165 | 0.37181 | 0.73063 | 0.78595 | 0.77124 | 0.82655
TCspectre | 0.40701 | 0.47999 | 0.46629 | 0.53927 | 1.102 | 1.175 | 1.1647 | 1.2376
WCspectre | 0.68415 | 0.79703 | 0.78717 | 0.90004 | 1.9021 | 2.0149 | 2.0124 | 2.1253

## Slide 31: Super Pixel Group dDLL
Area <0.75% of 16x2 SPG
Clock down
Clock up
~ 800 µm
70 µm

## Slide 32: dDLL controller (Periphery)
Synthesized using tcb65lp library
Digital “noise” filter  8 consecutive clocks with same UP or DN update direction
Lock detector with 8-bit programmable threshold
Programmable lock window
Bypass coarse and fine control words (not yet implemented)
DLL Controller
ckin
reset
bypass
bypassCode [7:0]
ckin
ckout
thresholdLock [7:0]
lockStatus
lockWindow [1:0]
coarse[3:0]
fine[3:0]
35.4 µm
19.8 µm
preliminary

## Slide 33: Digital-domain simulation @ TC corner with noise
Chart: [ns]
Chart: [ns]
Chart: [ns]
Chart: [ns]
0
25
50
75
125
150
175
200psrms
100
Δ(ideal-simulated)
Δ(ideal-simulated)
lockStatus
coarse & fine
time [µs]
* Specification for the absolute jitter of input clock is 5 psrms

## Slide 34: dDLL Power breakdown
22.7 mW/cm2 to distribute a 40MHz clock with a 100 psrms
Dynamic digital power consumption is distributed across the clock period
This not include all the local on-pixel clock distribution
Timepix3 @40MHz clock distribution is ~200mW (almost 4 times the area)
Block | Total power | Total
Single controller (TYP corner) | 32 µW | 
All controllers (1 per Double Column) | 224x2x20 µW | 14.3 mW
Single double Column delay line (TYP corner) | 16x20 µW | 
All double Column delay lines | 224x2x320 µW | 143.3 mW
Full chip dDLL | 224x2x352 µW | 157.6 mW

## Slide 35: On-Pixel TDC
Full digital Superpixel TDC implementation:
  VCO:
    full-custom designed (as in Timepix3) and characterized using Liberate
    640 MHz nominal frequency with 4-bit frequency adjustment (~40MHz range)
    Nominal consumption ~600 µW (enabled)
  Full digital integration with the rest of digital pixel
stage | Phase[3] | Phase[2] | Phase[1] | Phase[0]
0 | 1 | 1 | 1 | 1
1 | 1 | 1 | 1 | 0
2 | 1 | 1 | 0 | 0
3 | 1 | 0 | 0 | 0
4 | 0 | 0 | 0 | 0
5 | 0 | 0 | 0 | 1
6 | 0 | 1 | 1 | 1
7 | 0 | 0 | 1 | 1
VCO Oscillation pattern

## Slide 36: Timepix4 Super Pixel (4X2) TDC
33.6 µm
104 µm
VCO
8 pixels TDC:
  Hit synchronization
  Clock40 clock gating
  FTOA counters:
    5-bit FTOA_rise
    5-bit FTOA_fall
  UFTOA latches
    4-bit UFTOA_rise
    4-bit UFTOA_fall
~25% of available digital area

## Slide 37: Post-layout TDC simulations
VCO only enabled when required (~600 µW/hit/50ns)
@Max rate (178 Mhits/cm2/s)  ~15 mW average
Pix[0]
Pix[1]

## Slide 38: Post-layout TDC simulations (TYP)
Chart: [ns]
Chart: [ns]

## Slide 39: TDC post-layout digital simulations
Chart: samples
Chart: pixel number
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7
 | -10.2355 | -35.8355 | 12.5555 | 4.1255 | 13.3485 | 1.8585 | -15.6035 | 29.7865
Post-layout simulation (TYP)  57 psrms TOA resolution (VCO LSB/sqrt(12))
  Ideal FE
  Ideal 40 MHz
Simulation of 1 input random arrival
Limited pixel-to-pixel insertion delay mismatch < 80pspeak-to-peak

## Slide 40: Medipix4 Improvements
 | Medipix3RX | Medipix4 (2019)
Technology | CMOS 130nm | CMOS 130nm
Pixel Size | 55/110 µm | 70/140 µm
Pixel arrangement | 3-side buttable 256 x 256 | 4-side buttable ≥256 x 256
Thresholds | 2/4/8 Sequential RW 1/4 Continuous RW | Not yet defined
 | Sequential RW & Continuous RW | 
Maximum Count Rate | 376 x 106 ph/mm2/s (SPM 55) 29 x 106 ph/mm2/s (CSM 110) | 100 x 106 ph/mm2/s (CSM 140)
Dynamic Range | 25 Ke- | 32 Ke-
CSM energy resolution | ~ 4KeV (CSM 110) | ~ 2KeV (CSM 110)
Readout bandwidth | <1.6 Gbps (8x LVDS) | 1,2 Serializers @5.12 or 2.56 Gbps
Power | <1W @1.5V | <1W/cm2 @1.2V
Medipix4: A 4-side tillable multiple threshold charge summing imaging chip with improved energy resolution and count rate

## Slide 41: Expected Count Rate Medipix4 and Timepix4
Target Timepix4 in Imaging
Target Timepix4 in Data-Driven
Target Medipix4
R.Ballabriga, SpecXray 2015

## Slide 42: Conclusions
Hybrid Pixel Detectors (HPD) profit better from the continuous miniaturization in CMOS technologies independently from the sensor developments
The use of 3D-IC bonding techniques will bring to live 3D developments and large area detectors with no dead area
The Medipix4 Collaboration aims to design 2 4-side buttable ASICs:
  Medipix4 (2019/20)
    Pixel <70/140 µm
    Improved energy resolution (½) and count-rate (x5) compared to Medipix3RX
    Number of thresholds not defined
  Timepix4 (2018/19):
    Particle tracker AND Imaging chip with 1 threshold
    Improved energy resolution (x2), arrival time resolution (x8) and count-rate (x4) compared to Timepix3

## Slide 43: Acknowledgments
Medipix2/3/4 Collaborations
Linear Collider Detector group
LHCb Collaboration

## Slide 44: Frame Based CRW: [Minimum output bandwidth before counter overflow]
Chart: Bandwidth (Mbps) half chip
8-bit @ 1Gc/mm2/s > 10.24 Gbps
16-bit @ 8Gc/mm2/s > 640 Mbps
16-bit @ 1Gc/mm2/s > 80 Mbps

## Slide 45: TSV Process
Wafers thinned to 50-120 µm (limits the via aspect ratio)
TSV and RDL (Re-Distribution Layer) patterning
Final passivation and back side UBM (Under Bump Metallization)
Wafer is removed from temporary carrier and place on another specifically adapted for dicing and ejection
Sensor bump bonding to TSV processed chip
CEA-LETI

## Slide 46: TSV Yield
After foundry production and TSV added process the final yield is 45% of fully functional chips
Original wafer probing results (Foundry yield)

## Slide 47: CERN Medipix TSV PROJECT – First run – X-Ray Imaging
Hybrid Pixel Detector was positioned in front of the X-Ray beam
A biological sample (fish) placed before the detector
X-Ray chamber 35kV, 1mA
First irradiation setup using a Medipix 3.1 TSV assembly
First image obtained with a TSV processed hybrid pixel detector (flat field corrected)
R.Ballabriga and J.Alozy

## Slide 48: Latest TSV results
120 µm Medipix3RX with TSVs and bumped to a 500 µm thick P-on-N edgeless sensor
Sr90 Acquisition
400nA @150V
Full functional device !
Advacam
J.Alozy

## Slide 49: Jitter versus input charge
CIN=25fF-50fF
CL=15fF
CM=3fF
gm=90uS (I~3uA)
Noise ~70e- r.m.s.
Notes:
Input is a delta dirac
Slope of the time waveform is taken at maximum (at t=0) (as if threshold set to 0)
No limitation for the slope of the signal (no slew rate effects)
Preamplifier output (Discriminator not accounted for)
Conclusions:
For 5ke- signal (and 50fF CIN) the limit is 50ps
For 10ke- signal (and 50fF CIN) the limit is 24.5ps

## Slide 50: Jitter versus input charge
CIN=25fF-50fF
CL=15fF
CM=3fF
gm=90uS (I~3uA)
Noise ~70e- r.m.s.
Notes:
Slope of the time waveform is taken at maximum (at t=0) (as if threshold set to 0)
No limitation for the slope of the signal (no slew rate effects)
Preamplifier output (Discriminator not accounted for)
Conclusions:
For 5ke- signal (and 50fF CIN) the limit is 50ps
For 10ke- signal (and 50fF CIN) the limit is 24.5ps
Preamplifier output simulation (threshold at ~800e- (negative polarity)

## Slide 51: Jitter versus input charge
CIN=25fF-50fF
CL=15fF
CM=3fF
gm=90uS (I~3uA)
Noise ~70e- r.m.s.
Notes:
Slope of the time waveform is taken at maximum (at t=0) (as if threshold set to 0)
No limitation for the slope of the signal (no slew rate effects)
Preamplifier output (Discriminator not accounted for)
Conclusions:
For 5ke- signal (and 50fF CIN) the limit is 50ps
For 10ke- signal (and 50fF CIN) the limit is 24.5ps
Preamplifier output jitter simulation (threshold at ~800e- (negative polarity)
Comparator output jitter simulation