## Slide 1: Performance of a digital-SiPM readout chip.
Inge Diehl, Karsten Hansen, Christian Reckleben, Yuan Yao
DESY-FEC, Hamburg, Germany
11th Front-End Electronics Meeting
Jouvence, 20. May – 25. May 2018

## Slide 2: Readout concept: hybrid approach
dSiPM-array
(HLL-dSIMPl)
ultra-thin sensor + fast readout ASIC
sensor with small bulk resistor – fast recharging
low over-bias voltage (3.3-V mosfets)
pixel-wise readout – digital signal (no analog sum)
high fill factor
active quenching & recharging circuitry (AQRC)
adjustable recharge start/hold-time (pile-up rejection)
masking of noisy pixels
hit counting
wired-OR trigger
ASIC designed by DESY
    1st application:
    particle tracking
hold time
hit counter
2 bit
anode
quench
recharge
serializer
Vqt
Vrt
wired-OR
3.3 V
1.2 V
pixel electronics
AQRC
masking

## Slide 3: Readout concept: macro pixel
    16-by-16 pixel unit, 50-µm pitch
    interleaved row-wise 16-to-1 multiplexing
    12-bit TDC (5-bit fine, 7-bit coarse)
    validation logic with comparators
  adjustable thresholds
  row reference: 1 … >4 pixel
  column reference: 1 … >2 rows
    21-bit bunch counter
    serialized time stamp / validation flag / frame number
controller
  clock division & distribution
  status control register
≤ 3.3-MHz bunch clock scheme
≤ 450 MHz fast clock for MUX & TDC
sustained data throughput @ 3 MHz
  fast LVDS links
  816-Mbps hit data
  102-Mbps timing data
macro pixel
sensor (4x4 mm²) with handle wafer
ASIC (2x2 mm²) on top
130-nm GF 8M1P CMOS
30-µm SAC solder spheres by using the laser-assisted jetting technique
dSiPM-array
macro pixel
16x16
x16
fast
wired-OR
slow
wired-OR
16:1 MUX
12-bit
TDC
validation
logic
16 x pixel electronics
16 x pixel electronics
16 x pixel electronics
16
5
16
16
+
-
+
-
hit data
timing data
clocks
initials
controller
1
flag bit
bunch counter
21 bit
prototype

## Slide 4: Readout concept: detector system
HGF-AMC
interconnection board
LVDS repeater
clock distributor
decoupling
connector
flex cable
flexible carrier (termination/filtering)
ASIC (read-out)
sensor
cable
30-µm solder balls
MTCA.4
digital signal processing
(generic)
signal conditioning/distribution & power supply
(application specific)
DAQ-system
RTM
integration concept
final ASIC:
16 x macro pixel
    modular structure
    building of large-scale pixel detectors
    DAQ-system based on MicroTCA.4 standard
    application specific designed RTM
    ~14 Gbps data throughput per ASIC @ 3MHz
1HGF-AMC processes 128 macro pixel

## Slide 5: DAQ system: status
    board design and assembly ready
    FPGA programming in progress
    DAQ system available soon
    testboard with plug-in modules
testboard
plug-in module
with prototype
adapterboard
plug-in module
with prototype
testboard
MTCA.4

## Slide 6: DAQ system: beamline setup
    measurement of charged particles
    2 prototypes on top of each other
    coincidence measurements
    rejection of dark events
testboard
adapterboard
2 plug-in module
e-
MTCA.4

## Slide 7: Test setup: prototype
DG
DTG
scope
power Supply
testboard
plug-in module
with prototype
test setup at the moment
    lab instruments:
      pattern generator, oscilloscope, data-timing generator, power supply
    1 testboard with plug-in modules
    2 output streams:
      hit data
      timing data (valid bit, TDC out, frame number)
    ASIC-behavior test independently of the sensor possible by using a test input
    prototype performance tests only with dark events
ASIC behaves as expected
verification of timing characteristics, validation logic, in-pixel hit counting and data processing

## Slide 8: Results prototype: hit data
x 105 cps
dark-count rate as function of the sensor-bias voltage
dark-count rate per pixel
dark-count rate per area
    record time: ~12.5 ms
    1-hit counting per frame (3 MHz-frame rate)
red sample

## Slide 9: Results prototype: timing data
TDC value
valid bit
frame number
    time window: 333 µs = 1000 frames @ 3 MHz
    Vbias = 36.5 V
dark-count rate reduction
time stamp
timing data in every frame

## Slide 10: Results prototype: validation-logic performance
    record time: ~12.5 ms (~38000 frames @ 3 MHz)
    1-hit counting per frame
    Vbias = 36.5 V
    DCR can be significantly reduced (compare: DCR w/o validation = 180 kHz)
    crosstalk problematic shown
dark-count rate with consideration of valid bit
row: 2, col: 1
row: 3, col: 1
row: 4, col: 1
validation thresholds:

## Slide 11: Results prototype: TDC performance
    record time: ~12.5 ms (~38000 frames @ 3 MHz)
    Vbias = 36.5 V
    ‘0’ suppression by consideration the valid bit
    mean peak of TDC values at the beginning
      probability higher by enabling all pixel
      for 1 pixel enabled it is nearly equal distributed
histogram of the TDC values
time stamp of 1st event

## Slide 12: Results prototype: fine TDC performance
    in good agreement with the fine-TDC DNL
      determined by using the test input
      max. DNL error: +0.61/-0.57 LSB
Distribution of the fine TDC output values for random dark events
Mean value

## Slide 13: Test setup: ASIC - separate test circuits
DG
DTG
scope
power Supply
TC testboard
    lab instruments: pattern generator, oscilloscope, data-timing generator, power supply
    1 testboard with 1 ASIC sample
    test of separate circuits:
      1 TDC
      1 LVDS transceiver
additional performance tests

## Slide 14: Results test circuits: time-to-digital converter
valid
bunch clock
internal half clock
coarse TDC
fine TDC
trigger
timing data
reference clock
read
frame reset
requirements
3-MHz bunch clock
→ 408-MHz reference clock
timing resolution <90 ps
working principle
delay lock loop (fine TDC)
  32 delay elements
  ~77-ps timing resolution
  32-to-5 bit encoder
ripple counter (coarse TDC)
→ time stamp: 1st hit
clock
trigger
32-element diff. delay line Vcontrol
DLL
ripple counter
5 bit
7 bit
latch
latch
latch
latch
encoder
phase / frequency detector &
current pump
D
C
Q
_
Q
D
C
Q
_
Q
D
C
Q
_
Q
reg
reg
reg
reg

## Slide 15: Results test circuits: time-to-digital converter
linearity measurement
determination of INL, DNL and bin width
small bug, caused by registers setup/hold-time
→ can be improved by small design change
20-ps time conflict
coarse-TDC DNL
clock
trigger
32-element diff. delay line Vcontrol
DLL
ripple counter
5 bit
7 bit
latch
latch
latch
latch
encoder
phase / frequency detector &
current pump
D
C
Q
_
Q
D
C
Q
_
Q
D
C
Q
_
Q
reg
reg
reg
reg

## Slide 16: Results test circuits: time-to-digital converter
408-MHz reference clock
    TDC is designed for target frequency 408 MHz
    mean bin width is in good agreement with ideal value
    upper limit ≤450 MHz

## Slide 17: Results test circuits: LVDS transceiver
valid
working principle
voltage conversion
phase splitting
current switching with CMFB
termination RT = 100 Ω on TX and RX
Vcm = 1.2 V
Iout
RT-T
transceiver
differential probe
    eye-diagram measurements
    determination of bit-error rate (BER)
100 Ω
CMFB
current switching
V-conversion
phase splitting

## Slide 18: Results test circuits: LVDS transceiver
TX diff. output
816 Mbps
eye diagram measurements

## Slide 19: Summary
readout concept and detector system presented
successfully prototype assembly with 30-µm solder spheres at 50-µm pitch
functionality of sensor-ASIC prototype tested
  dark-count rate evaluated versus bias voltage
  hit detection and time stamping of first hit in macro pixel demonstrated at 3-MHz frame rate
  basic performance of validation logic using random dark events
  good agreement of fine-TDC DNL determined by dark image histogram and test input
12-bit TDC fulfills the targeted linearity requirements up to 3.3-MHz frame rate
bit-error rate remains below 10-9 up to 1.6 Gbps: link reduction by a factor of 2 is possible
next steps:
  test with DAQ system
  test with particles
outlook:
  improve TDC
  sensor with improved dark-count rate
  design transfer into new technology necessary

## Slide 20
Thank you!