## Slide 1: Frequency Hunting Application for Beam Noise Detection
DEELS’18
D. Yépez, A. Olmos
ALBA Synchrotron

## Slide 2
Content
The Problem
Frequencies tracking app
Tracked events
What’s next?

## Slide 3
Weird orbit distortions showing up and vanishing
We ca not correlate these perturbations directly with its source, because they appear at random times
VERTICAL Spectrum Amplitude
Freq (Hz)
0
50
100
150
300
350
400
17-Oct
18-Oct
450
297.5Hz
17-Oct
18-Oct
The Problem

## Slide 4
They do not have frequencies easily identifiable with a subsystem
We need a reliable way to keep track of this unusual frequencies
The Problem
~35min
Freq (Hz)
10sec beam spectrum slices
200
240
280
160
120
100
140
180
220
260
24-November

## Slide 5
Frequencies tracking app
Based on the analysis of BPM data from the Fast Archiving System
120 BPMs position data (10kHz) stored during ~13days
12Tb

## Slide 6
Matlab GUI to keep track of up to 3 frequencies using the data from the Fast Archiver
Not very reliable, uses a lot of computing power and needs a dedicated Matlab license (quite coveted at ALBA)
Frequencies tracking app
1st attempt

## Slide 7
Made in python  free
Memory handling under control
Stable  more that 3 months working without any problem
Keep track of 3 frequencies at the same time, together with many machine parameters (beam current, FOFB status, Injector status, …)
Frequencies tracking app
2nd attempt: “Frequencies Hunter”

## Slide 8
Frequencies tracking app
Machine Status:
SR current
Working mode
Date - Time
BPMs to check
Number of tries before error
Sleep time between retries
Time span to analyze
Time between captures
Number of captures
Central frequency and side bands to monitor
Data to save:
Full FFT
Partial FFT
Full BPMs raw data
Tracked frequencies Stats
Save data conditions:
Frequency peak value
Absolute frequency movement
Data view options
2nd attempt: “Frequencies Hunter”

## Slide 9
Frequencies tracking app
Log File
FFT values of the tracked frequencies saved in a log file
Log is updated on every acquisition loop, even if the save data conditions are not met
Off-line analysis of the file allow us to identify freq peaks vs. time and machine status

## Slide 10
Tracked events – Case1
Swinging 150-250Hz
23-Jan until 05-Feb
~14 days
Beam Current
Max Freq peak
Peak Amplitude
Huge noise
“The Calm”
Broad freq
swing (~100Hz)

## Slide 11
Swinging 150-250Hz
CORV0202 replaced on the 5th of February  Swinging peak gone since then
Retrieved data from the “noisy day”
CORV0202
Tracked events – Case1
From 180 to 600nm

## Slide 12
297Hz vertical noise
~20 days
Beam Current
Max Freq peak
Peak Amplitude
Huge noise
“The Calm”
???
19-Oct until 11-Nov
Tracked events – Case2
Tiny freq drift (3Hz)

## Slide 13
Tracked events – Case2
01-Mar until 21-Mar
19 days
~ 297Hz vertical noise showing up again from time to time
Beam Current
Max Freq peak
Peak Amplitude
Tiny freq drift (3Hz)
297Hz vertical noise

## Slide 14
0
Freq (Hz)
Current (mA)
150
200
100
50
300
298
296
302
Back to all OK
after reinjection
???
08:43:30 ???
14:07:04 ???
Beam noise
starts to show up
Tracked events – Case2
297Hz vertical noise
~ 297Hz vertical noise showing up again from time to time

## Slide 15
0
Freq (Hz)
Current (mA)
150
200
100
50
300
298
296
302
All OK
18:06:42
Noise again during
decay @ 120mA
No noise during reinjection …
but again there when >90mA
Tracked events – Case2
297Hz vertical noise
~ 297Hz vertical noise showing up again from time to time

## Slide 16
Freq (Hz)
Current (mA)
180
185
175
300
298
296
302
Noise
present
All OK
Noise
back
SCW quench +
beam down
SCW quench +
1mA loss
Tracked events – Case2
297Hz vertical noise

## Slide 17
CORV0303
But why related with beam reinjections, decaying beams, SCW quenches, RF cavities trips … and also when none of these happens?
CORV0303 replaced on the 5th of April  Noise gone since then
VER COR#
296
298
Freq(Hz)
301
303
Tracked events – Case2
297Hz vertical noise

## Slide 18
What’s next?
Multiple (infinite) frequencies tracking
Frequencies self-detection and tracking
Save not only mean FFT data but all BPMs FFT info
Create a TANGO Device Server to integrate it into Control System
Save data into ALBA archiving instead of logfile for easier events correlation
…
Thanks for your attention!!!