## Slide 1: FREIA Laboratory Facility for Research Instrumentation and Accelerator Development
News from FREIA
Status of the “Gersemi” Test Stand
Roger Ruber for the FREIA Team
2nd International Magnet Test Stand Workshop
BNL, 8-9 May 2018

## Slide 2: FREIA Laboratory at Uppsala University
cryogenics
- liquid helium
- liquid nitrogen
control room
- equipment controls
- data acquisition
radio-frequency (RF) power sources
3 bunkers
with test stands
horizontal cryostat
vertical cryostat
State-of-the-art Equipment
Facility for Research Instrumentation and Accelerator Development
Funded by KAWS, Government, Uppsala Univ.
Competent and motivated staff
collaboration of physics (IFA) and engineering (Teknikum).

## Slide 3: Cryogenics
Helium liquefaction
  150 l/h at 4.5K (LN2 pre-cooling)
  2000 l LHe dewar/buffer, 3+1 outlets
  cryostats connected in closed loop
Gas recovery
  100 m3 gasbag
  3x 25 m3/h compressor
  10 m3 200 bar storage
2K Pumping
  ~3.2 g/s at 10 mbar
  ~4.3 g/s at 15 mbar
  90 W at 1.8 K
  110 W at 2.0 K
Liquid nitrogen
  20 m3 LN2 tank
Sub-atmospheric pumping station
Helium gas recovery system

## Slide 4: “HNOSS” Horizontal Cryostat
HNOSS = Horizontal Nugget for Operation of Superconducting Systems
Test of superconducting cavities/devices
  3240 x ø1200mm inner volume
  up to two cavities simultaneously,
  each equipped with helium tank,
Low or High power RF testing
  fundamental power coupler (top, bottom, side)
  self excited loop or frequency generator driven
Operation in the range 1.8 to 4.5K.
Available as ARIES
Trans National Access Facility

## Slide 5: “Gersemi” Vertical Cryostat
Under construction
  May  installation
  June  start commissioning with simulator
  September/October  commissioning with inserts
Test of SC cavities & magnets
  3.2m x ø1.1m total volume
  2.65m x ø1.1m below lambda plate
    design includes joint for lambda plate
Three operation modes
  vacuum
  liquid (saturated) bath
  pressurized bath with 2K heat exchanger
Available as ARIES
Trans National Access Facility

## Slide 6: “Gersemi” Arrival & Positioning
05-Mar-2018 (081030)
Gersemi arrival at FREIA
15-Mar-2018 (P1040500)
13-Mar-2018 (P1040470)
16-Mar-2018 (P1040527)

## Slide 7: “Gersemi” Cryostat Re-assembly
21-Mar-2018 (P1040559)
22-Mar-2018 (P1040588)
27-Mar-2018 (P1040592)
28-Mar-2018 (P1040616)
28-Mar-2018 (P1040611)

## Slide 8: “µGersemi” Simulator & Valve Box
“µGersemi” simulator
2-May-2018 (100442)
2-May-2018 (100425)
“µGersemi” simulator
60 l cryostat for system test
helium distribution
cold box

## Slide 9: “Gersemi” Inserts
Liquid (Saturated) Bath Insert
(sub-)atmospheric liquid bath
mainly for SRF cavities
Pressurized Bath Insert
pressurized liquid bath
  2K heat exchanger
mainly for (small) SC magnets
  magnet + insert max 6 ton
  nominal operation 4 barG max.
    max. stored energy ~500 kJ
  ports for current leads
    4x 2kA current leads (ordered)
  central oval port for field measurement
    field measurement system itself not yet foreseen
No vacuum insert (yet) → use HNOSS
Liquid Bath Insert
2-May-2018
Magnet Insert
Delivery by end May

## Slide 10: Controls, Powering, Safety
Controls & Monitoring System
cryostat/valve box
  Siemens PLC
  NI cRIO (safety interlocks)
  EPICS interface with data archiver
  connecting different sub-systems: cryogenics, cryostats, ...
factory tests
  with valve box & simulator box performed during end 2017-begin 2018
magnet
  CERN type
Magnet Powering System
2x 2kA
  CERN type power converters
switch and power extraction unit
  CERN type
Opened valve box & electrical cabinet
26-Oct-2017
Talk by Kevin Pepitone
Tue. 5-May 17:05

## Slide 11: Summary
“Gersemi” at FREIA
commissioning after summer,
builds on a long tradition,
is a young player in magnet testing,
but working hard to contribute to the community.

## Slide 13: “Gersemi” Operation - “µGersemi”
LHe System
Vertical Cryostat
Valve Box

## Slide 14: “Gersemi” Operation – Saturated Bath
LHe System
Vertical Cryostat
Valve Box

## Slide 15: “Gersemi” Operation – Pressurized Bath
LHe System
Vertical Cryostat
Valve Box

## Slide 16: “Gersemi” Operation – Vacuum
LHe System
Vertical Cryostat
Valve Box