## Slide 1: European Scanning System
Andrey Alexandrov

## Slide 2: Toward the automatic emulsion scanning
Before 1974 – the only way to find the
charged particle tracks and decays in
the nuclear emulsions was the eye
inspection using manual microscopes
1974 K. Niwa: Track recognition by
superimposing tomographic images from different focal planes
This was the first idea of the automatic
scanning but the digital technology was not ready yet in that time (the first Digital Camera prototype from Kodak -1975)
1980 – First semi-automatic scanning (Nagoya)
1985 – “Track Selector” (TS) the first automatic scanning system based on tomographic image processing. Started TS-NTS-UTS-SUTS development line (Nagoya)
1994 – CHORUS data analysis – Italian groups enters into scanning business: two microscopes equipped with NTS systems arrived to Naples
2004 – the first prototype of the European Scanning System dedicated for OPERA scanning operational in Naples, developed in collaboration with other Italian groups

## Slide 3: OPERA ESS 20 cm2/h (2004 year components)
CMOS camera 1280×1024 pixel 256 gray levels 376 frames/sec (Mikrotron MC1310)
XY stage (Micos) 0.1 μm nominal precision
Emulsion Plate
Z stage (Micos) 0.05 μm nominal precision
Illumination system, objective (Oil 50× NA 0.85) and optical tube (Nikon)
Hardware performance of a scanning system for high speed analysis of nuclear emulsions NIMA568 (2006)

## Slide 4
16 Tomographic Images taken through 44-micron Emulsion Layer
Movable stage
Principle of the automatic emulsion scanning
Professional Image Processing board
Motor control card

## Slide 5
What the microscope CCD sees in one film..

## Slide 6: Tracks&vertices reconstruction in ECC
Field of view
Volume scan data (basetracks)
Vertex located in the brick
Track segment: aligned clusters
44 mm
15 tomographic views
Passing-through and short tracks rejected
390 μm × 310 μm
Track segment: aligned clusters
44 µm
Images -> microtracks
Microtracks->basetracks
Plate-to-plate alignment
Long tracks reconstruction
Vertex location
Event analysis

## Slide 7: OPERA tau event
Notes: Processing and reconstruction procedures has not changed since then. What has changed was the SW, the HW and scanning techniques.

## Slide 8: 2014 HW upgrade: ESS -> NGSS
Mikrotron MC-4082
DFCL interface
2336 x 1728 pixels
4 Mpixel @ 563 fps
Mikrotron MC-1310
SFCL interface
1280 x 1024 pixels
1.2 Mpixel @ 376 fps
SiliconSoftware mE5
Matrox Odyssey Xpro SFCL
PCI-X interface up to 680 MB/sec
PCIe 8x interface up to 2200 MB/sec
Nikon Plan Fluor
20X Oil
0.75 NA
Nikon Plan
50X Oil
0.9 NA
0.31 µm/pix
396x317 µm
0.34 µm/pix
794x586 µm
3.7x larger FoV
“A new generation scanning system for the high-speed analysis of nuclear emulsions” in JINST 11 P06002 2016, doi:10.1088/1748-0221/11/06/P06002
Scanning speed:
20 cm2/h
84 cm2/h

## Slide 9: Continuous Motion scanning technique
Proposed in 2011.
First published in NIM A 718 (2013) 184–185 . “A novel approach for fast scanning of nuclear emulsions with continuous motion of the microscope stage ”

## Slide 10: Distortion corrections for CM
XY-corrections: up to 1.5 µm near the view corners
Z-corrections: ~3 µm near the view corners
SG
CM
CM

## Slide 11: Microtracking performance in CM
“The Continuous Motion Technique for a New Generation of Scanning Systems” in Scientific Reports 7: 7310 2017, DOI:10.1038/s41598-017-07869-3
Scanning speed:
84
190 cm2/h

## Slide 13: LASSO Large Angle Scanning System for OPERA
LASSO
Main PC
Framegrabber Board
GPU
HDD
Motion Control Board
Z-stage
XY-stage
(WinXP 32bit)
Legend:
- GUI Module
- Guide Module
- Stage Module
- Camera Module
- Image Proc. Mod.
- Microtracker Mod.
- Commands flow
- Data flow
- Feedback
- External HW Connections
LASSO
Stage Ctrl PC
Videocamera
(Win7 64bit)
- Dispatcher Mod.
Proc. PC
LASSO
GPU
Proc. PC
LASSO
GPU
Proc. PC
LASSO
GPU
Proc. PC
LASSO
GPU

## Slide 14: How fast ESS can become?
Chart: 
ESS @ SG | 20
ESS @ CM | 40
x20+4M @ SG | 80
x20+4M @ CM | 200
x20+4M @ CMPZ | 300
2x cam @ CMPZ | 500
4x cam @ CMPZ | 700
4x cam @ new? | 1300
8x cam @ new? | 2600
ESS @ CM
New ideas to implement
New R&D
ESS @ SG (OPERA)
NGSS @ SG
NGSS @ CM
NGSS @ CM+PiezoZ
2011-2018 R&D

## Slide 15: Inclined Focal Plane Motion (IM)

## Slide 16: Possible IM implementations
Pros: Easy to implement
Cons: An objective must have the WD > 0.5 mm -> dry objectives only
Pros: Suitable for any objective type!
Cons:
The camera must be inclined at a relatively large angle
Pixel size may vary across the image

## Slide 17: Test setup

## Slide 18
SG
IM
IM
SG

## Slide 19: SG-IM microtracks matching
Dataset | Microtracks found
SG1+SG2+IM | 187
SG1+SG2 | 185
IM | 184
Microtracks selection:
Theta > 1.25 rad
Length > 90 µm
IM-SG
SG-SG

## Slide 20: Scanning Speed vs Number of Cameras
4M camera @ 563 fps
50 µm thick emulsion
30 µm overlap (X&Y)
Number of Cameras
Scanning Speed (cm2/h)
Inclined Motion
Con. Mot. + Piezo Z
Continuous Motion
HW limits:
Continuous illumination: 2 µs camera exposure -> 3500 cm2/h
Stroboscopic illumination: 30 cm/s stage speed -> 6000 cm2/h
XY-stage upgrade: 2 m/s stage speed -> 40000 cm2/h = 4 m2/h

## Slide 21: Holographic Microscopy
Collaboration with Institute of Applied Sciences and Intelligent Systems (ISASI)
Advantages
  Laser beam -> very strong illumination
  Includes phase information
  No need to move along Z or incline camera
Disadvantages
  Speckles generated by coherent laser light

## Slide 22: Summary
Successful 14 years history
20 cm2/h -> 190 cm2/h in last 7 years
Ongoing R&Ds
  High speed
  Super-resolution
Applications:
  OPERA
  FOOT & biomedical app (Cristina’s talk)
  SHiP (Antonia’s talk)
  Muography (Valeri’s talk)
  Directional DM search (Naka’s talk)

## Slide 23: Thank You!