## Slide 1: HL-LHC HEL BPM
Preliminary design considerations

## Slide 2: Outline
Specifications
First ideas
Beam signals
LF system
HF system
Acquisition system
Summary

## Slide 3: Specifications:
 | Proton beam | Electron beam | Comments
Intensity | ? - 2.5E11 (63A) | ? - 5A | 
Bunch / Pulse length | 1ns | 86us | P=4σ; E_beam off 3us; tr = 200ns
Relativistic β | 1 | 0.2 | EBeam Low β = 5* charge density
Resolution | 5um | 5um | 
Relative accuracy | 30um | 30um | 
Time resolution | 1s | 1s | TBD
Functional specifications not so clear (to me):
  E_beam on continuously and only off in abort gap?
  Injection and / or only stable beams?
  Pilots, batches, full machine…

## Slide 4: First ideas
Same BPMs for the two beams
  Minimize mechanical offsets
Same electronics for the two beams to minimize electrical offsets
  Frequency separation
    Different electronics for the two beams. Bigger offsets
    A lot of 11kHz lines will make it very difficult
  Time separation
    Measure EBeam in abort gap and PBeam outside, using same part of spectrum
    Same electronics for two beams, if same amplitude range
Low frequency (< 40MHz) or High frequency (40MHz++)
  Depends if EBeam modulation is possible?

## Slide 5
BPMs for Hollow Electron Lens

## Slide 6: BPMs considered
FNAL TEL 2 BPM
Present HEL BPM design
Button BPM
L= 100mm, ϕ=75⁰; C=20pF, R=50Ω
D= 20mm; C=20pF; R=50Ω
Zcoup_E = 15.9 ohms
Zcoup_P = 3.1 ohms
Zcoup_E = 1.1Ω
Zcoup_P = 0.2Ω
=

## Slide 7: Wide band Beam signals with electrostatic BPM
Frequency (Hz)
Response (V)
0
0
~160MHz
(20pF, 50Ω)
Zcpl_E = 15.9 ohms
Zcpl_P = 3.1 ohms
Protons :
I= 63A, 25ns
Electrons :
I= 5A,
Tr=200ns,
off 2us
EBeam
Off / on
EBeam spectrum heavily filtered
0.8Vpp
200Vp

## Slide 8: With 40MHz LP filter
6Vpp
Low frequency: PBeam signals with electrostatic BPM
Still a factor ~10 too high

## Slide 9: With 25MHz LP filter
With button BPM signal levels would be 15 times smaller ~25mV.
Electrostatic BPM seems better suited for the LF version
BP filter should be considered to minimize noise and offsets
Amplitudes shown for max currents
BPM coupling impedance for EBeam = 15.9Ω.
20% modulation @20MHz, 1A = ±8V/10=0.8V
(Factor ~10 less signal due to BPM high pass)
0.8Vpp
0.8Vpp

## Slide 10: High Frequency: With EBeam modulation @ 200MHz
Proton beam with 200MHz, 5MHz BP, 5th order, 2.5E11, σ= 0.25ns
Working in BPM pass band
Electrostatic BPM coupling impedance for EBeam = 15.9Ω.
20% sinewave modulation 1A = ±8V
Button BPM amplitudes ~ 500mVp
NEED 20% modulation
Higher resolution on EBeam
  Number of 5ns periods in 2us : 400
  20*higher amplitudes
β=0.2
100MHz no sensitivity change @ β=0.2

## Slide 11: Possible acquisition system
D
1
Abort gap
DOROS
Digitize abort gap timing to organize data in memory
Continuous digitization

## Slide 12: Summary:
We can tailor the signal conditioning (filter) to have only one acquisition system for the 2 beams, ADC included.
Low frequency system:
  Electrostatic BPM needed to get sufficient signal levels
  Resolution still to be estimated but ~100* worse than high frequency system
  Resolution can be significantly improved with high load resistance, but electrode charging could be a serious issue
High Frequency system
  In BPM passband > 200MHz
  20% Modulation is needed
  Can be very narrow band around harmonic / modulation frequency
  Low beta effect on BPM sensitivity to be verified. 100MHz seems OK
Acquisition system
  DOROS type seem a very good option