## Slide 1: Polarized target status 05/06/2018
Norihiro DOSHITA

## Slide 2: outline
Target material loading
TE calibration
Magnet commissioning
-- New field rotation procedure
Polarization test
Physics data taking
-- comparing to 2015
polarization and relaxation time
Issues to be followed
-- cooling water
-- EIO tubes
Man power

## Slide 3: Target material loading (17/4)
17/4 : loading material
Preparation
- loading platform
- nitrogen bath (forgot cleaning)
- leak detector
- LN2 dewar
- warming DR and magnet
Material
- upstream material in 2015  downstream in 2018
- downstream material in 2015  upstream in 2018
- added about 50 cc (20 -25 g ) of SMC material
- downstream socks in 2015 was cleaned in 2017

## Slide 4: TE calibration (20/4 -24/4)
Test of the communication between NMR and Magnet control system with Sylvain on Friday morning
( - dipole test was done from Thursday night to Friday morning)
Setting Trim coils  same as in 2015
Installed of He3 vapour pressure system
Filling liquid helium of 10 m^3 gas in the mixing chamber (from Thursday afternoon )
Started at 18:00 on Friday (1.44K)
Started 2nd temperature point on Saturday 16:30 (1.27K)
Started 3rd point on Sunday 16:10 (0.98 K)
Started 4th point on Monday 11:00 (2nd 0.98 K)
Target shift for TE calibration
Thanks to Vincent, counted in the normal shifts
Auto system stopped twice, otherwise working well

## Slide 5: TE calibration in Apr. 2018
21/04/2018 16:30
Start pump#4
22/04/2018 16:10
Start pump#1
1.45 K
0.98 K
1.27 K

## Slide 6: 4 temperature points

## Slide 7
July in 2015

## Slide 8: Temperature at TE calibration
He3 vapor system did not work well.
TTH4 and TTH5 available
Correction of resistance @ 2.5T
- data taken Tuesday morning

## Slide 9: Calibration constants (NMR signal area @ TE 1K)
coil1 = -54.4342 +/- 0.1816 (0.3335%)
coil2 = -22.7626 +/- 0.1136 (0.4988%)
coil3 = -48.4761 +/- 0.1896 (0.3911%)
coil4 = -27.1523 +/- 0.2678 (0.9863%)
coil5 = -43.8718 +/- 0.1447 (0.3298%)
coil6 = -13.3982 +/- 0.03178 (0.2372%)
coil7 = -18.3778 +/- 0.05888 (0.3204%)
coil8 = -33.131 +/- 0.105 (0.317%)
coil9 = -14.6842 +/- 0.07668 (0.5222%)
coil10 = -42.6571 +/- 0.2014 (0.4721%)

## Slide 10
 |  | Magnet Commissioning in 2018
20/3 | MSS Trim | Quench back heater for trim coils, control currents
20/3 | Solenoid | Power ON (10 A)
21/3 | MSS Sol | Slow discharge of solenoid
21/3 | MSS Dip | Slow discharge of dipole
22/3 | MSS Sol | Quench back heater for solenoid, fast discharge at 70 A (6 times)
22/3 | MSS | Resistance of current leads without helium flow
28/3 | Sol | Full current of 646 A
28/3 | Dip | Full current of 590A
28/3 | Sol | Stability test
29/3 | Dip | Stability test
29/3 | MSS | Safety for filed rotation (Logics of “Tiroir de Securité” )
 | Sol Dip | Present field rotation procedure (longitudinal transverse)
 | Sol Dip | New field rotation procedure (longitudinal transverse)
20/4 | Sol | TE mode procedure
 | Sol | Field homogeneity with material

## Slide 11: Issues on the magnet commissioning
Successfully done : Fabrice supported a lot.
Modification
- magnet coil voltage monitoring cards for current detection (23/3)
- new MSS program
- new power convertor control system (4/4)
Magnetic filed warning sign
Magnet field direction
- sol : + current  point to downstream
(opposite from 2015)
- dip : + current  point to up
( another intervention foreseen during LMD)

## Slide 12: New procedure test result
0.50 % loss in 2015 from sol. to dip. (from dip. to sol.) with trim coils OFF
< 0.10 % loss with trim coils ON oppositely in 2018 (new procedure)

## Slide 13
Dipole current
Solenoid current
Trim coils setting
+
+650 A
0 A
-
0
Trim coils operation in 2015 (Transverse mode)
Time

## Slide 14
Dipole current
Solenoid current
Trim coils setting
+
+650 A
0 A
-
0
Trim coils operation in 2018 (Transverse mode)
Time

## Slide 15: Polarization test
28/4 : polarization first in this year
8/5 -82.7 % (up), +81.3 % (dwn) for 40h
After several tests before data taking started.
Upgrade target parameter monitoring at DCS
done by Jaakko and Christophe
 see Christophe’s talk
-- still improving now

## Slide 16: First polarization build up in 2018
-78 % in 24h
+74 % in 24h

## Slide 17: Polarizations

## Slide 18: Polarization after 26h in 2015 and 2018
 | Upstream Cell | Downstream Cell
Positive Polarization | + 75.3 % + 75.0 % | +73.2 % + 78.2 %
Negative Polarization | - 73.5 % - 80.3 % | -72.2 % -79.4 %

## Slide 19: Relaxation time with beam in 2015 and 2018
 | Upstream Cell | Downstream Cell
Positive Polarization | ~ 1200 h ~ 1500 h | ~ 1000 h ~ 1500 h
Negative Polarization | ~ 900 h ~ 1300 h | ~ 700 h ~ 1100 h

## Slide 20: Relaxation time without beam in 2015 and 2018
 | Upstream Cell | Downstream Cell
Positive Polarization | 2070 h (3/9) 2890 h (4/6) | 1400 h (19/8)
Negative Polarization | 1250 h (19/8) | 1100 h (3/9) 2180 h (4/6)

## Slide 21: Cooling water
Tap water (microwave EIO tube, back up for pumps)
 temperature behavior known
Chilled water (Pump room air conditioner)
 pump room temperature during summer
 see Caroline’s talk
Raw water in pump room (Pumps )
Raw water inside hall (diffusion pump, NMR rack)

## Slide 22: EIO tube problem
23/5 Wed. night : Kaori realized frequency jumped more than 100 MHz.
- power also lost control
24/5 Thur. : Investigated by Yuri
- too much reflection : isolator problem
- blocked cavity control :
Cannot generate the freq. for positive pol.
Spare isolator : arrived from Yamagata (4/6)
Cathode voltage change??
Spare EIO tube ??

## Slide 23
Circulator
Circulator

## Slide 24: Cathode voltage vs Frequency
5.7 kV

## Slide 25
EIO -
oscill.
Calibration of
Trombone microm.
attenuators
Microwave Preparation

## Slide 26
Cavity
Wave guide
Motor

## Slide 28
Potential meter
readout

## Slide 29: EIO tubes in stocks

## Slide 30
Upstream waveguide
Measured by Yuri
69.953 GHz
69.984 GHz
70.209 GHz
70.273 GHz
Power
0.7 W
Power
0.7 W
Power
1.3 W
Power
1.3 W
Trombone attenuator position

## Slide 31: Man power availability
 | April | May | June | July | August | Sep | Oct | Nov | Dec
Jaakko |  |  |  |  |  |  |  |  | 
Michael |  |  |  |  |  |  |  |  | 
Genki |  |  |  |  |  |  |  |  | 
Kaori |  |  |  |  |  |  |  |  | 
Yuri |  |  |  |  |  |  |  |  | 
Nori |  |  |  |  |  |  |  |  | 
Stefan Runkel |  |  |  |  |  |  |  |  | 
28/6 – 5/7
16/5
??
26/5
??

## Slide 32: Back up

## Slide 33: After loading
Cleaned up He3 line by evacuation pumps
Restarted cooling dilution refrigerator (Tuesday evening)
- He4 pump system started on Thursday.
- He3 pump system : problem of relay
Leak check of o-ring on the flange
Leak check of the access tube (indium joint)
Leak check of isolation vacuum of the holder
Removed loading platform on Thursday

## Slide 34: NMR system preparation
Installed cooling water line
Installed NMR cables after loading
Connected NMR cables after loading
- one cable was cut but fixed it
Tuning Q-curves
Installed radiation shield

## Slide 35: Integrated area @ 1.27K (very preliminary)
coil1 = -43.1713 +/- 0.02362 (0.05471%)
coil2 = -18.4606 +/- 0.09139 (0.495%)
coil3 = -38.439 +/- 0.01839 (0.04783%)
coil4 = -21.8121 +/- 0.06017 (0.2759%)
coil5 = -34.7328 +/- 0.02014 (0.05798%)
coil6 = -10.5719 +/- 0.0169 (0.1598%)
coil7 = -14.5848 +/- 0.02518 (0.1726%)
coil8 = -26.272 +/- 0.01477 (0.05624%)
coil9 = -11.6647 +/- 0.02001 (0.1716%)
coil10 = -33.9287 +/- 0.02973 (0.08763%)

## Slide 36: Relaxation time of proton (NH3) at COMPASS
Year | Material | Program | Relaxation time
1998? SMC | OLD | Transverse | ~500 hours at 0.5T
2007 | OLD | Longitudinal /Transverse | ~4000 hours at 0.6T ~ 9000 hours at 1.0 T
2010 | OLD | Transverse | ~ 9000 hours at 0.6 T
2011 | NEW | Longitudinal | ~ 9000 hours at 2.5 T
2015 | NEW | Transverse | ~ 1200 hours at 0.6 T With hadron beam

## Slide 38
Sol Dip
-
Sol
Dip
+236 A
+465 A
+80 A
-80 A
+135 A
-465 A
-135 A
-236 A
Allowed region by MSS
Done
To be checked
MSS for field rotation