## Slide 1: Dimensional measurements Nb3Sn coils after impregnation
Salvador Ferradas Troitino (CIEMAT & CERN / TE-MSC-MDT)
Michela Semeraro (CERN / TE-MSC-LMF)
Alejandro Carlon Zurita (CERN / TE-MSC-SCD)
Jose Ferradas Troitino (UNIGE & CERN / TE-MSC-MDT)
Paris-11 October 2018

## Slide 2: Dimensional measurements review Outline
Equipment
Methodology
Results
Annex: Backup slides

## Slide 3: Equipment: CMM and Software
CMM Portable device: Faro Arm Edge 2.7
  Faro Arm Edge 2.7 volumetric accuracy: ± 41μm
  Faro Arm Edge 2.7 repeatability: 29 μm
Polyworks Inspection Software 2016 IR21.1 & 2017 IR6.1 & 2018 IR3.1
The CMM is sent to maintenance and is calibrated once a year
Technical datasheet: http://www.b3-d.com/files/96731539.pdf

## Slide 4: Equipment: Marble & Tooling
Coils clamped using a specially designed support
Dedicated area for metrology inside the workshop
  Marble @ short coils (2.5 m x 1.5 m) Class 0
  Marble @ long coils (10 m x 1.2 m)
Arm fixed to the table.
Parts of the support are aligned and screwed to the marble before operations

## Slide 5: General Procedure
Coil clamping
Data acquisition
Analysis

## Slide 6: Procedure
Distance between supports : 1200 mm
Distance between lead end and support : 180 mm
Coil clamping (I/IV): MQXFS
Rubber pads

## Slide 7: Procedure
Distance between supports : ~1400 mm
Distance between lead end and support : 250 mm
Coil clamping (II/IV): MQXFB

## Slide 8: Procedure
Distance between supports : 1500 mm
Distance between lead end and support : 120 mm
Coil clamping (III/IV): DS 11T
Rubber pads
Notes: The 11T dipole has to branches that we consider separately.

## Slide 9: Procedure
Distance between supports : ~1000 mm
Coil clamping (IV/IV): D 11T

## Slide 10: Procedure Data acquisition (I/II)
Device compensation (already warmed-up)
Device validation
Pre-alignment
General geometry probing
Best fit alignment to precisely determine the position of the cross sections
Usual results:
Hole compensation 2-sigma: Below 10μm
Calibrated pieces:
Deviation below 8μm
1Compensation
The process by which a measurement device is optimized to perform accurate measurements. This may be done through mechanical adjustments, as well as software corrections. Although the FaroArm is factory compensated, you have the ability to change probes and individually compensate each probe after mounting to the FaroArm. This process determines the centre of the probe tip relative to the arm coordinate system.
The value reported by 2 sigma is twice the standard deviation of all the points taken while performing the compensation. 2 Sigma is used to determine whether the arm passes or fails.
1Calibration
The process by which one proves that a device is performing within factory specification. After calibration, a certificate is issued to authenticate the process, thus the use of the term certification.
Cross section probing
1https://knowledge.faro.com/Hardware/FaroArm_and_ScanArm/Gage/Compensation_and_Calibration_Standards_for_the_FaroArm_and_Gage
Notes: Specialists at the FARO production facility perform the factory compensation. Using patented proprietary hardware and software, the true dimensions and kinematics of the FaroArm are determined. The results of this factory process are stored in the arm. The FaroArm is permanently compensated, such that it does not require periodic mechanical adjustments or software corrections during normal operation in the field, unless mechanical or electrical damage occurs.
Although the FaroArm is factory compensated, you have the ability to change probes and individually compensate each probe after mounting to the FaroArm. This process determines the center of the probe tip relative to the arm coordinate system.

## Slide 11: Procedure Data acquisition (II/II)
MQXFS
11 Cross sections. The straight section is probed each 6 inches.
MQXFB
33 Cross sections. The straight section is probed each 200mm
11T 2m
15 Cross sections. Probed each 100mm. The two branches are independent.
11T 5.5m
33 cross sections. Probed each 150mm. The two branches are independent.

## Slide 12
Normal vectors extraction
Each CS is aligned individually
Aligned point cloud is exported as .TXT
.TXT point cloud is imported as separate data
Results
Point cloud clean-up
Procedure Post-processing (Polyworks) (I/IV)

## Slide 13
Procedure Post-processing (II/IV)
Coil azimuthal size and asymmetry
Coil length
Coil width
Outer diameter of each cross section
Inner diameter of each cross section
Results

## Slide 14
Deviation vectors left
Deviation vectors right
Procedure Post-processing (III/IV)
MQXF Coil azimuthal size and asymmetry
Only results from cross sections along the keyway are used because these values depend on the alignment
Sample results:

## Slide 15
Deviation vectors left
Deviation vectors right
Procedure Post-processing (IV/IV)
11T Coil azimuthal size and asymmetry
Only results from cross sections along the loading plate are used
Chart: Coil size (impregnated trace)
| 50 | 150 | 250 | 350 | 450 | 550 | 650 | 750 | 850 | 950 | 1050 | 1150 | 1250 | 1350 | 1450
120 | 0.07 | 0.188 | 0.174 | 0.16 | 0.112 | 0.122 | 0.157 | 0.079 | 0.08 | 0.087 | 0.098 | 0.135 | 0.208 | 0.135 | 0.099
117 | 0.051 | 0.221 | -0.005 | 0.109 | 0.197 | 0.261 | 0.365 | 0.338 | 0.251 | 0.277 | 0.298 | 0.271 | 0.321 | 0.221 | 0.267
116 | 0.022 | 0.282 | 0.318 | 0.23 | 0.172 | 0.322 | 0.4 | 0.402 | 0.298 | 0.368 | 0.318 | 0.362 | 0.323 | 0.247 | 0.244
115 | 0.158 | 0.226 | 0.308 | 0.35 | 0.425 | 0.542 | 0.532 | 0.363 | 0.185 | 0.144 | 0.299 | 0.078 | 0.198 | 0.108 | 0.081
114 | 0.221 | 0.279 | 0.346 | 0.347 | 0.329 | 0.386 | 0.474 | 0.408 | 0.388 | 0.287 | 0.331 | 0.303 | 0.305 | 0.323 | 0.24
118 | 0.402 | 0.677 | 0.766 | 0.84 | 1.005 | 1.032 | 1.073 | 1.049 | 0.926 | 0.842 | 0.806 | 0.748 | 0.674 | 0.646 | 0.322
119 | 0.426 | 0.551 | 0.585 | 0.638 | 0.617 | 0.655 | 0.736 | 0.708 | 0.652 | 0.578 | 0.515 | 0.328 | 0.394 | 0.305 | 0.28
122 | 0.255 | 0.302 | 0.279 | 0.248 | 0.246 | 0.193 | 0.232 | 0.229 | 0.226 | 0.196 | 0.187 | 0.178 | 0.149 | 0.129 | 0.074
Sample results:

## Slide 16
Results:
Production plots

## Slide 21
MEASUREMENTS REVIEW
Annex:
Backup slides

## Slide 22
Procedure: Measurements Device compensation
Performed following FARO hardware calibration procedures:
  Hole compensation method (600 points)

## Slide 23
Procedure: Measurements Device validation

## Slide 24
Procedure: Measurements Reference system (MQXF Quadrupole)

## Slide 25
Procedure: Measurements Reference system (11T dipole)

## Slide 26
Procedure: Measurements Pre-alignment
Starting point: Pre-alignment of real coil to CAD model
  Using 6 “Surface point alignment”
  6 points are defined in the CAD model, operator must probe the same 6 points in real coil
  For improving reproducibility, singular points are chosen (Corners, edges…)

## Slide 27
Procedure: Measurements Best fit alignment (MQXF) (1/3)
After the general geometry is probed, we perform a better alignment to precisely locate the cross sections
Best-fit alignment of real point cloud to CAD model is performed using:
  Lead end plane
  Outer cylinder
  Left key plane
  Right key plane
All degrees of freedom are fixed
Coil is aligned to Lead end in order to define the cross section distances

## Slide 28
Procedure: Measurements
Best fit alignment (MQXF) Cross section alignment (2/3)
Example of cross section alignment

## Slide 29
Procedure: Measurements Best fit alignment (MQXF) Cross section alignment (3/3)
Example of cross section alignment
Original CS
Banana shape present
CS aligned using special CAD
Resulting CS
Aligned points compared with full CAD

## Slide 30
Outer Diameter
Pole end lead end (where the reference system is)
Outer Diameter
Loading plate Right
Loading plate Left
Procedure: Measurements Best fit alignment (11T Dipole) (1/3)
After the general geometry is probed, we perform a better alignment to precisely locate the cross sections
Best-fit alignment of real point cloud to CAD model is performed using:
  Pole end Lead end
  Outer cylinder
  Left loading plate plane
  Right loading plate plane

## Slide 31
Procedure: Measurements Best fit alignment (11T Dipole) (2/3)
After first part of the probing process, information is enough to perform a better alignment
Best-fit alignment of real point cloud to CAD model is performed using:
  Pole end Lead end
  Outer cylinder
  Left loading plate plane
  Right loading plate plane
Coil is aligned to Lead end in order to define the cross section distances

## Slide 32
Procedure: Measurements Best fit alignment (D11T Cross section alignment (1/3)
Example of cross section alignment
Original CS
Banana shape present
CS aligned using special CAD
Final “CS
Aligned” points and deviation vectors to nominal

## Slide 33
Procedure: Measurements Extra: Post-processing (V/IV)
Coil length
The point cloud used is the one corresponding
to the general geometry.
Ref.

## Slide 34
Procedure: Measurements Extra: Post-processing (VI/IV)
Coil width
The individually aligned cross sections are duplicated and then trimmed. Only the points belonging to the ID remain.
Only results from cross sections along the keyway//loading plate are used because the result depends on the alignment
Name | Dev
Δ T3 | -0.177
Δ T4 | -0.100
Δ T5 | -0.120
Δ T6 | -0.058
Δ T7 | -0.097
Δ T8 | -0.101
Δ T9 | -0.135
Sample results:

## Slide 35
Procedure: Measurements Extra: Post-processing (VII/IV)
Outer diameter and Inner diameter
Circumferences are best-fitted to the points belonging to the inner and outer diameter (wo outliers). The algorithm makes minimum the square of the deviation of the points belonging to the arc. Then the center and the radius are computed.
The radius does not depend on the alignment. The center does.
Name | Nominal | Measured | Dev
ID 1 | 74.750 | 74.664 | -0.086
ID 2 | 74.750 | 74.737 | -0.013
ID 3 | 74.750 | 74.811 | 0.061
ID 4 | 74.750 | 74.884 | 0.134
ID 5 | 74.750 | 74.755 | 0.005
ID 6 | 74.750 | 74.863 | 0.113
ID 7 | 74.750 | 74.759 | 0.009
ID 8 | 74.750 | 74.780 | 0.030
ID 9 | 74.750 | 74.725 | -0.025
ID 10 | 74.750 | 74.683 | -0.067
ID 11 | 74.750 | 74.525 | -0.225
RADIUS
Sample results:

## Slide 36: Thanks for your attention,