## Slide 1: SNS Ring BPM Old and New
Richard Dickson
Nov 12, 2018

## Slide 2
Topics:
Facility overview
Old system and requirements
New system topology and functional diagrams
Data compression
Jitter reduction
New capabilities

## Slide 3
1 GeV LINAC
Accumulator Ring: compress 1-msec long pulse to ~700 nsec
2.5 MeV
Warm LINAC
Front-End
RTBT
HEBT
Injection
Extraction
RF
Collimators
945 ns
1 ms macropulse
Current
mini-pulse
Chopper system makes gaps
Current
1ms
Liquid Hg Target
1000 MeV
1 ms macropulse
1 ms
<1 msec
Cold LINAC
186 MeV
38mA
38A
Spallation Neutron Source (SNS) Accelerator Complex
Front-End:
Produce a 1-msec long, chopped, H- beam
Courtesy of A. Aleksandrov
The SNS is located in Oak Ridge, Tennessee (U.S.)

## Slide 4
Ring Beam Structure
Nice square pulses: Not!
Beam varies within and between each turn in:
Intensity
Position
Energy
Beam in 4/5 of beam turn.
(~800nS of ~1000nS)

## Slide 5
SNS Ring Beam Position Monitors (BPMs)
44 floated strip-line BPMs of aperture: (28) 21cm, (8) 26cm, (8) 30 cm
Position Resolution and Accuracy of +/- 1% half aperture
Original requirement of ~5E10 to 2E14 protons per pulse (~8mA to 35A).
  Now ~6E9 to 6E14 ppp (~1mA to 100A)
  Dynamic range is of major importance
Custom made PCI AFE and digital cards (no in-house expertise)
LV 2001 SW with Embedded Windows XP on individual PCs (one per BPM)
Meets accuracy specs but reliability is poor

## Slide 6
New System Motivation
Hardware obsolescence is major problem
  Parts, cards, PC motherboards
Software obsolescence is major problem
  Impossible to upgrade OS, security concerns
Lost ability to repair boards with personnel attrition
Original system limited to 1Hz. Would like to include 60Hz acquisition.
Long term solution: New System

## Slide 7
New Analog Front End

## Slide 8
Analog Front End System Diagram
(Each AFE hosts 2 BPMs)
2 BPMs
(2 X 4 Wires)
Relay Switch
Each wire
Cal Oscillator
(1 MHz)
X8
Low Gain
Amp
High Gain
Amp
5 MHz
Filter
5 MHz
Filter
X8
AFE
To Digitizer
To Digitizer
CPU Control & Status

## Slide 9
PXIe system with four FPGA/Digitizer(center blue), processor running LabVIEW RealTime (left) and SNS timing card (right).
1/4 space, 1/5 power consumption of previous system.

## Slide 10
System Diagram
SNS Timing Interface
Real Time CPU
FPGA (X4)
PXI Bus
Timing Decode
Logic
Event Link/
Real Time Data Link
Gate
Gain Mixing
PXI Bus
Analog Front End
(2 X 8 Ch High/Low Gain)
2 BPM Pickups
(8 Wires)
Beam Envelope Calculation
Data Formatting
EPICS
Network
PXI System
Digitizer
(50 MHz)
RTDL
Triggers
Envelope
Data
NI PXIe-8880 CPU
NI PXIe-7971 FPGA (x4)
NI 5751B ADC (x4)

## Slide 11
Data Rates (per BPM, after compression):
1 Hz rate 50 MHz full accumulation: 1.31 mS (512 kB/s)
60 Hz rate single turn position (22.5 kB/s)
60 Hz rate orbit data (30 kB/s)
ADC resolution of 14 bits delivered in 16 bits.
Each BPM wire has high and low gain analog channels (2 ADC channels)
Data is compressed by FPGA for transfer to RT CPU as follows:
7 Bits (X128)
1
0
Low Gain
High Gain (X128)
S
S
S
Re-Combined
21 Bits Range (not resolution)
Each FPGA supports two BPMs (16 ADC channels)
Each system supports four FPGAs (64 ADC channels)
Data Compression

## Slide 12
Jitter Reduction

## Slide 13
Ring turns as seen by BPM
Stacked
Turns

## Slide 14
1.047115 MHz Ring
50 MHz ADC
47 Samples
48 Samples
47.75 Samples
With (for example):
50/1.047115 = 47.75 samples/turn
Jitter Type 1: Sampling Mismatch Jitter

## Slide 16
Turn 63/64
ADC n-2
ADC n-1
%
%
% = %
Turn 64/64
Turn 65/64
Turn 0, ADC 0
N is ring turn number
n is closest integer number of samples
ADC n
ADC data is interpolated to force
it to 64 samples per ring turn.

## Slide 17
Before:
After:

## Slide 20
Trigger Window
50 MHz ADC Clock
(20 nS)
200 MHz Trigger Clock
(5 nS)
0
1
2
3
1
3
2
4
Sample N-1
Sample N
50 MHz ADC
200 MHz Timing
SNS Beam Trigger
Start
Jitter Type 2: Trigger Jitter
Jitter
0
3

## Slide 21
50 MHz ADC Clock
200 MHz Timing Clock
0
1
2
3
0
2
1
3
Sample N-1
Sample N
Actual Trigger
0
3

## Slide 23
New Features
Orbit synthesis
Centroid/Peak phase display
Single turn motion
Combined screen

## Slide 24
Multi Pulse Beam
Single Pulse Beam (many Stored Turns)
(Used during tuning to compute “Orbit” of particles in ring)
Orbit Synthesis

## Slide 25
 | 1 | 2 | 3 | 4 | 5 | M |  |  |  | 
1 | 0 |  |  |  |  | 0 |  |  |  | 
2 | 10 | 0 |  |  |  | 5 |  |  |  | 
3 | -4 | 10 | 0 |  |  | 2 |  |  |  | 
4 | -10 | -4 | 10 | 0 |  | -1 |  |  |  | 
5 | -6 | -10 | -4 | 10 | 0 | -2 |  |  |  | 
On1 = Om1 |  | M1 = O11 |  |  |  |  |  |  | O11 = M1 | 
On2 = Om2 |  | M2 = (O12+O21)/2 |  |  | 2*M2 = O12+O21 |  | O12 = 2*M2 - O21 |  | O12 = 2*M2 - M1 | 
… |  | M3 = (O13+O22+O31)/3 |  |  | 3*M3 = O13+O22+O31 |  | O13 = 3*M3-(O22+O31) |  | O13 = 3*M3 - 2*M2 | 
Ont = Omt |  | M4 = (O14+O23+O32+O41)/4 |  |  | 4*M4 = O14+O23+O32+O41 |  | O14 = 4*M4-(O23+O32+O41) |  | O14 = 4*M4 - 3*M3 | 
* Ont is the orbit (position) of the n'th injected pulse on its t'th turn around the ring. |  |  |  |  |  |  |  |  | O1t = t*M(t) - (t-1)*M(t-1) | 
Mt is the measured agragate position of all mini-pulses on turn t. |  |  |  |  |  |  |  |  |  | 
Chart: Five Pulse Accumulation
Pulse 1 | 0 | 10 | -4 | -10 | -6
Pulse 2 | 0 | 10 | -4 | -10
Pulse 3 | 0 | 10 | -4
Pulse 4 | 0 | 10
Pulse 5 | 0
Measured | 0 | 5 | 2 | -1 | -2
Turn
Pulse
Measured
Single Pulse Orbit De-Convolution
Pos

## Slide 26
Chart: BPM C08 Hor
Measured | -20.2306 | 23.5818 | 16.6733 | -18.0398 | -42.226 | -3.7949 | 32.7218 | 16.9993 | -25.2382 | -17.2072 | 1.6583 | 24.4207 | -2.4894 | -28.0869 | -10.2584 | 19.3325 | 26.808 | -7.4516 | -24.8871 | -2.7242 | 14.8018 | 13.3023 | -9.1712 | -20.2093 | -8.5258 | 21.7849 | 10.1794 | -10.2823 | -19.9493 | 4.9921 | 2.6772 | 2.8258 | -19.9512 | -7.8526 | 5.1255 | 11.5953 | -6.5582 | -15.887
Calculated | 20.5378 | 71.5265 | 36.7801 | -20.5506 | -1.8203 | 32.6739 | 34.5056 | -1.256 | -25.1601 | -8.7274 | 22.2526 | 21.6365 | -3.0349 | -15.1517 | 5.4915 | 28.1404 | 26.9511 | 0.14689 | -14.9498 | 4.6365 | 30.9803 | 18.6268 | -3.3025 | -0.048687 | 12.1253 | 17.1714 | -0.87287 | -2.0372 | 7.3956 | 16.1856 | 4.1265 | 2.6824 | 6.1516 | 9.0557 | 6.4729 | 3.3562 | 6.9443 | 7.9739
Chart: BPM C08 Ver
Measured | -17.2596 | -8.6152 | -8.4643 | -10.6659 | -12.5452 | -28.727 | -16.4235 | -7.4764 | -2.4682 | -9.6755 | -20.5274 | -23.3835 | -12.8082 | -6.1686 | -16.3898 | -20.9368 | -13.9129 | -24.2636 | -28.3534 | -15.8124 | -10.1606 | -8.8894 | -4.2634 | -27.9277 | -21.0786 | -10.7789 | -3.285 | -8.9417 | -4.1222 | -15.9362 | -23.7912 | -16.6345 | -9.202 | -10.5372 | -15.4998 | -16.8981 | -11.3883 | -28.5284
Calculated | 9.2957 | 13.2415 | 27.2295 | 6.0267 | -2.6158 | -0.65626 | 6.2563 | 6.4495 | 2.7536 | -4.2191 | -5.7568 | -4.5849 | 1.0419 | 4.7851 | 6.0033 | 2.5412 | -0.4782 | -3.91 | -4.2319 | -6.1845 | -6.3255 | -0.13551 | -1.6377 | -5.108 | -9.0024 | 0.092155 | 3.5006 | 5.4775 | 6.9351 | -2.4443 | -8.2644 | 1.1125 | 5.6451 | 12.5024 | 10.7813 | 4.7523 | -2.289 | 1.988

## Slide 27
Centroid Phase

## Slide 28
Single Turn Position

## Slide 29
Position and Swing Bars

## Slide 30
Combined Screen

## Slide 31: SNS Ring BPM Old and New
Questions?
System development thank you to:
Alexander Aleksandrov
Taylor Davidson
Alysse Dawson
Anthony Webster

## Slide 32
The following slides are not part of this presentation.
They are included only in anticipation of being needed to answer questions.