## Slide 1: BPM Signal Processing using Time-Multiplexed Electrode Signals
M. Wendt
CERN

## Slide 2: Outline
Introduction to the Beam Position Measurement
  Symmetry in beam position monitoring
  Calibration methods to ensure a high measurement stability
  Single channel heterodyne receiver
BPM electronics utilizing time-multiplexed electrode signals
  DESY: BPM electronics at HERA-e and FLASH
    Design principle introduced by R. Neumann
  CERN: LHC interlock BPM R&D
    Based on the thesis activities of Oskar Bjorkquist, and with help of Irene Degl’Innocenti and Jan Posipil

## Slide 3: Introduction: A typical BPM setup
symmetry plane
beam
position
Aelec signal conditioning
ADC
DAQ
Belec signal conditioning
ADC
digital
analog
A
B
large common mode
signal contents
large common mode
signal contents
small difference mode
signal contents
small difference mode
signal contents
Aelec signal
Belec signal

## Slide 4: Beam Position Measurement
BPMs are based on a symmetric measurement setup
  Detect the beam asymmetry, i.e. the beam position,
    by a perfect symmetric arrangement of 2 identical read-out channels
    to suppress the common mode
    to simplify the normalization to the beam intensity
    Beam position signals: An AM signal with the bunch response as carrier
Asymmetries in the BPM read-out system channels
  Will result in an (electronic) offset of the reported beam position
    can be tolerated and calibrated if the asymmetric effect does not change over time
    Is often linked to tolerances of RF / analog electronic components, RF cables and connectors, etc.
      Can also be design choices, e.g. different BPM electrodes or asymmetric arrangement, different gains to electronically center a permanent large beam offset
Time varying asymmetries result in uncontrollable position offsets, and are a major limitation of the BPM performance!
  Caused by a variety of effects in the analog and RF signal processing, e.g. ambient temperature, humidity, aging and radiation effects of components
  Also external EM-fields (pulsed RF, kicker signals), or uncontrolled grounding can break the symmetry

## Slide 5: Example: Aging of RF components
RF connectors and coaxial cables also undergo aging effects!

## Slide 6: Online Calibration Methods
Calibration tone technique
  only applicable in narrowband operation
  Requires a separate detector channel
courtesy N. Eddy
Crossbar switching technique
  e.g. used for the Libera BPMs from Instrumentation Technology
  <100 nm stability over 14 hours
  No position measurement during switching transition
courtesy P. Leban
Online calibration methods
  require a few stable, high performance “reference” components, like
    analog switches
    RF couplers & connectors
    transmission-lines
  Adds complexity to the BPM system

## Slide 7: Single channel BPM: MPX Receiver
Narrowband RF heterodyne receiver with multiplexed inputs:
  Downmix, demodulate and normalize the BPM signals
    Classical RF radio-receiver technology applied to process BPM signals
    Supplies the individual BPM electrode signals and the hor./ vert. position signals
  Single channel signal processing with T&H at the analog outputs
    Improves the stability due to drift of electronics components
MPX
Pre-
Amp
Active
Matrix
A
B
C
D
X
Y
AGC
S
BP
Filter
BP
Filter
VCO
PLL
Limiter
BP
Filter
IF
Amp
Mixer
VLSI
Freq.
Synth.
Mixer
MPX
A
B
C
D
courtesy J. Hinckson, K. Unser, J. Bergoz

## Slide 8: Time-multiplexed BPM signals
bunched beam current
signal conditioning
ADC
DAQ
Electrode A
Electrode B
Belec signal,
direct
Aelec signal,
delayed
signal delay,
e.g. coaxial cable
requires every 2nd or more RF buckets empty!
combined A & B signal
+
B
B
A
A

## Slide 9: DESYs early MPX BPM System
Developed and operated at the DESY HERA electron ring
  Every 48th 2 ns bucket filled (96 ns bunch-to-bunch distance)
“Exported” to the Fermilab A0-Photoinjector test facility linac

## Slide 10: DESY MPX BPM Electronics @ FNAL

## Slide 11: DESY FLASH MPX BPM System
courtesy B. Lorbeer

## Slide 12: DESY MPX BPM Electronics
100 ns delay-line signal combination (2 electrodes)
  80 m long 3/8” high quality coaxial cables
40 dB RF pre-amplifier to cover 15 pC single-bunch operation
  NF = 3.3 dB, 600 MHz BW
  Resolution <10 μm for single bunch operation >10 pC
Analog electronics based on μTCA RTM
  600 MHz BW (-3 dB) & double-peak detector
    VRF peak min = 5 mV
  Gain switching (RF step attenuator)
  4 input channels and test pulse generators
Commercial μTCA digitizer Struck SIS8300
  10-ch, 125MS/s, 16-bit
  External RF synchronous 108 MHz clock
Unfortunately: No long-term drift analysis data available, …yet.
  However, this BPM technology is routinely and successfully used for beam energy calibration, which is cross calibrated to the photon energy of the FLASH FEL.

## Slide 13: CERN LHC Interlock BPM R&D
25 ns
Beam
25 ns
Stripline pick-up
Delay 12.5 ns
Delay 2 ns
Filter output
Electronics rack
Tunnel
anti-aliasing LPF
~50 m coaxial cable
comb BPF

## Slide 14: LHC Interlock BPM: Key Elements
Stripline BPM
  2 vertical and 2 horizontal, 120 mm long electrodes
High isolation, balanced high-power signal combiner
4-stage delay-line based, comb (FIR) band-pass filter
ADC digitizer and digital signal processing
12.5 ns
Stripline BPM
Amplitude balanced combiner
Delay line filter
∆2 ns
∆4 ns
∆6 ns
Digitizer
&
FPGA

## Slide 15: Stripline BPM Bunch Signals
CST wakefield simulation and oscilloscope measurement
  1.35e11 protons per bunch
  Measurement captured after ~70 m ½” Heliax cable
    Oscilloscope LeCroy Waverunner SDA 18000 (60 GS/s, 18 GHz BW)
Time domain comparison: Simulation and measurement
Frequency domain comparison: Simulation and measurement
2 ns
Amplitude difference: 10 %
~500 MHz

## Slide 16: Amplitude-balanced Power Combiner
port 2
Σ-port 1
port 3
Ferrite beads
~ coaxial cable @ 500 MHz
~ coaxial cable @ 500 MHz
Ferrite beads
~6 dB insertion loss
10 W CW power handling
40 dB isolation @ 2 GHz
±0.4 dB amplitude, ±40 phase balance
  in the range 10-5000 MHz

## Slide 17: RF FIR Filter Theory
4x 2 ns delays:
  maxima at: n x 500 MHz
Ideal filter impulse response
Ideal filter transfer function

## Slide 18: High Isolation Power Combiner
Prototype design
Port 1
Port 2
Port 3

## Slide 19: Stripline PCB Comb BPF
Frequency response of PCB filter.
LHC beam measurement of PCB filter at 1.35e11 bunch intensity.
PCB layout and initial power divider design.
First manufactured prototype.
1.35e11 proton bunch via ~70 m long coaxial cable
  Acquired at 60 GS/s with 18 GHz BW
Stripline PCB prototype
  Center frequency off by ~5 %
  Substrate: Rogers RO4360G2 (εr = 6.15)

## Slide 20: Beam Measurements
60 GS/s, 8-bit
3.2 GS/s, 12-bit
2.6 GS/s, 14-bit

## Slide 21: LHC Doublet Bunches (Simulation)
Single bunch (measured)
A1+
A2-
A2+
A1-
A+
A-
25 ns
25 ns
Doublet bunch (simulated)

## Slide 22: Estimated Performance
Meets the LHC interlock BPM resolution requirement
  <500 μm RMS bunch-by-bunch for a range of 5e9…2e11 ppb w/o gain switching!
  including a beam displacement range of ±7.5 mm
Keeps the reported mean value beam position over the entire bunch intensity range
Operates also with 5+20 ns doublet bunches
  at a reduced performance

## Slide 23: Summary
Single channel, time-multiplexed BPM electronics can be an alternative to a multi-channel BPM read-out technique
  Requires empty RF buckets
Time MPX BPM technologies are based on
  Precise, stable time delays utilizing high quality coaxial cables and power combiners
  Low-pass integration or comb-style FIR band-pass filters in connection with track&hold circuits, peak detectors, or fast digitizers
Time MPX BPMs performance is successfully demonstrated in ring and linear accelerators
  DESY HERA-e, Fermilab A0-Photoinjector, DESY FLASH
  In future: CERN LHC interlock BPMs
Long term drift stability could not yet be quantified: TBD!
  However, the concept omits the needs of online calibration or channel switching schemas, thus appears to be more simple and straightforward.

## Slide 24: Backup Slides…
Thank you!

## Slide 25: Stripline electrodes: Characteristics
Coupled signal is dependent on geometry of electrode
One can say that a certain impulse response relates the beam current to the pickup voltage in the time domain through convolution:
Beam current
Electrode voltage

## Slide 26: Stripline electrodes: Characteristics
… And in the frequency domain, a certain transfer impedance relates the beam current and the pickup voltage through multiplication (ohms law):
Beam current
Electrode voltage

## Slide 27: Stripline electrodes: Characteristics
Given the beam current and the electrode voltage signal, the transfer impedance can be calculated:
FFT

## Slide 28: Wilkinson power divider
‘Typical’ choice of high isolation power divider + Cheap and simple to manufacture - Poor bandwidth
Isolation
Reflection
Transmission

## Slide 29
A Balun is a 1:1 transformer
  Signal amplitudes on primary and secondary are equal
Primary and secondary signals are 1800 out of phase
  Bal-Un: Balanced-Unbalanced transformer
High isolation power combiner: How does the isolation work?

## Slide 30
High isolation power combiner: How does the isolation work?
Example: Assume the input signal at port 2
  All resistors are 50 Ohm
The 1800 phase shift of the balun forces the node at port 3 to receive zero voltage
  port 2-3 isolation

## Slide 31: Amplitude balanced power combiner: Coaxial balun design
Notes: I want to talk a bit about the balun design:
The balun transformer that I am going to show you has the exact same function as the balun that I was talking about previously. It should produce equal amplitudes on the balanced outputs and also should produce a perfect 180 degree phase shift. This balun however works a very differently and is instead made using a set of coaxial cables. It is actually possible to create a balun using a single piece of coaxial cable as is suggested in the top picture. However, due to the skin effect, electrons are concentrated to inner side of the coaxial jacket which leads to the outer side acting as a separate signal path, effectively acting as a short circuit stub that is connected in parallel with the 180 degree output. This means that the two outputs will see different loads and thus there will be an imbalance in amplitudes. This problem is solved by connecting a second coaxial sleeve to the zero degree output, which makes each of the two outputs see the same load.

## Slide 32
Amplitude balanced power combiner: Balun design
Amplitude: Balanced output 2
Amplitude: Balanced output 1
Amplitude: Balanced output 1
Amplitude: Balanced output 2
Second coax balances amplitudes, but balanced outputs are still frequency dependent
Notes: I made measurements that show this effect: In the graph to the left I have measured a balun using only a single coaxial cable, and on the right I have added the second, balancing cable. The graphs show the amplitudes on the two differential outputs and should in the ideal case be completely equal. As expected we can see very big imbalances in the left graph as the two outputs are effectively loaded with different impedances. On the right, the two outputs have approximately the same loads and the amplitudes are thus much more even. However, there are still these notches for certain frequencies at which no signal at all is transmitted through the balun. This is still an unwanted effect, and we can counteract this by introducing a ferrite material onto the coaxial sleeve. This will make the currents flowing on the outer jacket see a very high inductance that instead makes this signal path look more like an open circuit.

## Slide 33
Amplitude balanced power combiner: Balun design
Amplitude: Balanced output 2
Amplitude: Balanced output 1
Ferrite beads nearly remove frequency dependency

## Slide 34
Amplitude balanced power combiner: Balun design
Max phase deviation: 4 degrees
Max amplitude difference: < 0.8 dB

## Slide 35
Amplitude balanced power combiner: Balun design
Balun amplitude: Balanced output 2
Balun amplitude: Balanced output 1
Coaxial cable based balun
Cables covered in ferrites to prevent currents from flowing on the outside of coaxial jacket
Produces a very well performing amplitude balance over a large frequency range

## Slide 36: Delay line filter: Theory
An ideal filter (no losses) can be modelled very easily using basic signal theory

## Slide 37: Filter theory
What is the frequency domain behavior?
Impulse response:
Standard transform pairs:
Transfer function:

## Slide 38
Surface mount power dividers/combiners: Why high isolation?
Why can we not split and combine the signals directly on the transmission line? - No isolation causes filter to ring!

## Slide 39: Amplitude balanced power combiner
The surface mount combiner is not well balanced for some frequencies
This performance is effectively limited by performance of the balun transformer
The other circuit is in principle the same as the surface mount version
  uses a different type of a balun, based on distributed transmission-line elements
Surface mount combiner transmission: Input 2
Surface mount combiner transmission: Input 2

## Slide 40: Power combiners: Dissipation
Software developed specifically to estimate the power dissipation in a S-parameter network that is fed with a real stripline pickup signal