## Slide 1: Challenges in medical accelerator design
Jarno Van de Walle
Accelerator physicist
Jarno.vandewalle@iba-group.com

## Slide 2: Outline
Cyclotrons in proton therapy
Major future challenges
  Energy degrader and beam losses
  Variable energy accelerators
Beam diagnostics in the compact IBA ProteusONE

## Slide 3: Cyclotrons in proton therapy

## Slide 4: Current cyclotrons for proton therapy
IBA C230
230 MeV protons
4.3 m Diameter
CW beam
Normal conducting
Magnet: 200 kW
RF: 60 kW
Varian-Accel Probeam
250 MeV protons
3.1 m Diameter
CW beam
Superconducting (NbTi)
Magnet: 40 kW
RF: 115 kW
Mevion SC250
250 MeV protons
~1.5 m Diameter (shield)
Superconducting (Nb3Sn)
IBA S2C2
MeV protons
2.2 m Diameter
Rep. rate: 1 kHz
Superconducting (NbTi)
RF: 11 kW

## Slide 5: Ongoing cyclotron developments : fixed energy
Varian/Antaya
230 MeV protons
2.2 m Diameter
CW beam
Superconducting (Nb3Sn)
30 tons+
5.5 T (extr.)
“Flutter” coils
Pronova/Ionetix
250 MeV protons
2.8 m Diameter
CW beam
Superconducting (Nb3Sn)
60 tons
3.7 T (extr)
Hefei/JINR
200 MeV protons
2.2 m Diameter
CW beam
Superconducting
30 tons
3.6 T (extr.)
SHI
230 MeV protons
2.8 m Diameter
CW beam
Superconducting (NbTi)
55 tons
4 T (extr.)

## Slide 6: Size and weigth versus field

## Slide 7: Limits for isochronous cyclotrons
Minervini, MIT, DTRA-TR-12-40
< 0 : dfdt < 0 = synchro cyclotron
> 0 : f=cte = isochronous cyclotron
 Continuous beam
 Pulsed beam

## Slide 8: Superconducting (SC) challenges
Fabrication of SC coils on industrial scale
Cryogenics installation : cryocoolers (“dry”) or He bath (“wet”, ex. fast ramping)
In synchro cyclotrons the SC coil position is crucial in extracting the correct energy and direction of the beam
  S2C2 (5.7 T central field) :
    Horizontal positioning precision down to 0.1 mm needed
    Vertical beam angles sensitive to sub 0.1 mm vertical tilt/shift of the coil
  MEVION (9 T central field) :
  Cyclotron rotates with gantry : active tie rod system needed

## Slide 9: Major future challenges

## Slide 10: Major future challenges
Minimize beam losses
  Reduce decommissioning costs
  Reduce shielding requirements = reduce size
  Synchro-cyclotron : largely asymmetric emittances
  No more degrader …
Variable energy accelerators
  Linacs
  Synchrotrons
  FFAG’s
  Superconducting (ironless) synchro-cyclotrons
  + achromatic gantries

## Slide 11: The degrader in the IBA ProteusONE
Cyclotron vault
Treatment room
degrader

## Slide 12: The degrader in the IBA ProteusONE
Symmetric emittance in front of rotating gantry needed
BPM + beamstop on degrader position
Air filled IC in front of degrader

## Slide 13: The degrader in the IBA ProteusONE
Transmission from cyclotron exit to isocenter
Al
Graphite
Be

## Slide 14: The degrader in the IBA ProteusONE
Transmission from cyclotron exit to isocenter
Al
Graphite
Be
Horizontal beam tracks
230 MeV
70 MeV

## Slide 15: Major future challenges
Minimize beam losses
  Reduce decommissioning costs
  Reduce shielding requirements = reduce size
  Synchro-cyclotron : largely asymmetric emittances
  No more degrader …
Variable energy accelerators
  Linacs
  Synchrotrons
  FFAG’s
  Superconducting (ironless) synchro-cyclotrons
  + achromatic gantries

## Slide 16: Variable energy options
Hitachi
70-250 MeV protons
Slow (>1s) or fast cycling (50 ms)
7 m Diameter
“PIMMS” (CERN) design
Up to Carbon
25 m Diameter
Rep. rate: 5 Hz
Installed @CNAO, MedAustron

## Slide 17: Variable energy options
Protom
Up to 330 MeV protons
5 m Diameter, ~16 tons
Being installed @MGH

## Slide 18: Variable energy cyclotron (development)
MIT/ProNova
250 MeV protons
(2.4-)2.8 m Diameter
Pulsed beam
Superconducting (Nb3Sn)
4 tons
Cost …. ?
Variable-energy possible
 2.8 m

## Slide 19: Beam diagnostics in the IBA ProteusONE

## Slide 20: S2C2 and ProteusONE : time structure
62
68
74
86
92
80
RF Frequency [MHz]
0
5
10
RF voltage [kV]
0
200
400
600
800
1000
Time [ms]
Diamond detectors (in collaboration with Cividec)
Fourier transform of diamond signal

## Slide 21: Beam monitor devices in the CGTR (*) : BPM’s
degrader
1. Beam Position Monitor (BPM)
Air filled ionization chamber with H & V wires
 60 mm
(*) Compact Gantry, part of the ProteusONE system

## Slide 22: Beam monitor devices in the CGTR (*) : BPM’s
degrader
1. Beam Position Monitor (BPM)
Air filled ionization chamber with H & V wires
 60 mm
(*) Compact Gantry, part of the ProteusONE system
@ end of “energy selection system” (ESS)
Disperion function maximized

## Slide 23: Beam monitor devices in the CGTR (*) : BPM’s
degrader
1. Beam Position Monitor (BPM)
Air filled ionization chamber with H & V wires
 60 mm
(*) Compact Gantry, part of the ProteusONE system
@ entrance of scanning magnets

## Slide 24: Beam monitor devices in the CGTR : BPM’s
degrader
1. Beam Position Monitor (BPM)
Air filled ionization chamber with H & V wires
 60 mm
Horizontal beam tracks
230 MeV
70 MeV

## Slide 25: Beam monitor devices in the CGTR : IC CYCLO
degrader
2. IC CYCLO (Ionization chamber) + BEAMSTOP
Measures beam pulses coming out the S2C2 : 0.100 to 150 pC/pulse (1e6-1e9 protons)
IC CYCLO : 2 IC’s with 1 mm and 2.5 mm gap (asymmetric ionization chamber)
1. Beam Position Monitor (BPM)

## Slide 26: Beam monitor devices in the CGTR : IC CYCLO
degrader
3. Nozzle IC’s
Measures beam pulses in the nozzle : 0.100 to 4 pC/pulse
Asymmetric IC’s (2 gap sizes : 3 and 5 mm)
2. IC CYCLO (Ionization chamber) + BEAMSTOP
1. Beam Position Monitor (BPM)

## Slide 27: Nozzle beam diagnostics
Upstream scanning
Large area air-filled IC (30x30 cm2)
HV : 1.3 kV
Position and charge (dose) measurement
Recombination is major issue

## Slide 28: The asymmetric IC
62
68
74
86
92
80
RF Frequency [MHz]
0
200
400
600
800
1000
Time [ms]
Proton pulse (10 ms)
Ionization current (200 ms)

## Slide 29: The asymmetric IC
beam
62
68
74
86
92
80
RF Frequency [MHz]
0
200
400
600
800
1000
Time [ms]
Proton pulse (10 ms)
Ionization current (200 ms) : integrated charge =
d1 = 3 mm
d2 = 5 mm

## Slide 30: The asymmetric IC
beam
62
68
74
86
92
80
RF Frequency [MHz]
0
200
400
600
800
1000
Time [ms]
Proton pulse (10 ms)
Ionization current (200 ms) : integrated charge =
d1 = 3 mm
d2 = 5 mm

## Slide 31: The asymmetric IC
beam
62
68
74
86
92
80
RF Frequency [MHz]
0
200
400
600
800
1000
Time [ms]
Proton pulse (10 ms)
Ionization current (200 ms) : integrated charge =
d1 = 3 mm
d2 = 5 mm
Charge amplification
Charge collection efficiency

## Slide 32: Boag theory for pulsed beams
DETAILS :
p = free electron fraction, depending on type of gas, gap size and voltage (ex. p=1 for N2)
parameter with
= incoming proton charge !
Boag et al., Phys. Med. Biol. 41, 885 (1996)
PARAMETERS :
- type of gas (air, N2, …)
- type of particle (protons)
- gap size
- voltage
- spot size

## Slide 33: The asymmetric IC

## Slide 34: Conclusions
Reduce footprint of proton therapy facility :
  Reduce cyclotron size  trend towards synchro-cyclotrons
  Reduce beam losses  reduce shielding
  Variable energy option : superconducting ironless synchro-cycloton
Beam diagnostics :
  Non-interceptive …
  Consider difference pulsed vs continuous beam

## Slide 35: Thank you
Jarno Van de Walle
Jarno.vandewalle@iba-group.com

## Slide 36: The asymmetric IC : example (IC CYCLO)

## Slide 37: Beam losses : degrader
Emittance increases a lot for lower energies.
A circular collimator in front of the gantry reduces losses inside the gantry considerably