## Slide 1: Design and Optimisation of Ultra-Compact, High-Resolution X-Ray Imaging Systems
Thomas Primidis
3rd OMA Topical Workshop – Medical Accelerator Design and Diagnostics

## Slide 3: Digital Tomosynthesis vs Computerised Tomography
Much higher dose than X-ray but offers great 3D detail.
Around double dose than X-ray but offers depth information.

## Slide 4: The benefit of the emitter array

## Slide 5: Contents
Electron generation and acceleration
X-ray production from deceleration of electrons in a high Z target.
Filtration of the X-ray beam.
Transport of X-rays in the phantom-patient.
Radiation collection in a pixelated detector and generation of an image.
3D image reconstruction using generated images.
Discussion on the methods, approximations and next steps.

## Slide 6: X-ray generation
Ta
a few μm
Al filter
1.5mm
60 keV e- pencil beam
0 emittance
90 degrees incident angle
X-rays
No magnetic field
No electric field
Gaussian beam spot
6σ=100μm

## Slide 7: Compare the unfiltered spectrum with literature
Taken from
www.xrfresearch.com
on 7/12/18

## Slide 8: The spectrum after the Al filter.
Approximate this with a Gaussian:
μ=30keV
3σ=20keV

## Slide 9: The emitter array
45 point X-ray sources
7x7 configuration
1cm pitch
Uniform cone beam each of
40 degrees full cone angle
Spectrum is Gaussian approximation of the previous simulation.

## Slide 10: The phantom human head
89 voxels x0.488=43.5mm
35 voxels x0.488=17.1mm
Voxel size: 0.488mmx0.488mmx1.25mm
184 slices x1.25=230 mm
Anonymous phantom in DICOM format taken from:
Patient Contributed Image Repository www.pcir.org

## Slide 11: The ideal X-ray detector
100% X-ray detection efficiency
100μm x 100μm pixel size
360 pixels, 36 mm total width
260 pixels, 26mm total height

## Slide 12: The simulation geometry
Emitter array plane
Detector 1mm away from phantom at teeth height.
Source Detector Distance = 12cm
Phantom selected volume

## Slide 13: X-Ray images
Beer-Lambert Law

## Slide 14: Image reconstruction

## Slide 15: 3D Shepp-Logan phantom
Ellipsoid parameters taken from:
HU to material conversion by

## Slide 16: Shepp-Logan phantom digital tomosynthesis

## Slide 17: Next steps: Use realistic description of the following
Assumed monoenergetic e- pencil beam with zero emittance.
Neglected electric and magnetic fields in the source geometry.
Fitted a Gaussian instead of sampling energies from the actual spectrum.
Assumed point sources.
Assumed uniform cone beam of X-rays.
Used an ideal detector.
Did not include all the components of the system like collimators etc.
This is only a proof of principle that the image quality of DT can be assessed with Monte Carlo techniques and offer system optimisation potential.

## Slide 18: Most important challenge: Phase space sampling
Source simulations only generate a few 100.000 X-rays per day on cluster.
Imaging with high resolution detector requires trillions of X-rays to reduce stat. uncert.
5 orders of magnitude difference

## Slide 19: Most important challenge: Phase space sampling
Source simulations only generate a few 100.000 X-rays per day on cluster.
Imaging with high resolution detector requires trillions of X-rays in MC.
A fast way to generate million times more particles must be found.
  Discrete multivariate distributions E(x,y,x`,y`)
  Multivariate Kernel Density Estimation
  Generative Adversarial Neural Networks
  Biasing techniques like particle splitting or even rare event.
  Apply some reasonable approximations.

## Slide 20: Thank you
Thomas.Primidis@liverpool.ac.uk
linkedin.com/in/thomasprimidis