## Slide 1
Beginning of training is associated with a pronounced spike in Vtap and QA signals at the onset of every quench. The most striking example is quench 5 where such spike also coincides with a loss of Vtap A06 in Q1. This points to the mechanical origin of the precursor spike.
Post-spike “ringing” is seen in Vtap and antenna signals, and its frequency is fairly high (~4 kHz), suggesting the motion is associated with a coil or structural part that is not overly large or massive (as that would give rise to much lower frequencies of vibration)
The antenna signals suggest that mechanical activity in quench 5 is well confined in the RE section below element 13 of the antenna
The summary table of QA-based quench locations to quench initiation in-between element 14 and 16 for the majority of training quenches, independently of the quenching coil
In the last OL quench #13 some weak antenna signals are observed in the LE for the first time; quench onset appears to be close to element 3.

## Slide 2
Large mechanical event in quench 5 results in a loss of voltage tap A06 which coincides with the onset of the quench

## Slide 3
Frequency of post-spike voltage ringing is at least 4 kHz :
not likely to be of entire coil or anything massive
- Likely due to motion of a small-sized coil fraction or structural element

## Slide 4
E1
QA1/2
E2
QA3/4
E3
E4
E13
QA26/27
E14
QA28/29
E15
QA30/31
E16
QA32/33

## Slide 5
Quench antenna raw signals of quench 5
Quench is centered at the element 14 (QA 27/28) where largest signal amplitude is seen. Elements 15 and 16 (towards RE from element 14) are active as well, while elements towards LE (13, 12…) show very little or no activity. One can guess that Q1 coil motion was well confined to the RE region down from element 13.

## Slide 6
Quench #13: QA6 (element 3, LE) - OL quench
Quench #12 : ?? (no discernible signals) - OL quench
Quench #11 : QA30,QA31 -> QA29,QA28 (element 15 -> element 14, RE)
Quench #10 : QA26 -> QA28 (elements 13 -> element 14, RE)
Quench #9 : QA28 & QA30 (?, weak signals)
Quench #8 : QA28 (?, weak signals) (element 14, RE)
Quench #7 : QA30 -> QA32; QA30 -> QA28 (element 15 -> elements 16 and 15 -> element14)
Quench #6 : ?? (no discernible signals)
Quench #5 : QA27/QA28 (element 14, RE)
Quench #4 : QA32 (element 16, RE)
Quench #3 : QA29 (element 14, RE)
Quench antenna localization summary

## Slide 7
Quench antenna signal example (quench 10)
QA26 (element 13)
QA28 (element 14)
Wavelet-filtered and rms (5-point) averaged antenna voltages are plotted to improve signal/noise

## Slide 8
Quench antenna signal example (quench 13)
QA6 (element 3)
Wavelet-filtered and rms (5-point) averaged antenna voltages are plotted to improve signal/noise