## Slide 1: MQYYM magnetic MEASUREMENT
MQYYM
INSTRUMENTATION
TECHNICAL REVIEW
Damien Simon for the MQYYM team
CEA: H. Felice, D. Simon, M. Segreti, J.M. Rifflet, S. Somsom, R. Machado-Correia, R. Godon
CERN: M. Guinchard, P. Grosclaude, J.C. Perez, N. Bourcey, A Foussat, E. Todesco

## Slide 2: INtroduction
The MQYY project:
  Collaboration CERN-CEA for the HL-LHC
  Design of the MQYY a double aperture cos(2θ) Nb-Ti matching section magnet
  Fabrication and test of the MQYYM single aperture magnet short model of the MQYY magnet
  Fabrication of the MQYYM coils at CEA
  Assembly, instrumentation and warm magnetic measurements at CERN
  Cold magnetic measurements at CEA
The MQYY
The MQYYM

## Slide 3: Strain gauges position
Collars instrumentation:
Strain gages

## Slide 4: ScenariI
Collars instrumentation:
  The goal is to measure the azimtal stress on the pole
  The radial stress should be low to measure only the azimutal stress
  3 possibilities:
    0 cut
    1 cut
    2 cuts
Strain gages

## Slide 5: Drawings
Collars instrumentation:
  The goal is to measure the azimtal stress on the pole

## Slide 6: Introduction to the simulation
Collars instrumentation:
  PATH: 1 path
    NOSE2
  Notation: e(x/y)StepYZCuts
    X Radial strain & Y azimuthal strain
    Y number of the load step (1 collaring, 2 cooling, 3 enregization)
    Z number of cuts (0 cut, 1 cut, 2 cuts)

## Slide 7: Introduction to the simulation
Collars instrumentation:
  The idea is to canceled the azimuthal stress
  The poisson ratio of the stainless steel is 0,3
  To canceled the azimuthal stress we need a ratio between the radial strain and the azimuthal strain of -0,3.

## Slide 8: RESULTS 0 Cut
Radial and azimuthal strain in the collar nose for 0 cut
(Step 3)
azimuthal strain
Radial strain

## Slide 9: RESULTS 0 Cut
Radial and azimuthal stress in the collar nose for 0 cut
(Step 3)
azimuthal stress
Radial stress

## Slide 10: RESULTS 0 Cut
The radial strain along the « nose 2 » path is huge and not constant especialy at the end of the nose (x = 18 mm)
The radial stress is not canceled
0 mm
18 mm

## Slide 11: RESULTS 1 Cut
Radial and azimuthal strain in the collar nose for 1 cut
(Step 3)
azimuthal strain
Radial strain

## Slide 12: RESULTS 1 Cut
The radial and azimuthal strain along the « nose 2 » path is almost contant.
The radial stress is not canceled at the middle of the path.

## Slide 13: RESULTS 2 Cuts
Radial and azimuthal strain in the collar nose for 2 cut
(Step 3)
azimuthal strain
Radial strain

## Slide 14: RESULTS 2 Cuts
The radial and azimuthal strain along the « nose 2 » path is almost contant.
The radial stress is almost canceled all along the path
2 CUTS: BEST OPTION

## Slide 15: RESULTS 2 Cuts
Path (m) | Step 1 |  |  | Step 2 |  |  | Step 3 |  | 
 | ey (me) | ex (me) | ex/ey | ey (me) | ex (me) | ex/ey | ey (me) | ex (me) | ex/ey
0 | -204,6 | 55,7 | -27% | -137,0 | 37,2 | -27% | -54,4 | 14,8 | -27%
9,21E-04 | -266,9 | 67,2 | -25% | -190,8 | 48,7 | -26% | -76,5 | 20,3 | -26%
1,84E-03 | -294,5 | 66,6 | -23% | -216,3 | 49,6 | -23% | -86,4 | 21,4 | -25%
2,76E-03 | -303,1 | 66,6 | -22% | -224,6 | 50,0 | -22% | -88,4 | 22,2 | -25%
3,68E-03 | -306,2 | 71,6 | -23% | -228,1 | 54,2 | -24% | -88,6 | 24,6 | -28%
4,60E-03 | -308,6 | 79,7 | -26% | -231,3 | 60,8 | -26% | -89,1 | 28,0 | -31%
5,52E-03 | -312,0 | 88,8 | -28% | -235,6 | 68,2 | -29% | -90,8 | 31,2 | -34%
6,45E-03 | -316,1 | 97,0 | -31% | -241,2 | 74,9 | -31% | -93,7 | 33,4 | -36%
7,37E-03 | -313,2 | 101,1 | -32% | -242,5 | 78,9 | -33% | -98,1 | 33,9 | -35%
8,29E-03 | -271,3 | 86,4 | -32% | -214,1 | 68,2 | -32% | -101,5 | 30,9 | -30%
9,21E-03 |  |  |  |  |  |  |  |  | 
1,01E-02 | -374,3 | 94,1 | -25% | -291,9 | 73,4 | -25% | -145,4 | 36,9 | -25%
1,10E-02 | -341,1 | 86,4 | -25% | -264,1 | 67,1 | -25% | -135,5 | 33,1 | -24%
1,20E-02 | -330,8 | 90,2 | -27% | -255,8 | 70,3 | -27% | -132,9 | 34,0 | -26%
1,29E-02 | -326,3 | 96,3 | -30% | -252,7 | 75,2 | -30% | -132,7 | 36,4 | -27%
1,38E-02 | -321,0 | 101,3 | -32% | -248,5 | 78,7 | -32% | -132,1 | 39,1 | -30%
1,47E-02 | -313,0 | 103,8 | -33% | -241,8 | 80,1 | -33% | -130,3 | 41,3 | -32%
1,57E-02 | -302,4 | 103,3 | -34% | -232,4 | 78,8 | -34% | -126,7 | 42,5 | -34%
1,66E-02 | -286,8 | 97,9 | -34% | -219,0 | 74,0 | -34% | -120,7 | 41,2 | -34%
1,75E-02 | -265,3 | 87,1 | -33% | -200,9 | 65,8 | -33% | -113,7 | 36,9 | -32%
1,84E-02 | -235,4 | 68,8 | -29% | -170,7 | 50,6 | -30% | -111,5 | 31,9 | -29%
 | -299,6 | 85,5 | -29% | -227,0 | 65,2 | -29% | -107,4 | 31,7 | -30%
Path (m) | Step 1 |  |  | Step 2 |  |  | Step 3 |  | 
 | sy (MPa) | sx (MPa) | sx/sy | sy (MPa) | sx (MPa) | sx/sy | sy (MPa) | sx (MPa) | sx/sy
0 | -41,9 | -1,3 | 3% | -28,1 | -0,9 | 3% | -11,1 | -0,3 | 3%
9,21E-04 | -55,0 | -2,9 | 5% | -39,3 | -1,9 | 5% | -15,7 | -0,6 | 4%
1,84E-03 | -61,2 | -4,9 | 8% | -44,9 | -3,4 | 8% | -17,8 | -1,0 | 6%
2,76E-03 | -63,2 | -5,4 | 9% | -46,8 | -3,9 | 8% | -18,2 | -1,0 | 5%
3,68E-03 | -63,5 | -4,5 | 7% | -47,3 | -3,2 | 7% | -18,1 | -0,4 | 2%
4,60E-03 | -63,5 | -2,9 | 5% | -47,5 | -1,9 | 4% | -18,0 | 0,3 | -2%
5,52E-03 | -63,7 | -1,1 | 2% | -48,0 | -0,6 | 1% | -18,2 | 0,9 | -5%
6,45E-03 | -64,0 | 0,5 | -1% | -48,8 | 0,6 | -1% | -18,7 | 1,2 | -6%
7,37E-03 | -63,1 | 1,6 | -3% | -48,8 | 1,4 | -3% | -19,6 | 1,0 | -5%
8,29E-03 | -54,7 | 1,1 | -2% | -43,2 | 0,9 | -2% | -20,6 | 0,1 | 0%
9,21E-03 |  |  |  |  |  |  |  |  | 
1,01E-02 | -77,2 | -4,1 | 5% | -60,2 | -3,2 | 5% | -30,0 | -1,5 | 5%
1,10E-02 | -70,3 | -3,6 | 5% | -54,4 | -2,7 | 5% | -28,0 | -1,7 | 6%
1,20E-02 | -67,8 | -2,0 | 3% | -52,4 | -1,4 | 3% | -27,4 | -1,3 | 5%
1,29E-02 | -66,3 | -0,4 | 1% | -51,3 | -0,1 | 0% | -27,2 | -0,7 | 3%
1,38E-02 | -64,8 | 1,1 | -2% | -50,2 | 0,9 | -2% | -26,8 | -0,1 | 0%
1,47E-02 | -62,9 | 2,2 | -3% | -48,6 | 1,7 | -3% | -26,3 | 0,5 | -2%
1,57E-02 | -60,5 | 2,8 | -5% | -46,6 | 2,0 | -4% | -25,4 | 1,0 | -4%
1,66E-02 | -57,4 | 2,7 | -5% | -43,9 | 1,9 | -4% | -24,2 | 1,1 | -5%
1,75E-02 | -53,4 | 1,7 | -3% | -40,4 | 1,2 | -3% | -22,9 | 0,6 | -3%
1,84E-02 | -47,9 | -0,4 | 1% | -34,7 | -0,1 | 0% | -22,7 | -0,4 | 2%
 | -61,1 | -1,0 | 1% | -46,3 | -0,6 | 1% | -21,8 | -0,1 | 0%
Strain
Stress
sx=(E/(1-n2))*(ex+ney) | 
E | 2,03E+11
n | 0,3

## Slide 16: Strain Gages
1-XC11-1.5/350
a | b | c
1,5 mm | 1,5 mm | 6 mm
Orientation of the gages:
Orientation 1
Orientation 2

## Slide 17: Strain Gages Orientation
1-XC11-1.5/350
a | b | c
1,5 mm | 1,5 mm | 6 mm
Orientation 1
Orientation 2

## Slide 18: Strain Gages orientation
Strain Gages parallel to the pole

## Slide 19: Strain Gages orientation
Strain Gages perpendicular to the pole

## Slide 20: Strain Gages orientation
Strain max: - 297 me
Strain min: - 307 me
Strain mean: - 302 me
Nodes
Time 1 (layer 1)

## Slide 21: Strain Gages orientation
Strain max: 92 me
Strain min: 48 me
Strain mean: 70 me
Nodes
Time 1 (layer 1)

## Slide 22: Strain Gages orientation
Strain max: - 300 me
Strain min: - 320 me
Strain mean: - 310 me
Nodes
Time 1 (layer 2)

## Slide 23: Strain Gages orientation
Strain max: 105 me
Strain min: 80 me
Strain mean: 92 me
Nodes
Time 1 (layer 2)

## Slide 24: Strain Gages orientation
Strain max: - 125 me
Strain min: - 133 me
Strain mean: - 129 me
Nodes
Time 3 (layer 2)

## Slide 25: Strain Gages orientation
Strain max: 45 me
Strain min: 35 me
Strain mean: 40 me
Nodes
Time 3 (layer 2)

## Slide 26: Strain Gages orientation
BEST SOLUTION: Solution 1

## Slide 27: Strain Gages orientation
30%
Time 3
(layer 2)