## Slide 1
ETH to fiber converters and
fiber bundle radiation hardness
tests at CHARM and IRMA
BE-BI-PM
Ack. S.Danzeca*, C. Cangialosi*, J. Lendaro*, H. Desjonqueres**, S. Poirier** (* CERN, ** IRMA-CEA Saclay)
RADWG Meeting 13th December 2018 D.Celeste BE-BI-PM

## Slide 2
Introduction
Beam profile imaging nowadays: BTV – Vidicon
  Advantages:
    Most Rad Hard so far…
    No active components inside.
  Drawbacks:
    Limited life.
    Limited sensitivity (observation instruments).
    Issues of procurement (no production since > 5years, still ≈100 pieces in stock).
Beam profile imaging upgrade: imaging fiber system
  Principle: Transport an image away from a high radiation
  level area up to the camera located in a ‘safer’ place.
  Advantages: Increased life time.
  Procurement of 2x2m and 2x10m long/bunch of 10k fibers.
Radiation hardness tests on:
1. Digital cameras @ CHARM
(RADWG 13th Avril 2017, S. Burger)
2. Eth/fiber converters @ CHARM
3. Fiber bundle @ IRMA
(RADWG 23rd May 2018, D. Celeste)
Camera
RADWG Meeting 13th December 2018 D.Celeste BE-BI-PM

## Slide 3
Summary
Introduction
Tests at CHARM: introduction and setup
Results on ethernet to fiber converters RadHard test
Tests at IRMA introduction and setup
Results on fiber bundle RadHard test
Conclusions
RADWG Meeting 13th December 2018 D.Celeste BE-BI-PM

## Slide 4
Increase of the use of high performance digital cameras in BI.
Radiation effect issue mitigated with optical line, shielding and RadHard innovation.
Tests at CHARM: introduction and setup (1)
Distance to its control front end.
The Gigabit Ethernet allows up to 100m connections.
Not enough in the case of the SPS BTV dump future camera based (BTV) system.
Ethernet to Optical Fiber converter
Eth/fibre
1G/Single ch.
Digital Camera
Eth/fibre
1G/Single ch.
Ethernet
(max 100m)
Dual Fibre SC single mode
Ethernet
PowerPC, Ethernet network, etc.
As SPS machine…test in CHARM facility
RADWG Meeting 13th December 2018 D.Celeste BE-BI-PM

## Slide 5
Request:
  To reach approximately 50 Gy TID
  RUN1: from 08/08/2018 to 14/08/2018
  RUN2: from 16/08/2018 to 21/08/2018
Tests at CHARM: introduction and setup (2)
Monitoring tool by CHARM
Extract dose/HEH/1MeV eq. fluence
24GeV
~4E11p
400ms spill
Every ~30s
Location of the modules
RADWG Meeting 13th December 2018 D.Celeste BE-BI-PM

## Slide 6
Radiation area CHARM G0
Safe area
Eth/fibre
1G/Single ch.
Digital Camera
Ethernet
Power supply1
Function | Desc. | Connector | Qty
Ethernet |  | RJ-45 | 10
Power Supply1 | +12V | 2 wires or coax | 2
Fibre type1 | SC Single Mode | SC type fiber connector (EKI-2741SX/LX) | 2x2
Summary of cables and fibres needed
Layout for Ethernet/Fibre module irradiation test @ CHARM
Eth/fibre
1G/Single ch.
x2
Fibre type 1
‘GUDE‘ power switch CTRL
Ethernet
CHARM installation
RADWG Meeting 13th December 2018 D.Celeste BE-BI-PM

## Slide 7
Analysis:
Logging cameras
Logging GUDE power switch
Checking connection failures
CHARM results (1)
Test procedure:
Image acquisitions every minute
Power cycles at each connection loss
Module1: 28 failures
Module2: 22 failures
RADWG Meeting 13th December 2018 D.Celeste BE-BI-PM

## Slide 8
CHARM results (2)
Parameter Value | 
POT [#] | 4.64973 E+11
Dose [Gy] | 6.71596 E-04
1 MeV Neutron eq. Fluence [cm-2] | 1.629182257 E+07
HEH eq. Fluence [cm-2] | 1.372112079 E+06
Average impact on modules per spill:
 | Module 1 | Module 2
Reset number | 28 | 22
HEHeq | 8.91E+10 | 8.91E+10
[cm2] | 3.14E-10 | 2.47E-10
Φfailure [cm-2] | 3.18E+09 | 4.05E+09
Φfailure Basler camera [cm-2] | 1.5 E+09 | 
 | Module 1 | Module 2
Reset number | 28 | 22
HEHeq | 8.91E+10 | 8.91E+10
 | 3.14E-10 | 2.47E-10
Φfailure [cm-2] | 3.18E+09 | 4.05E+09
Φfailure Basler camera [cm-2] | 1.5 E+09 | 
Results module power cycles:
SETUP STILL WORKING AT
THE END OF THE TWO RUNS
OK!
RADWG Meeting 13th December 2018 D.Celeste BE-BI-PM

## Slide 9
Isotropic 60Co source facility
Dose rate to reach at least 500 kGy total dose
Reasonable measurement time (days/month)
Reasonable cost
Tests at IRMA: introduction and setup
Simulating 10 years of operation of fiber bundle
Choice of gamma radiation
On-line results to monitor behaviour
IRMA facility at CEA Saclay
RADWG Meeting 13th December 2018 D.Celeste BE-BI-PM

## Slide 10
IRMA introduction and setup (2)
Dose Rate = 3.6kGy/h
Irradiation Time = 8 d
Relaxation time = 1 d
Acq. Image delay = 15’
Temperature = 24-27 °C
Date/Time Integrated Dose [Gy] | 
Start acquisition: 10/09/2018 14:41:03 | 0
Start irradiation: 10/09/2018 15:14:05 | 0
Stop irradiation: 18/09/2018 15:14:31 | 694.076 E+03
Stop Acquisition: 19/09/2018 14:36:25 | 694.076 E+03
RADWG Meeting 13th December 2018 D.Celeste BE-BI-PM

## Slide 11
IRMA results
Before Irradiation (0 Gy) After Irradiation (700 kGy) After relaxation (700 kGy)
1
2
3
1
2
3
NORMALIZATION
RADWG Meeting 13th December 2018 D.Celeste BE-BI-PM

## Slide 12
Conclusions
Ethernet to fiber radiation hardness test:
Single events can stop camera acquisitions
Power cycles reset correctly the system
System keeps working up to 45 Gy TID
Shielding for ETH to fiber converters foreseen
Failure cross section lower than cameras -> Non limiting factor
Fiber bundle radiation hardness test:
700 kGy TID reached in 8 days
Integrated light loss around 50%
Non-uniform light loss -> to be investigated
Temperature effects
Still able to reconstruct a qualitative image
RADWG Meeting 13th December 2018 D.Celeste BE-BI-PM

## Slide 13
Analysis:
Images selected every 10 kGy
Up to Dtot = 240 kGy (3 days)
Normalized to reference
Results:
Center stable
Edges down to 50%
IRMA results (2)
RADWG Meeting 13th April 2017 S.Burger BE-BI-PM
Analysis:
Images selected every 10 kGy
From Day 3 to Day 8
From 240 kGy to 700 kGy
Normalized to reference
Results:
Center starts decreasing faster
down to 70 %
Edges down to 30 %

## Slide 14: Relaxation time profiles

## Slide 15: Spikes vs Temperature

## Slide 16: Temperature corrected profiles