## Slide 1
The National High Magnetic Field Laboratory is supported by the National Science Foundation through NSF DMR-1644779 and the state of Florida.
A part of analysis work by U. Bong and S. Hahn was supported by the National Research Foundation of Korea as a part of Mid-Career Research Program (No. 2018R1A2B3009249)
Design and Performance Estimation of a 20 T No-Insulation all-REBCO User Magnet
1National High Magnetic Field Laboratory, 1800 East Paul Dirac Drive, Tallahassee, FL 32310, USA
2Department of Electrical and Computer Engineering, Seoul National University, Gwanak-ro 1, Gwanak-gu, Seoul 08826, South Korea
3GE Healthcare, 3001 W Radio Drive, Florence, SC, 29501, USA
1. Introduction
Tue-Af-Po2.14-02
Kwangmin Kim1, Kabindra R. Bhattarai1,3 ,Kwanglok Kim1, Hongyu Bai1, Iain R. Dixon1, Thomas A. Painter1, Uijong Bong2, David Larbalestier1 and Seungyong Hahn1,2
Fig. 1. To-scale drawing of the 20 T NI magnet.
This paper reports a design and performance estimation of a 20 T no-insulation (NI) high temperature superconductor (HTS) standalone user magnet currently being developed at the National High Magnetic Field Laboratory, hereinafter named ‘NI20T’. It consists of a stack of 18 double pancake coils wound with tapes from two different vendors, SuNAM and SuperPower, in consideration of their complementary in-field critical current and mechanical properties. The inner and outer winding diameters and overall winding height of the magnet are 46 mm, 102.6 mm, and 163.4 mm, respectively. The magnet is designed to generate a center field of 20 T at its nominal operating current of 316 A in a bath of liquid helium at 4.2 K. The inductance and stored energy are 1.89 H and 94.5 kJ, respectively. First, we report key design parameters followed by numerical simulations on: (1) Magnetic field performance, stress analysis; (2) NI charging behaviors and estimation of liquid helium consumption; (3) post-quench behaviors including a local temperature rise.
Parameter | Unit | Module-1 | Module-2 | Module-3
REBCO tape |  |  |  | 
Manufacturer |  | SuNAM | SuperPower | 
Width; thickness | [mm] | 4.1; 0.12 | 4.1; 0.096 | 6.1; 0.096
Substrate thickness | [mm] | 0.1 | 0.05 | 0.05
Cu stabilizer thickness | [μm] | 7.5 (per side) | 20 (per side) | 20 (per side)
Ic at 77 K, self-field | [A] | 150 |  | 
Equivalent Young’s modulus (Er, Eh, Ez) | [GPa] | 107; 168; 168 | 72; 130; 130 | 
Equivalent Poisson’s ratio (νrh, νhz, νrz) |  | 0.21; 0.33; 0.21 | 0.18; 0.33; 0.18 | 
Magnet configuration |  |  |  | 
Inner; outer diameter | [mm] | 46; 102.6 |  | 
Overall height | [mm] | 163.4 |  | 
Turns per Pancake |  | 236 | 295 | 
Number of DP coil |  | 10 | 6 | 8
Total conductor per module | [m] | 1,102 | 828 | 276
Estimated Rca | [mΩ] | 24.7 | 18.9 | 4.1
Operation |  |  |  | 
Center field, Bc | [T] | 20 |  | 
Operating current, Iop | [A] | 316 |  | 
Operating temperature, Top | [K] | 4.2 |  | 
Conductor current density, Je | [A/mm2] | 642.3 | 802.8 | 539.6
Inductance, L | [H] | 1.89 |  | 
Storage energy at Iop | [KJ] | 94.4 |  | 
Estimated time constant, τ | [sec] | 39.6 |  | 
aContact surface resistance: Rct=50 μΩ∙cm2
2. Magnet design
a. Key Parameters of NI20T
Performance Estimation of NI20T
b. Configuration of the NI20T
c. Magnetic Field and Critical Current Distribution
Fig. 2. Magnetic field distribution and critical current estimation of NI20T; (a) normal magnetic field distribution, (b) Field angle (θ) distribution, (c) critical current distribution.
d. Stress and Strain
(a)
(b)
Fig. 3. Mechanical characteristics on axial midplane of NI20T; (a) magnetic hoop stress and radial stress, (b) overall calculated strains including magnetic and bending strains.
Fig. 4. Lumped circuit model for the charging characteristic and the quench simulation of the 20 T NI magnet.
L: coil inductance
Ir: leak current
M: mutual inductance
Ips : power supply current
Rc: characteristic resistance
Rq: quench resistance
Fig. 5. Charging delay analysis results of NI20T (ramp rate: 0.2 A/sec).
a. Charging behaviors
b. Post-quench behaviors
Fig. 6. Azimuthal current pattern of each DP coil during quench; (a) module-1, (b) module-2 and 3.
Fig. 7. Temperature variation of each DP coil during quench; (a) module-1, (b) module-2 and 3.
Force equilibrium and generalized Hook's laws were used.
Since NI20T is expected to be “dry-wound” without using epoxy, the so-called “self-supporting-turn" characteristic of NI20T is also considered in the stress analysis.
Estimated critical current (Ic) in consideration of Bnorm, θ, and tape width at 4.2 K.
The minimum Ic of modules-1, 2 and 3 are calculated to be 565 A, 436 A and 493 A, respectively.
M: module
Min: Minimum Ic
r-axis
z-axis
Magnet center
Bnorm
r-axis
z-axis
Magnet center
(a)
(b)
M1
M2
M3
Min: 493 A
Min: 436 A
Min: 565 A
r-axis
z-axis
Magnet center
1400 A
400 A
1300 A
1100 A
900 A
700 A
500 A
(c)
90o
0o
80o
70o
60o
50o
40o
30o
20o
10o
20.3 T
18 T
14 T
10 T
6 T
2 T
0 T
A lumped circuit model was used for charging and quench behavior analysis of the NI20T.
The calculated τc was 39.6 seconds for NI20T. With a constant ramping rate of dIps/dt = 0.2 A/sec, the time to reach the target magnetic field after the power supply current reached 316 A is calculated to be ≈3.3 minutes for NI20T.
The quench propagates to all pancake coils in about 0.32 s
During the quench propagation, some DP coils experience overcurrent (Fig. 6). The peak current is calculated to be 830 A in DP18 at t = 0.28 s.
After NI20T is completely discharged, the highest and lowest final temperatures of DP coils of NI20T are calculated to be 117 K (DP coil-16) and 88 K (DP coil-5), respectively.
larger amount of energy is dissipated in the end DP coils.
CONTACT INFORMATION: Kwangmin.kim@asc.magnet.fsu.edu