## Slide 1
Beam imaging system with radiation tolerant fibre bundle
D. Celeste, S. Burger, E. Bravin, F. Roncarolo (BE-BI-PM)
Ack. S.Danzeca*, C. Cangialosi*, J. Lendaro*, H. Desjonqueres**, S. Poirier** (* CERN, ** IRMA-CEA Saclay)

## Slide 2
Introduction
Beam profile imaging nowadays: BTV – VIDICON tube cameras
  Advantages:
    Most Rad Hard so far…
    No active components inside.
  Drawbacks:
    Limited life.
    Limited sensitivity (observation instruments).
    Issues of procurement (no production since > 5years, still ≈100 pieces in stock).
Beam profile imaging upgrade: imaging fiber system
  Principle: Transport an image away from a high radiation
  level area up to the camera located in a ‘safer’ place.
  Advantages: Increased life time.
  Procurement of 2x2m and 2x10m long/bunch of 10k fibers.
RADWG Meeting, 04th Dec 2018 D. Celeste, BE-BI-PM
Camera

## Slide 3
Summary
Introduction
Optical characterization setup
Radiation hardness tests on fibre @ IRMA
Beam imaging test @ TT2
Conclusions
RADWG Meeting, 04th Dec 2018 D. Celeste, BE-BI-PM

## Slide 4
Beam profile imaging with optical fiber bundle
Lens 1
Opt fiber in
Opt fiber out
Lens 2
CCD camera
Fiber bundle choice:
10m Fujikura RadHard
FIGR10 – 10k fibers in
1.5mm diameter.
Lens choice:
9mm lens to focus a
65mm diameter
circular area
@ 500mm into the
fiber bundle.
25 or 50mm lens to
magnify the image on
the camera sensor

## Slide 5
Optical characterization results
Advantage:
Digital cameras allow many options
(exposure time, etc..) that can
significantly compensate optical losses.
Reference image:
Target @ 500mm
Texp = 20ms
Cam + Lens 25mm
Acquired image:
Target @500mm
texp=20ms
1st lens 9mm
10k fiber bundle
2nd lens 25mm
Analysis | Results
Light loss due to 2nd lens | From 100% to 26.2%
Light loss due to Fujikura fiber interfaces | From 26% to 2.9%
Light attenuation in fibers | 0.65 dB/km
Contrast loss due to fibers with 9mm lens | From 34% to 16%

## Slide 6
Isotropic 60Co source facility
Dose rate to reach at least 500 kGy total dose
Reasonable measurement time (days/month)
Reasonable cost
Fibre radiation hardness: intro
Simulating 10 years of operation of fiber bundle
Choice of gamma radiation and 10m bundle
On-line results to monitor behaviour
IRMA facility at CEA Saclay
RADWG Meeting, 04th Dec 2018 D. Celeste, BE-BI-PM

## Slide 7
Fibre radiation hardness: IRMA setup
Dose Rate = 3.6kGy/h
Irradiation Time = 8 d
Relaxation time = 1 d
Acq. Image delay = 15’
Temperature = 24-27 °C
Date/Time Integrated Dose [Gy] | 
Start acquisition: 10/09/2018 14:41:03 | 0
Start irradiation: 10/09/2018 15:14:05 | 0
Stop irradiation: 18/09/2018 15:14:31 | 694.076 E+03
Stop Acquisition: 19/09/2018 14:36:25 | 694.076 E+03
RADWG Meeting, 04th Dec 2018 D. Celeste, BE-BI-PM

## Slide 8
Fibre radiation hardness: results (1)
Before Irradiation (0 Gy) After Irradiation (700 kGy) After relaxation (700 kGy)
1
2
3
1
2
3
NORMALIZATION
RADWG Meeting, 04th Dec 2018 D. Celeste, BE-BI-PM

## Slide 9
Analysis:
Images selected every 10 kGy
Up to Dtot = 240 kGy (3 days)
Normalized to reference
Results:
Center stable
Edges down to 50%
Fibre radiation hardness: results (2)
Analysis:
Images selected every 10 kGy
From Day 3 to Day 8
From 240 kGy to 700 kGy
Normalized to reference
Results:
Center starts decreasing faster
down to 70 %
Edges down to 30 %
RADWG Meeting, 04th Dec 2018 D. Celeste, BE-BI-PM

## Slide 10
Beam imaging test: intro & setup
RADWG Meeting, 04th Dec 2018 D. Celeste, BE-BI-PM
Installation of a imaging fibre on FT16.BTV218 instrument in TT2
Camera
Optics
Fibre ‘IN’
Fibre ‘OUT’
Screen photons
Optics
PROTON BEAM
Equipment:
Analog camera connected to control software
Screen CHROMOX
8mm focal length input
25mm focal length output
10m, 10k optical fibers bundle
Measurements:
Beam LHC_INDIV, intensity ≈ 1 E+11
No control on exp. time nor gain
Modifying trigger delay and filters to avoid saturation

## Slide 11
Beam imaging test: results
RADWG Meeting, 04th Dec 2018 D. Celeste, BE-BI-PM
Acquisition with analog camera system
Acquisition with analog camera and optical fiber
σx = 3.48mm
σx = 4.60mm
σy = 1.41mm
σy = 2.34mm
2018_09_14
OD1
2018_11_06
OD3
100x less light transmitted with fibres
Accuracy of beam size measurement to be checked

## Slide 12
Conclusions and next steps
Conclusions:
Light loss up to 98 % due to optical coupling
Able to compensate with digital parameters
Fiber radiation hardness confirmed
Non-uniform light loss up to 50% due to radiations
Able to reconstruct qualitative images
Tests in TT2 confirmed 100 times less transmitted light
System able to acquire beam images
To be clarified:
Non uniform light loss in RadHard fibres:
  Due to high rate/TID?
  Due to other factors: temperature, cladding,…?
Different beam sigmas:
  Laser measurements to simulate beam?
  Issue with resolution fiber?
RADWG Meeting, 04th Dec 2018 D. Celeste, BE-BI-PM

## Slide 13: Relaxation time profiles

## Slide 14: Spikes vs Temperature

## Slide 15: Temperature corrected profiles

## Slide 16
First and last image form irradiation test in IRMA September 2018
Slice fibre profile

## Slide 17: Picture of the setup: example with two 25mm lens
S1=517mm
S2=38mm
S3=42mm

## Slide 18: Study of the effect of 180° rotation
1st configuration
2nd configuration
Int light: 15858842 a.u.
Int light: 15993512 a.u.
<1%
Reference image:
Target @500mm
CCD camera
One 25mm lens

## Slide 19: Study of the effect of 180° rotation
Reference image:
Target @500mm
CCD camera
One 25mm lens
Chart: Pixels
From 38%
To 52%
Light profile and spatial resolution
Central square

## Slide 20: Beam profile imaging with optical fibers
Lens 1
Opt fiber in
Opt fiber out
Lens 2
CCD camera
Set up description:
LED screen at 500mm from the 1st lens (25mm or 9mm)
The optical fiber interface is placed at a distance d=f to catch the focalized image of the screen, created by the 1st lens.
A 2nd lens is placed at the fiber output to obtain a certain magnification factor of the image on the CCD camera.

## Slide 21: Acquired images: same MAG on camera
Reference image:
Cam + Lens 25mm
Cam + 2 Lens 25mm
Cam + 2 Lens 25mm
Old 10m opt. fiber
New 2m opt. fiber
New 10m opt. fiber
Cam + 1 Lens 9mm + 1 Lens 25mm
Old 10m
opt. fiber
New 2m
opt. fiber
New 10m
opt. fiber

## Slide 22: X-integrated light profile (a.u.)
#Pixel
Reference image one 25mm lens
Image with two 25mm lens

## Slide 23: X-integrated light profile (a.u.)
Reference image one 25mm lens
Image with two 25mm lens
a.u.
Old 10
New 2
New 10
#Pixel

## Slide 24: X-integrated light profile (a.u.)
Reference image one 25mm lens
Image with two 25mm lens
a.u.
Old 10m fib.
New 2m fib.
New 10m fib.
#Pixel
#pixel
a.u.
Old 10
New 2
New 10

## Slide 25: Effect of the presence of the optical fiber
Integral value:
1217222
319084
43197
35316
24443
1
2
3
4
5
Normalized integral value:
1
0.2621
0.0355
0.0290
0.0201
Central square
#Pixel
a.u.
Old 10
New 2
New 10
Reference
Two lens

## Slide 26: Light attenuation in fibers
#pixel
a.u.
New 2m
New 10m
#pixel
a.u.
-12.6%
-15.5%
Light loss ~ 14% in 8m (1.75%/m if linear -> 0.68 dB/km)
New 2m
New 10m
9mm lens
25mm lens

## Slide 27: Reproducible measurements (camera texp = 20 ms)
1) 6820955 a.u. 1) 5743142 a.u.
2) 7005164 a.u. 2) 5921904 a.u.
Differences < 3%
Comparison between two new 2m opt. fibers by Fujikura
Comparison between two new 10m opt. fibers by Fujikura
Differences < 3%

## Slide 28: Resolution loss: analysis of the central square
Cam + Lens 25mm
Cam + 2xLens 25mm
1 Lens 9mm + OptFib+ 1 Lens 25mm
2m New fiber
10m Old fiber
10m New fiber
1 Lens 25mm +
OptFib +
1 Lens 25mm
2m New fiber
10m Old fiber
10m New fiber
Conclusions:
The first main factor worsening the resolution is the presence of the 2nd lens: with respect to the mean value, we go from 90% to 34% approximately (graph A).
The presence of the optical fibers does not drive further resolution losses in the case with two 25mm lens (graph B).
In the case of the 9mm lens before the optical fiber, a further decrease of the spatial resolution is noticed: 16%
A
B
C

## Slide 29: Comparison with Eyepiece
Main goals:
  Increase light collection
  Improve spatial resolution
Choice of the eye piece: f=24mm
Comparison with the 25mm lens

## Slide 30: Comparison with eye-piece
Set-up and measurements:
The 9mm lens is used as 1st optical input in the fiber.
The 2nd camera lens (f=25mm) is substituted with an eye-piece (f=24mm) and the distance with the camera is adjusted to obtain approximately the same MAG.
Two comparative measurements were performed with both the 2m and 10m new optical fibers, with texp=20ms.
Eye-piece 24mm
Camera lens 25mm
2m Opt fib
10m Opt fib
X-Profiles
Integral values [a.u.]
EyeP_2m: 9066809
Lens_2m: 7005164
EyeP_10m: 7515769
Lens_10m: 5921904
+ 29.4%
+ 26.9 %

## Slide 31: Comparison with eye-piece
Eye-piece 24mm
Camera lens 25mm
10m Opt fib
X-Profiles
Spatial resolution value:
From 20% to 30%
(in 2.5mm gaps)
Chart: Pixels
From 10% to 15%
(in 1mm gaps)

## Slide 32: Best images
Target @500mm, texp=20ms, 1st lens 9mm, 2nd lens 25mm
Target @200mm, texp=40ms, 1st lens 9mm, 2nd lens 25mm

## Slide 33: Fiber radiation hardness: mapping test set up
60Co
irradiator
window
source
Optical fiber
Radiation cone
d
Where d is the distance chosen to maximize the dose rate
Ion chamber
Ion chamber
35 cm
40 cm

## Slide 34: Choice of the distance
Distance [cm] | Dose rate [Gy/h]
72 | 4.16
68 | 4.29
66 | 4.73
65 | 4.54
63 | 3.12
60 | 2.30
55 | 0.98
Chart: 
| 72 | 68 | 66 | 65 | 63 | 60 | 55
Dr [Gy/h] | 4.16 | 4.29 | 4.73 | 4.54 | 3.12 | 2.3 | 0.98
Radiation
cone limit
Dose rate
≃ 1/r2
Region out of
radiation field
The distance of 66 cm between the fiber support and the source is chosen
Gy/h
cm

## Slide 35: Uniformity check
Position [deg] |  | Gy/10s | Gy/s | mean | std%
0 | 1 | 14.39 | 5.1804 | 5.27652 | 0.766482
 | 2 | 14.84 | 5.3424 |  | 
 | 3 | 14.64 | 5.2704 |  | 
 | 4 | 14.64 | 5.2704 |  | 
 | 5 | 14.6 | 5.256 |  | 
 | 6 | 14.69 | 5.2884 |  | 
 | 7 | 14.76 | 5.3136 |  | 
 | 8 | 14.61 | 5.2596 |  | 
 | 9 | 14.7 | 5.292 |  | 
 | 10 | 14.7 | 5.292 |  | 
For each point, 10 dose measurements integrated
over 10s were taken and then the avg and the std dev is calculated, as follows:
angle = 0° Dose Rate = 5.43 Gy/h
angle = 180° Dose Rate = 5.03 Gy/h
angle = 90° Dose Rate = 4.40 Gy/h
angle = 270° Dose Rate = 4.89 Gy/h
angle = 45° Dose Rate = 5.20 Gy/h
angle = 315° Dose Rate = 5.23 Gy/h
angle = 225° Dose Rate = 4.94 Gy/h
angle = 135° Dose Rate = 4.94 Gy/h
The obtained map is presented below, where the average dose rate is approximately 5 Gy/h over the all fiber. In this situation at the limit of the radiation cone and very close to the source, even 1-2 cm of displacement can drive a big difference in the dose rate.
The best obtained result, looking for a compromise between
the maximum achievable dose rate and the irradiation uniformity,
is ≃ 5Gy/h
angle = 0° Dose Rate = 5.27 Gy/h
angle = 180° Dose Rate = 5.05 Gy/h
angle = 90° Dose Rate = 4.56 Gy/h
angle = 270° Dose Rate = 4.71 Gy/h
angle = 45° Dose Rate = 5.14 Gy/h
angle = 315° Dose Rate = 5.07 Gy/h
angle = 225° Dose Rate = 4.98 Gy/h
angle = 135° Dose Rate = 4.99 Gy/h

## Slide 36: Starting irradiation test
Testing the new 10m fiber by Fujikura
10 TBq active Co60 source
Integrated dose over 3 weeks: 2 kGy
Reference image taken at the beginning
Acquired image every hour
Light-tight box with all optical components and supports
9mm lens
25mm lens
LED screen
Opt fib OUT
Opt fib IN
CCD cam
Opt fib
bundle

## Slide 37
Increase of the use of high performance digital cameras in BI.
Radiation effect issue mitigated with optical line, shielding and RadHard innovation.
ETH/Fiber converters radiation hardness: intro
Distance to its control front end.
The Gigabit Ethernet allows up to 100m connections.
Not enough in the case of the SPS BTV dump future camera based (BTV) system.
Ethernet to Fibre Optic converter
Eth/fibre
1G/Single ch.
Digital Camera
Eth/fibre
1G/Single ch.
Ethernet
(max 100m)
Dual Fibre SC single mode
Ethernet
PowerPC, Ethernet network, etc.
As SPS machine…test in CHARM facility
RADWG Meeting, 04th Dec 2018 D. Celeste, BE-BI-PM

## Slide 38
Request:
  To reach approximately 50 Gy TID
  RUN1: from 08/08/2018 to 14/08/2018
  RUN2: from 16/08/2018 to 21/08/2018
ETH/Fiber converters radiation hardness: setup
24GeV
~4E11p
400ms spill
Every ~30s
Location of the modules
RADWG Meeting, 04th Dec 2018 D. Celeste, BE-BI-PM
Radiation area CHARM G0
Safe area
Eth/fibre
1G/Single ch.
Digital Camera
Ethernet
Power supply1
Eth/fibre
1G/Single ch.
x2
Fibre type 1
‘GUDE‘ power switch CTRL
Ethernet

## Slide 39
Analysis:
Logging cameras
Logging GUDE power switch
Checking communication failures
ETH/Fiber converters radiation hardness: results
Test procedure:
Image acquisitions every minute
Power cycles at each connection loss
RADWG Meeting, 04th Dec 2018 D. Celeste, BE-BI-PM
SETUP STILL WORKING AT
THE END OF THE TWO RUNS
OK!
 | Module 1 |  | Module 2
Reset number | 28 |  | 22
HEHeq | 8.91E+10 |  | 8.91E+10
[cm2] | 3.14E-10 |  | 2.47E-10
Φfail [cm-2] | 3.18E+09 |  | 4.05E+09
ΦfailCam [cm-2] |  | 1.5 E+09 | 
 | Module 1 |  | Module 2
Reset number | 28 |  | 22
HEHeq | 8.91E+10 |  | 8.91E+10
 | 3.14E-10 |  | 2.47E-10
Φfail [cm-2] | 3.18E+09 |  | 4.05E+09
ΦfailCam [cm-2] |  | 1.5 E+09 |