Cameras at GSI / FAIR
Hardware – Software – Operation
Harald Bräuning
Beam Instrumentation, GSI

---
Stone Age Beam Diagnostics
Scintillating Screen Readout at GSI prior to 2013
Transport Protocol
from Experiment Cave
to Main Control Room
Image Readout
Lots of room for improvement . . .

---
GSI – Overview
GSI accelerates all ions from protons up to uranium
Unilac
SIS18
ESR
Cryring
100m
High Energy Beam Transport
Experiment
Caves

---
GSI – Overview
Unilac
Unilac
SIS18
energies up to 11.4 MeV/u
up to 1012 particles per ms pulse
up to 50 Hz repetition rate
beam size: mm – cm
SIS18
energies up to 4 GeV/u
electron cooler
up to 1011 particles
fast extraction (1µs)
slow extraction up to 10 s
beam size (incl. HEBT): mm – cm
ESR
Cryring

---
GSI – Overview
ESR
Unilac
SIS18
primarily storage ring
stochastic and electron cooling
deceleration to 4 MeV/u
fast extraction to Cryring
slow extraction up to 60s and more
primarily storage ring
stochastic and electron cooling
deceleration to 4 MeV/u
fast extraction to Cryring (few µs)
slow extraction up to 60s and more
Cryring
recent addition (from Stockholm)
test bench for FAIR control system
primarily storage ring
energy range 0.3 – 14 MeV/u
electron cooling
2 injectors
Linac for light ions at 0.3 MeV/u
cooled heavy ions from ESR at 4MeV/u
ESR
Cryring

---
Up..., Up..., Upgrade
2013: Upgrade Cave A
driven by atomic physics experiments
improve beam diagnostics on beamline from ESR to Cave A
low energy ions
low intensity
required for optimizing beam transport and emittance / collimation
6 new scintillating screens on path from ESR to Cave A
5 experiment screens in Cave A (only temporarily installed)
digital readout
Cave A

---
Up..., Up..., Upgrade
Cave A: 190MeV/u U89+
collimated beam on 1st experiment screen
2013: Upgrade Cave A
Requirements from experiments:
slow and fast extraction from ESR
ions up to uranium
down to about 104 particles per second
energies 4 MeV/u up more than 200 MeV/u
simultaneous view in experimental hut and main control room
storage of images for later analysis
measurement of beam position and approximate FWHM in SI units
free run and triggered image acquisition
charge state distribution on screen after analysing magnet

---
Upgrade – Contraints by FAIR
2013: Upgrade Cave A
Prototype for scintillating screen readout in the FAIR accelerator complex!
Constraints imposed by the FAIR Control System:
Client – Server architecture
Server (Frontend Controller)
Linux operating system
based on Cern's FrontEnd
Software Architecture (FESA)
C++
Client (GUI, ...)
based on Java / JavaFX
number of clients is indefinite

---
General Architecture
Client Applications
LHC Software Architecture (LSA)
Ion Optics, Trim, Settings, Scheduling
Automatic setting of beam parameters
(charge state, energy, etc)
Machine Timing Events
Actions
FESA
Hardware Access
Hardware
Common Middelware
LSA Middle Tier

---
Standard Assembly
Linear Feedthrough
mostly driven by pressurized air
some driven by stepper motors
Camera
Remote Controlled Lens
LED for target illumination
Scintillating screen in standardized holder

---
Camera
iDS UI-5240SE-M (monochrome)
	specification	remarks
interface	GigE
sensor	CMOS (e2v)	selection criteria
pixel	1280 x 1024	prog. AOI
pixel pitch	5.3µm
shutter	global
max. framerate	50fps	≤ 10fps
exposure time	9µs – 2s	1ms – 1s
binning	h + v (2x)
analog gain boost	2x
ADC	10bit	8bit readout
I/O	trigger input	trigger by event
	strobe output	monitor
e2v Image Sensors:
high sensitivity
good radiation tolerance
used in a space environment
http://ids-imaging.com

---
Camera
iDS UI-5240SE-M (monochrome)
Standard scintillating screen for FAIR:
chromium doped aluminum oxide Al2O3:Cr (0.04%mass)

---
Camera Readout Hardware
Industrial PC
data acquisition for cameras
via 10 Gbit ethernet switch (local network)
camera configuration
camera readout
via standard LAN (accelerator network)
control of CPS8 controllers
Fair Timing Receiver Node
hardware trigger (LVTTL)
based on machine event
distributed via CPS8 controller
currently: all cameras = one trigger
future: one camera = one trigger
10 Gbit Ethernet Switch

---
Camera Readout Hardware
New Systems: µTCA based
compact modular system
8 Port Gbit ethernet switch
data acquisition from cameras
multiple switches in principle possible
Fair Timing Receiver Node
disadvantages:
complicated configuration of µTCA backplane ethernet lines
at first suboptimal configuration:
system worked
30s discovery time of cameras after camera power cycle
(typical with optimal configuration: 2 – 3s)
Fair Timing Receiver Node
8 Port Gbit Ethernet Switch

---
Camera Control Hardware
CPS8 Controller
build in-house
Atmel Microcontroller (Arduino) based
simple ethernet based access
camera power control
camera power status
camera trigger
software trigger
hardware trigger distribution
currently: all cameras = one trigger
future: one camera = one trigger
further options like burst mode could be implemented
front panel indicators
camera power
camera exposure (camera strobe output)

---
Remote Controlled Lens
Remote Controlled Lenses
mainly to adjust image intensity via iris
occasionally focus in case of different screen positions
i.e. Cryring: different correct beam position on injection screen depending on source path
Pentax -ER Series
iris control by DC voltage (1.5 – 5.5V)
iris closes when power is cut off
manual focus
(usually sufficient)
DISCONTINUED
(but we have 60 pcs in reserve)
Linos MeVis-Cm
intended as replacement
motorized iris
motorized focus
encoders for precise and reproducible adjustment
DISCONTINUED

---
Remote Controlled Lens
Remote Controlled Lenses
Controller
Satellites
Digital Out
Power Stage
PID Controllers
MeVis-Cm Control
Programmable Logic
Controller (PLC)
Pentax Control
Analog Out
Digital Out

---
Complete System Overview
Control Rooms
Electronics Room
Radiation
Save Area
Experiment /
Beamline
camera image / settings
trigger
power
timing network
profinet
CPS8
PLC
Main Controller
PLC Satellites
User Clients
10 Gbit Switch
(local network)
Industrial PC / mTCA
camera data acq.
PLC control
Timing Receiver
accelerator network
local 10 Gbit network

---
Frame Grabber
Analog Cameras
legacy cameras already installed
Thermo Fischer RadHard camera (MegaRAD3)
Pleora iPort Analog-Pro External Frame Grabber
Gbit ethernet based
GigE Vision and GenICam compatible
Analog cameras run continuously
rate reduction in software to 10Hz
trigger by software
all frames are discarded but next frame after trigger

---
Client – User Application

---
Server – Data Acquisition
Data Acquisition on the Frontend: Server for Multiple clients
Image Acquisition and Processing
image acquisition
software crop
rotation (90°, 180°, 270°)
mirror
simple image scaling
projections
simple analysis neglecting further perspective distortion
histogram
position of brigthest pixel
center-of-brightness
FWHM
Camera / Software Settings (if supported)
acquisition mode (free running / triggered)
exposure time
binning
gain boost
contrast
brigthness
rotation / mirror
area-of-interest (fixed during runtime)

---
Perspective Distortion
Perspective Distortion
scintillating screen is typically mounted under 45° with respect to the optical axis
short distance leads to visible perspective distortion
position (x,y) in SI units can be calculated for each pixel by
we do not rectify the image!
x'_p ~=~ x_p - x_p0
y'_p ~=~ y_p - y_p0
x ~=~ ( s_x + %DELTA s_xx x'_p + %DELTA s_xy y'_p) x'_p
y ~=~ ( s_y + %DELTA s_yx x'_p + %DELTA s_yy y'_p) y'_p
xp,yp = pixel position; xp0,yp0 = pixel position of screen center
s,Δs = scale factors

---
Perspective Distortion
Perspective Distortion
scintillating screen is typically mounted under 45° with respect to the optical axis
standard target holder contains alignment marks
all scale factors are determined from the alignment marks
vertically mounted assemblies
Δsxx = 0 and Δsyx = 0
horizontally mounted assemblies
Δsxy = 0 and Δsyy = 0
x'_p ~=~ x_p - x_p0
y'_p ~=~ y_p - y_p0
x ~=~ ( s_x + %DELTA s_xx x'_p + %DELTA s_xy y'_p) x'_p
y ~=~ ( s_y + %DELTA s_yx x'_p + %DELTA s_yy y'_p) y'_p

---
Perspective Distortion
Perspective Distortion
perspective distortion is solely handled by the client
overlay a grid showing the perspective distortion
display pixel coordinates in SI units taking perspective into account
x'_p ~=~ x_p - x_p0
y'_p ~=~ y_p - y_p0
x ~=~ ( s_x + %DELTA s_xx x'_p + %DELTA s_xy y'_p) x'_p
y ~=~ ( s_y + %DELTA s_yx x'_p + %DELTA s_yy y'_p) y'_p

---
Experiences
iDS Ueye
operational: 34 cameras
stable operation
no degradation due to radiation visible so far
well supported under Linux
firmware update in situ possible
Problem (more a really not understood / further investigated observation)
saturation with max pixel value <8bit (255) for short pulses (few µs length)

---
Experiences
Pleora iPort Analog-Pro
operational: 3 cameras
well supported under Linux
some stability problems
frame grabber stops sending images
maybe due to too many dropped packets?
bad µTCA configuration?
frame grabber internal counter, which stops data receiver on too man packet resends?

---
END
Special thanks to...
André Petit – PLC programming
Rolf Lonsing – PLC and electronics
Christian Schmidt – CPS8 controller box
Christoph Dorn and group – mechanics
Beata Walaseck-Höhne - Coordination
view into the Cryring ion source
Thank You for Your Attention...

---
Full Width Half Maximum
Full Width Half Maximum is calculated on the profiles
1. Simple Algorithm
find position and value of brigthest pixel
find value of dimmest, non-zero pixel
go to both sides of brigthest pixel until value drops below half of distance between maximum value and mininimum value
Fails on older scintillating screens with etched-on grid!

---
Full Width Half Maximum
Full Width Half Maximum is calculated on the profiles
2. Standard Deviation
find position and value of brigthest pixel
find value of dimmest, non-zero pixel
calculate standard deviation of profile's center-of-brightness
use only data points with
v ~>~ v_min + ( v_max - v_min) / 4

---
GSI accelerates all ions from protons up to Uranium
SIS18
ESR
Experiment
Caves
Cryring
Unilac
100m
High Energy Beam Transport

---
Stone Age Beam Diagnostics
Scintillating Screen Readout at GSI prior to 2013
Lots of room for improvement . . .