## Slide 1: Beam Instrumentation developments at HiRadMat
T. Lefevre on behalf of the BI Team involved
Thanks to the contributions of J. Albertone, M. Bergamaschi, C. Boccard, T. Bogey, S. Burger, A. Curcio, L. Jensen, R. Kieffer, M. Krupa, S. Mazzoni, A. Sounas, A. Topaloudis, E. Senes, M. Turner and M. Wendt
International HiRadMat Workshop, 10-12 July 2019, CERN

## Slide 2
Outline
Existing beam instrumentation
Beam instrumentation R&D
  Test-stand
  Example of future developments
Conclusions

## Slide 3: BI close to experiments
In October 2014, a request came to improve the Beam monitoring at experiment
Motivation :
  Measure the Beam position w.r.t experiments better than 0.1mm
  Measure shot-to-shot beam position fluctuations
  Measure transverse beam size on target
Decided to built a combined BTV/BPM station
  3 ensembles built and mounted on HiRadMat experiments

## Slide 4: BTV/BPM assembly
Al2O3:CrO2, SiC and Al2O3 (target) screens

## Slide 5: BTV/BPM assembly
BTV is used for setting up and absolute calibration
  Calibration Screen pre-aligned to the experiments on surface – Used for reference
  Offset between BTV and BPKG measured using low beam intensities
BPM providing bunch-by-bunch on-line measurements up to highest beam intensities
No beam profile for high intensity beams

## Slide 6: Particle showers affecting BTV
Scintillating screen - Image acquired synchronously with beam
back-scatterred shower from the target
Scintillating screen - Image acquired with 100ms delay after beam passage
No noise on the image
Observing the light emitted during the decay time of the screen
Scintillating screen - Al2O3:CrO2
  Non linearity above 1E12 p/mm2
  Damage above 1E13 p/mm2

## Slide 7: Particle showers affecting BTV
Use of OTR screens is limited as acquisition must be synchronized with beam (backscattering particles gives high background on camera !)
12 bunches (1.2E12p)
SiC screen
T 0.001%

## Slide 8: Cherenkov background in -air
Strong background from Cherenkov light emitted in air
Cherenkov emission
Energy threshold in air: 37.4GeV
Vacuum
Air
OTR screen
Proton Beam
OTR Light

## Slide 9
Strong background from Cherenkov light emitted in air
Cherenkov emission
Energy threshold in air: 37.4GeV
Vacuum
Air
Proton Beam
‘Blocking foil’
OTR screen
OTR Light
Cherenkov background in -air

## Slide 10: Cherenkov background in -air
Strong background from Cherenkov light emitted in air
Oct 2015 – HRMT30
June 2016 – HRMT32
* Scintillation and Otr Screen Characterization with a 440 GeV/c Proton Beam in air at the Cern HiRadMat Facility
S. Burger†, M.Turner2, B. Biskup1, S.Mazzoni, CERN, Geneva, Switzerland 1also at Czech Technical University, Prague, Czech Republic 2also at Graz University of Technology Theoretical Physics Institute, Vienna, Austria

## Slide 11: Cherenkov background in -air
Strong background from Cherenkov light emitted in air
Oct 2015 – HRMT30
June 2016 – HRMT32
* Scintillation and Otr Screen Characterization with a 440 GeV/c Proton Beam in air at the Cern HiRadMat Facility
S. Burger†, M.Turner2, B. Biskup1, S.Mazzoni, CERN, Geneva, Switzerland 1also at Czech Technical University, Prague, Czech Republic 2also at Graz University of Technology Theoretical Physics Institute, Vienna, Austria
OTR

## Slide 12: BPM Performance
Some issues observed in fall 2015 when testing high-Z targets
  Traced to be due to backscattered particles affecting the electronic
  Mitigation found by relocating the electronic in a radiation safe zone during YETS 2015-2016
Still some perturbations observed on the beam position signal resulting in large fluctuations in the BPM readings (±mm)

## Slide 13: BPM Performance
Some issues observed in 2015 when testing high-Z target
  Traced to be due to backscattered particles affecting the electronic
  Mitigation found by relocating the electronic in a radiation safe zone
Still some perturbations observed on the beam position signal resulting in large fluctuations in the BPM readings
Development of a BTV system capable of measuring in all beam conditions
Small beam size at high intensities

## Slide 14: BTV for all beam conditions
Long optical line to install the camera in a radiation-friendly location (showers and ageing)
Thermal resistant OTR screen in primary vacuum to suppress Cherenkov background light (<10-3 Torr)
Aging of the Camera damaged by radiation by the end of HRMT27

## Slide 15: BTV for all beam conditions
Camera installed in TT61 parallel tunnel
New BTV installed on Table A
Camera
Shielding
TT61
TNC

## Slide 16: BTV for all beam conditions
> 5m long optical line
Optical magnification 0.2
Achromatic Lens Focal 1m, Diam: 80mm
Screen position IN
Beam
Mirror
Diameter 4”
Camera
WATEC ULT902H3
~5m
~0.9m
~0.3m
Resolution [um/px] | 
Horizontal | 72
Vertical | 79

## Slide 17: BTV for all beam conditions
p BEAM
Photons
Installation of the new BTV HRM from 2017 Located end of experimental table A
Image from V. Clerc
Beam
Photons
CATIA 3D drawing of the new BTV for HRM

## Slide 18: BTV for all beam conditions
Sigradur G Glassy C
0.5mm
Chromox
0.5mm
Ti
0.1mm
SiC
0.5mm
Fixed 75um Carbon foil
Beam

## Slide 19: BTV for all beam conditions
SiC broken after 3 impacts at full beam power density
288 bunches at 1E11 with σH = 0.26mm and σV = 0.28mm
Chart: Sensitivity
| CHROMOX 400um | CHROMOX 250um | Glassy C | Diamond
 | 100 | 28.3199 | 0.60638 | 0.018499

## Slide 20: BTV for all beam conditions
SiC broken after 3 impacts at full beam power density
Glassy C has shown reliable measurements down to smallest beam sizes of 0.25mm
  Systematic over-estimation due to Forward OTR from blocking foil simulated using ZEMAX to be less than 10%
288 bunches at 1E11 with σH = 0.26mm and σV = 0.28mm

## Slide 21
BI developments at HiRadMat

## Slide 22
BI developments at HiRadMat
Using Table a as an in-air BI test-stand

## Slide 23
BI developments at HiRadMat
1m long in-air testing area on CLEAR/CERN (200MeV electrons – low charge)
  Equipped with beam position monitor and Screen/camera
  Equipped with movers and motor controller

## Slide 24
BI developments at HiRadMat
Testing Directivity of Hilumi new cold Stripline BPM
  Expected directivity better than > 25dB
  Prototype under fabrication

## Slide 25
BI developments at HiRadMat
Testing Electro-optical BPM for Hilumi
  Prototype tested in SPS for measuring intrabunch instability
  New compact design underway – prototype expected in 2020
25
SPS protoype

## Slide 26
BI developments at HiRadMat
R&D on Incoherent Cherenkov-Diffraction Radiation
  Beam profile monitoring for highly relativistic particles
Positron Beam
Positron Beam
Cherenkov
Diffraction
Radiation
R. Kieffer et al., “Direct Observation of Incoherent Cherenkov Diffraction Radiation in the Visible Range”, PRL 121 (2018) 054802
5.3GeV Positrons at CESR propagating at 3mm from the surface of a 2cm long fused silica prism

## Slide 27
BI developments at HiRadMat
R&D on Incoherent Cherenkov-Diffraction Radiation
  Beam position monitoring in long(er) dielectrics
Motivation
Increasing the photon yield to detect particles sitting at larger distances (h > 1cm)
Investigating measuring ChDR in longer wavelength (NIR) for low(er) beam energies (e-@200MeV or H+@400GeV)
L ????
h
Incoherent Cherenkov DR
photons
Dielectric
E Field
Vacuum

## Slide 28
BI developments at HiRadMat
R&D on Incoherent Cherenkov-Diffraction Radiation
  Beam position monitoring in long(er) dielectrics
20cm long stripline in Fused Silica

## Slide 29
BI developments at HiRadMat
R&D on Incoherent Cherenkov-Diffraction Radiation
  Beam position monitoring in long(er) dielectrics
First test in dec. 2018 at CLEAR using a train of 50pC electron bunches (1nC total) 0 charge)

## Slide 30
BI developments at HiRadMat
R&D on Incoherent Cherenkov-Diffraction Radiation
  Beam position monitoring in long(er) dielectrics
First test using train of 50pC electron bunch (1nC total charge)
R&D for high energy collider studies (LC or FCC)
But also instrumentation for crystal collimator @ LHC

## Slide 31: Conclusions
Since 2014, significant improvements in measuring high intensity beam properties in HiRadMat
Very useful R&D performed already on screen and imaging systems (Many thanks to the HiRadMat team for help and support)
Instrumenting Table A with an in-air test-stand is an efficient and unique opportunity for beam instrumentation developments !
Most of the testing can be run almost parasitically to other HiRadMat experiments.
LoI under preparation

## Slide 32: Thanks for your attention
Many Thanks to all colleagues involved !

## Slide 33: BTV/BPKG assembly

## Slide 34: Measurement setup
Diaphragm
beam
Al foils
Oct 2015 – HRMT30
June 2016 – HRMT32

## Slide 35: Screens tested
Scintillator:
  Alumina (Al2O3) – Purity: 96% & 99%
  Chromox (Al2O3:Cr2O3)
  thicknesses: 3 mm, 1 mm, 0.5 mm
  YAG – thicknesses: 0.1 mm, 0.5 mm
  both coated and uncoated
OTR
  Titanium, Aluminium, Silicon coated by aluminium & silver

## Slide 36: Results (Nov. 15-HRMT30)
Different luminescent/scintillating screens
Chromox 1 mm thick screen was chosen to be used for the AWAKE halo measurement (as the best compromise between a high light yield and sufficient resolution)
For reflective OTR screens, measurements limited by some parasitic light of unknown origin (cherenkov in air ?)

## Slide 37: BTV for all beam conditions
Chart: Sensitivity
| CHROMOX 400um | CHROMOX 250um | Glassy C | Diamond
 | 100 | 28.3199 | 0.60638 | 0.018499

## Slide 38: Issues with BPKG (1/4)
Cross-calibration between BTV ad BPKG went well
But we then discovered an issue while using long(er) bunch trains with large
unreal beam offsets measured
Published BPM data is the average of the beam position of all bunches
Target in
Target out
12 bunches
12 bunches

## Slide 39: Issues with BPKG (2/4)
Target in
Target out
BPM508
BPM517
BPM408
Strong perturbations coming after the first 9 bunches
Seen on all BPMs down to the BPM408 in TJ7

## Slide 40: Issues with BPKG (3/4)
The electronics of all these BPMs sit in TJ7 underneath the BPM408
Perturbations compatible with backscattered shower affecting the electronic. Confirmed by time of flight
Relocation of BPM electronic in less exposed area
  Tunnel access to TJ7
  Done during EYETS 15-16

## Slide 41: BTV/BPKG cross-calibration @HRMT28
Cross-alignment between BPKG.524 & BTV.524
Bumps ±1,2 mm
Intensity ~ 1e11 p
Identifying central offsets and scaling issues

## Slide 42: Correcting BPM geometrical offsets
First shots before changing offsets
Offset on the horizontal position
Shots after changing offsets

## Slide 43: Reproducibility study
One nominal bunch

## Slide 44: Steering along the line
Beam Direction

## Slide 45: Reproducibility study – longer trains
72 -bunches
288 -bunches

## Slide 46: Steering along the line
72 -bunches
288 -bunches

## Slide 47: Test with Collimator in – long trains
BPKG.524
BPM.517
BPM.408
No wrong reading anymore due to radiation effects

## Slide 48: BLM test @ HRMT19-BLM2
LHC ionization chamber (IC) : 1.5l active vol.@ 1.1bar
Little Ionization Chamber (LIC) : 0.05l active vol.@ 1.1bar
for LHC in high losses region
Flat Ionization Chamber (FIC) : 0.05l active vol.@ 1.1bar
new detector for Booster
Tested with 3-5 shots using respectively
1, 36, 72, 144, 216, 288 nominal bunches
Measuring linearity and dynamic range

## Slide 49: BLM test @ HRMT19-BLM2
IC
FIC
LIC
FIC
LIC

## Slide 50: Diamond Detector (1/2)
Requirements:
Position, and size precision of 0.1mm
Simple real-time acquisition of beam parameters
Single bunch resolution
intensity range 1E9 to 5E13 p+/pulse
Based on Beam halo measurement

## Slide 51: Diamond Detector (2/2)
Beam Size
Measuring BEAM SIZE
Assuming circular spot
Measurements for beam sizes of 1mm and 2mm
  2 mm data taken with W-target in beam
Standard deviation 0.07 mm for 1 mm optics, 0.36 mm for 2 mm optics
Measuring BEAM POSITION
Comparing with BTV
± 1s standard deviation is plotted
1s = 0.12mm
Working as well with W-target inserted with reduced resolution (temporal signal distorted by showers)

## Slide 52: R&D performed in 2017-18
Development of non-invasive beam instruments using incoherent Cherenkov Diffraction Radiation in long(er) dielectrics
LHC 450GeV
LHC 7TeV
FCC
Beam Energy (MeV)
Nph
per protorn
h
Incoherent Cherenkov DR
photons
Dielectric
E Field
Vacuum
L = 1m & h = 2mm
For comparison Nph from OTR is 10-2 -10-3

## Slide 53: R&D performed in 2017-18
Detection
in the IR
Wavelength (m)
Photon
Spectrum (a.u)
Development of non-invasive beam instruments using incoherent Cherenkov Diffraction Radiation in long(er) dielectrics
h
Incoherent Cherenkov DR
photons
Dielectric
E Field
Vacuum
L = 1m & h = 2mm
LHC 450GeV
LHC 7TeV
FCC

## Slide 54
BI developments at HiRadMat
R&D on Coherent Cherenkov-Diffraction Radiation
  BPM for Electrons which should be insensitive to Protons
Current system
not working !!
Simulations by E. Senes

## Slide 55
BI developments at HiRadMat
R&D on Coherent Cherenkov-Diffraction Radiation
  Bunch length and Beam position monitor for short bunches (sub-ns)
Bunch spectrum
Bunch position monitor using coherent Cherenkov-Diffraction radiation at >20GHz (in 2018)
E. Senes , A. Curcio, M. Bergamaschi

## Slide 56: Developing Beam diagnostics at CERN
Testing directly on the Operational Machines themselves
  It works..but may lead to unpleasant surprises
    e.g. Beam position dependency of CERN Fast Beam Current Transformer on LHC
  Limited time for hardware installation/modification in the tunnel (i.e. Technical stops)
  Limited beam time available for tests during MDs
  R&D is rarely compatible with the strict requirements for Operational Machines
  e.g. Testing gas jet monitor and their performance
  as function of gas pressure would conflict with
  vacuum requirements