## Slide 1: Operation of a Cryogenic Current Comparator for continuous beam intensity measurements in the AD
Miguel Fernandes
Beam Instrumentation Group Seminar
BI seminar
30/11/2018

## Slide 2: The Antiproton Decelerator (AD)

## Slide 3: Antiproton Decelerator (AD)
C = 182.5 m

## Slide 4: Antiproton Decelerator (AD)
C = 182.5 m
ELENA
C = 30.4 m
ELENA
Reduce the energy of antiprotons coming from AD from 5.3 MeV to 100 keV
Should allow the experiments to trap
( ~ x 100) more antiprotons

## Slide 5: AD beam and cycle parameters
AD beam parameters | 
Beta | (0.97 - 0.11) c
Cycle | ~110 s
Frev | (1.59 - 0.17) MHz
N particles | (5 - 1) x 107
Current | (12 - 0.1) μA

## Slide 6: Beam intensity monitoring
  DCCT:
  Insufficient resolution: > 1µA
  Fast BCTs:
  Limited to bunched phases
  Schottky monitor:
10
20
30
40
50
60
70
80
90
time [s]
0
1
2
3
4
N [x 107 pbar]
0.0
0.2
0.4
0.6
1.0
β [c]
0.8

## Slide 7: Beam intensity monitoring
  DCCT:
  Insufficient resolution: > 1µA
  Fast BCTs:
  Limited to bunched phases
  Schottky monitor:
10
20
30
40
50
60
70
80
90
time [s]
0
1
2
3
4
N [x 107 pbar]
0.0
0.2
0.4
0.6
1.0
β [c]
0.8
Coasting beam plateaus
  Un-bunched:
    time resolution of
    accuracy error > 10%

## Slide 8: Beam intensity monitoring
  DCCT:
  Insufficient resolution: > 1µA
  Fast BCTs:
  Limited to bunched phases
  Schottky monitor:
10
20
30
40
50
60
70
80
90
time [s]
0
1
2
3
4
N [x 107 pbar]
0.0
0.2
0.4
0.6
1.0
β [c]
0.8
Bunched beam ramps
  Bunched:
    time resolution of
    accuracy error of <10%
    Bunch shape dependent

## Slide 9: Specifications for new monitor
Current/intensity measurement:
Measure beam: Bunched and debunched
Current resolution: < 10 nA
Intensity resolution: < 5 x 105 charges
Bandwidth: DC - 1 kHz
Operations ready:
Integrated acquisition: FESA based
Automatic operation: synchronized with AD cycle
Requirements for the cryostat
“Zero-boil off” using a pulse tube cryocooler as He reliquefier unit
Stand-alone long term availability
Collaboration:

## Slide 10: Overview of the CCC

## Slide 11: CCC functioning overview
Magnetic shield:
  Suppresses all field components except azimuthal beam component
Pickup coil:
  Soft ferromagnetic material with high-permeability concentrates flux
Flux transformer:
  Couples magnetic flux (down to DC) to SQUID
SQUID + Electronic readout:
  Superconducting QUantum Interference Devices
  Measures the magnetic field induced in the SQUID’s input coil
Superconducting
Room temperature

## Slide 12: SQUID’s in a nutshell
Superconductors:
  Zero resistance to electrical currents
  Magnetic field expulsion from bulk material
  Conservation of magnetic flux in closed loops
Current carriers are bounded states of two electrons (Cooper pairs)
From: https://dc.edu.au/hsc-physics-ideas-to-implementation/
Copper pair

## Slide 13: SQUID’s in a nutshell
Ic
Josephson junction:
  Mixed superconducting and resistive current flow
Tunneling of electron-pairs allows for a superconducitng current through a barrier
Tunneling of electrons resulting from pair break-up approximates an ohmic current **
*
* Only for certain parameters of Josephson junction. In general I-V curve is hysteretic.
** For small V

## Slide 14: SQUID’s in a nutshell
  dc-SQUID
  Two Josephson junctions in parallel
2Ic
Ibias
No external flux applied:
Applying external flux:
Constant magnetic flux threading closed loops:

## Slide 15: SQUID’s in a nutshell
for A=1mm2

## Slide 16: SQUID’s in a nutshell

## Slide 17: SQUID devices
SQUID’s are very sensitive magnetometers
Periodic voltage-flux transfer function with period (flux-quanta)
Voltage output of the order of
Periodic transfer function strongly limits its dynamic range

## Slide 18: SQUID devices
Constant Working Point
Linearization of the transfer function using a feedback flux lock loop (FLL)
Total flux in SQUID is kept constant at a fixed working point
Possible to resolve variations of
Dynamic range increased to up to 120 dB

## Slide 19: SQUID devices
Flux jumps of working point may occur
If feedback loop is too slow to track input signal
If too much flux noise is coupled to the SQUID
Increasing FLL bandwidth makes it faster, but also adds more flux noise
Necessary to limit slew-rate of coupled signal to avoid flux-jumps

## Slide 20: SQUID devices
Flux jumps of working point may occur
If feedback loop is too slow to track input signal
If too much flux noise is coupled to the SQUID
Increasing FLL bandwidth makes it faster, but also adds more flux noise
Necessary to limit slew-rate of coupled signal to avoid flux-jumps

## Slide 21: Adaptation of the CCC to the AD beam

## Slide 22: AD beam slew-rate
= 10 φ0/μA
Filtering | Slew-rate
None | 86 Gφ0/s
Core permeability | 1.2 Gφ0/s
Flux slew-rate at SQUID

## Slide 23: Filtering beam current signal
SC
Low pass filtering to reduce signal slew-rate
Maximum possible bandwidth

## Slide 24: CCC components
Magnetic shield had previously been fabricated by GSI and Univesity Jena as prototype for FAIR

## Slide 25: CCC components

## Slide 26: CCC components
Magnicon SQUID and FLL electronics system

## Slide 27: Cryostat design

## Slide 28: Cryostat design
New custom cryostat was designed and fabricated to host CCC monitor
Insulating ceramic integrated to prevent mirror currents from shielding the beam signal
Low heat-load to allow stand-alone operation
Cooling power entirely provided by pulse-tube cryocooler
HV Support designed to mitigate vibration transmission
VV: Vacuum Vessel
TS: Thermal Shield
HV: Helium Vessel
BC: Bayonet Connection

## Slide 29: Cryostat design
LHe vessel support
Closed circuit cooling of termal shield

## Slide 30: Cryostat installed in AD

## Slide 31: CCC Acquisition and Control

## Slide 32: Acquisition and controls
FLL Electronics
Magnicon (20MHz)
Settings:
  RFLL = 1 kOhm
  GBW = 1.04 GHz
Advantech
Serial Server
Calibration Source
Triax
VME Crate
CPU L866
CTRV timing
VD80 ADC
  16 channels
  16 bit
  +/- 10V
  800kOhm
  Isolated diff. inputs
RS-232
Ethernet
Calibration loop
NIM Crate
AD B-field cycle
Current/Nparticlenormalization
Acquisition / Controls
VME architecture
FESA SW + Expert GUI
Signal acquisition at end of cycle and on-line
Control of multiple SQUID/FLL parameters
Low-pass
filter

## Slide 33: AD cycle synchronization

## Slide 34: FESA Class
σtotal = 5.8 nA
BEAM CURRENT [uA]
NUMBER OF PARTICLES
AD MAGNETIC CYCKE

## Slide 35: Expert GUI

## Slide 36: Beam measurements and performance

## Slide 37: Beam current measurement
σ = 5.8 nA
Resolution within < 10 nA (with cryocooler and vacuum pumps running)
Flux jump at injection causes loss of baseline
SQUID/FLL stable throughout the rest of the cycle

## Slide 38: Beam intensity measurement
σ = 1.2x104
Baseline corrected at the end of the cycle
Beam intensity calculated by normalizing against fRev
Intensity resolutions at low-energy ~105 charges
σ = 1.3x105

## Slide 39: Comparison with Schottky monitor

## Slide 40: Comparison with Schottky monitor

## Slide 41: Comparison with Schottky monitor

## Slide 42
AD
B. Lefort
Ref: https://indico.cern.ch/event/641870/contributions/2619208/

## Slide 43: Performance analysis
Measurement resolution
Computed as the signal standard deviation in 1s interval
Analysis of ~25000 cycles
With cryostat insulation vacuum pumps on/off
Corresponds to resolution in terms of number of particles of:
  N = (1-3) x 104 at injection energy
  N = (1-3) x 105 at ejection energy
Current resolution

## Slide 44: Performance analysis
Baseline variation
Slower baseline drifts are dominant measurement error
These are mainly caused by helium gas pressure variations
Additionally perturbation induced by cycling the acccelerator
Baseline drift obtained from cycles with no injected beam

## Slide 45: Extra: Noise performance improvements
σ = 5.5 nA
Perturbation from magnetic cycle
Perturbation from cryocooler

## Slide 46: Conclusions
Measurement noise performance under initial specification
Able to stably cope with AD bunched beams (except at injection)
Very good immunity to mechanical vibrations
Cryogenic system enables “long” term operations (~3/4 month)
Assessed long term stability over entire year run
Automatic system control with FESA class and expert GUI
On-going work
Investigate source of flux jump at injection
Improve availability of cryogenic system
Mitigate perturbations by controlling and stabilizing cryostat pressure
Proposal of a new CCC monitor for the ELENA

## Slide 47: Acknowledgements
J. Tan, S. Pederson, D. Alves, E. Oponowicz, M. Ludwig, T. Koettig, A. Lees, A. Vacca, F. Carrasco,
J.P. Brachet, D. Lombard, G. Foffano, R. Speroni
C. Welsch
T. Schwickert; T. Sieber; F. Kurian;
R. Geithner; R. Neubert , J. Golm
Thank you for your attention!

## Slide 48: Backup

## Slide 49: Coupling circuit low-pass filter
Filtering | Slew-rate
Core permeability + RC-parallel | 0.97 Mφ0/s
Additional flux noise at low freq:

## Slide 50: Calibration vs. Temperature
0.3 (nA/φ0) / K
Ib = 10uA
σ(Ib) < 10nA
δT < 300 mK

## Slide 51: Flux jumps

## Slide 52: Cryogenic availability
2016
Started actively pumping the insulation vacuum
Air contamination at end of year
2017
Power glitch stopped turbo pump that compromised vacum
Still possible to work with gas
~1 month unavailable
2018
New controls for vacuum pumps and vacuum valves
New remotely controlled gas flow-valve was installed
~2 month unavailable

## Slide 53: Stabilize cryostat operation

## Slide 54: Stabilize cryostat pressure

## Slide 55: Immunity to mechanical vibrations
Main vibration modes as captured by SQUID:
  65 Hz
  115 Hz
  78 Hz
  5 Hz