## Slide 1: Experimental results in 2018 from Cockcroft
H. Zhang, A. Salehilashkajani, C. Welsch

## Slide 2
Gas jet in fluorescent mode v2
  System highlights
  Vacuum condition test
  Beam imaging using N2 gas jet and Characterization of N2 jet.
  Beam imaging using Neon gas and comparison to Nitrogen
  Vacuum condition test aiming for LHC installation
Future plan
Outline
Notes: The motivation for this project is to develop a non-invasive beam profile monitor for hollow Electron lens and HLLHC. Although it is covered before, we will briefly talk about the principle of this supersonic gas jet beam profile monitor based on beam induced fluorescence mode.
Then we will discuss the progress of a prototype monitor we have built last recently as well as the experimental program.

## Slide 3
3
Building a prototype in CI
63 mm
63 mm
Notes:   Better pumping 2hour to start E-gun. Overnight to start experiment program.
Resolution up to 20 lp/mm
i.e. 50 µm for the beam
– Magnification @ MCP: 1.2 ± 0.1
Note: The DOF has been estimated to be 4.5 ± 0.5 mm

## Slide 4: Nozzle and skimmers
~45 degree tilted
Notes: The new nozzle design allow bake-out

## Slide 5: Blackening of the interaction chamber
Notes: The graphite spray is used for black coating. The test did in CERN shows a reflection of 13%. This only applied to the interaction chamber.

## Slide 6: Electron gun
Energy: 100 eV to 10 keV
Current: 200uA to 10mA
Spot size: 1.5mm to 20 mm
Notes: From Kimball physics. For neon measurement, extra aperture is used to reduce the light from the hot filament of the electron gun.

## Slide 7: Electron beam diagnostics
Phosphor screen
Faraday Cup
20 mm
115 mm
20 mm

## Slide 8: Vacuum test of current system
Chart: Pressure in the interaction chamber vs inlet pressure
| 0.5 | 1 | 2 | 5 | 8 | 9.9
 | 5.82e-09 | 5.86e-09 | 5.98e-09 | 6.08e-09 | 6.22e-09 | 6.33e-09
Chart: Pressure in the interaction chamber vs inlet pressure
| 0.5 | 1 | 2 | 5 | 8
 | 1.06e-08 | 1.03e-08 | 1.05e-08 | 1.07e-08 | 1.07e-08
E-gun off
E-gun on
Notes: Need a curve of pressure for interaction chamber and dump chamber

## Slide 9: Nitrogen gas jet test
2.59 mm
Electron energy: 5keV
Beam current:0.69mA
Integration time: 2 s
Gas jet type: N2
Inlet pressure: 5bar
2s integration time to give a profile
More time to give a 2D detailed image
Gas jet direction
Electron beam direction
200s integration time
Background pressure: 1.56E-08mbar

## Slide 10: Moveable gauge chamber
224 mm
265 mm
4 mm
0.4 mm

## Slide 11: Interaction chamber
Assume linear expansion from the 3rd skimmer
The expansion is decoupled in both transverse direction
Average density = 9.0 e14 m-3
224 mm
265 mm
4 mm
0.4 mm

## Slide 12: Residual gas calibration
Fitting equation:
Gaussian equation + linear background
Beam size
Photon number
Background pressure 2.5e-7mbar
Number of photons = 4.60E4
Number of photons per pixel per second = 0.82
140 pixels

## Slide 13: Gas jet image of a vertical 3rd skimmer
FWHM = 1.24 mm
Number of photons = 1.57e4
Number of photons per second per pixel = 0.117
Fitting equation:
Gaussian equation + linear background
Beam size
Photon number
1.24 mm
24.3mm
Gas jet image with Inlet pressure of 5bar
Integration time = 1600seconds
Electron beam energy = 5kev
electron beam current = 0.66mA
Third skimmer = vertical
~0 degree tilted
Notes: Need to do a flat top analysis to get the FWHM.

## Slide 14: Photon Gauge test
Chart: Pressure VS photon count
| 9.2888269752821842E-2 | 0.26503786255179312 | 0.82154593513359064 | 1.6459494213459067 | 2.5342906129447069 | 5.6650950135733682
 | 2.35e-08 | 7.55e-08 | 2.5e-07 | 5e-07 | 7.5e-07 | 2.25e-06
Chart: Number density VS photon count
| 9.2888269752821842E-2 | 0.26503786255179312 | 0.82154593513359064 | 1.6459494213459067 | 2.5342906129447069 | 5.6650950135733682
 | 5.6763e+14 | 1823671497584541 | 6038647342995168 | 12077294685990336 | 18115942028985508 | 54347826086956520
For gas jet image: The photon rate per pixel = 0.117
Gas jet average number density = 1.12e15 m-3 or 4.62e-8 mbar

## Slide 15: Cross section measurement
0.96e15 ±0.2e15 m-3
0.98e-3 ± 0.2e-3 m
0.0059
0.75 ± 0.05
0.45 ± 0.05
0.75 ± 0.15
ηPC = 0.19 ± 0.02
N = 15700 ± 100
Δt = 1600 s
Cross section: 1.07±0.55 e-18 cm-2

## Slide 16: Comparison to the theory
Prediction
Experiment: 1.07±0.55 e-18
1.59e-18 cm-2 for 5keV electron beam

## Slide 17: Preliminary results of Neon gas jet
5 mm
y
x
Electron energy: 3 keV
Beam current: 0.50mA
Integration time: 200 s for N2, 4000 s for Neon
Inlet pressure: 5bar
Image from the phosphor screen

## Slide 18: Comparison between N2 and Neon
Gas jet condition: inlet 5 bar
Photon number per second
  Nitrogen: 5.63
  Neon: 0.77
Ratio of Nitrogen/Neon: 7.32

## Slide 19: Possible reason
 | Neon | Nitrogen |  | Ratio of Nitrogen and Neon |  | 
 | 3keV | 3keV |  | 3keV |  | 
Degree of freedom | 3 | 5 |  |  |  | 
Heat capacity ratio | 1.4 | 1.67 |  |  |  | 
Atomic mass | 20 | 28 |  |  |  | 
Pressure in the interaction from theory | 3.29E+17 | 2.62E+16 |  | 7.96E-02 |  | 
(inlet 5 bar) |  |  |  |  |  | 
Cross section | 4.60E-20 | 2.37E-18 |  | 5.15E+01 |  | 
Eff_PC | 0.09 | 0.19 |  | 2.111111 |  | 
T_filter | 9.00E-01 | 4.50E-01 |  | 5.00E-01 |  | 
Equal density |  |  |  | 5.43E+01 |  | 
Take density into consideration |  |  |  | 4.32E+00 |  | 
Notes: Add curve of gas jet density for Nitrogen and Neon.

## Slide 20: Vacuum test of modified system (lower pumping speed)

## Slide 21: Pressure data when e-gun is off
Chart: Pressure in the interaction chamber vs inlet pressure
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
 | 1.95e-09 | 1.97e-09 | 2e-09 | 2.02e-09 | 2.04e-09 | 2.05e-09 | 2.08e-09 | 2.1e-09 | 2.12e-09 | 2.13e-09
Chart: Pressure in the dump chamber vs inlet pressure
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
 | 3.17e-10 | 3.34e-10 | 3.44e-10 | 3.54e-10 | 3.63e-10 | 3.7e-10 | 3.78e-10 | 3.83e-10 | 3.87e-10 | 3.9e-10
Detailed data for each chamber will be loaded to indico

## Slide 22: Pressure data when e-gun is on
Chart: Pressure in the interaction chamber vs inlet pressure
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
 | 1.4e-08 | 1.4e-08 | 1.4e-08 | 1.38e-08 | 1.38e-08 | 1.39e-08 | 1.36e-08 | 1.37e-08 | 1.38e-08 | 1.39e-08
Chart: Pressure in the dump chamber vs inlet pressure
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
 | 4.39e-09 | 4.41e-09 | 4.38e-09 | 4.3e-09 | 4.3e-09 | 4.35e-09 | 4.27e-09 | 4.36e-09 | 4.38e-09 | 4.38e-09
Detailed data for each chamber will be loaded to indico

## Slide 23: Vacuum test of modified system (lower pumping speed and back flow blocker)

## Slide 24: Pressure data when e-gun is off
Chart: Pressure in the interaction chamber vs inlet pressure
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
 | 1.95e-09 | 1.97e-09 | 2e-09 | 2.02e-09 | 2.04e-09 | 2.05e-09 | 2.08e-09 | 2.1e-09 | 2.12e-09 | 2.13e-09
Chart: Pressure in the interaction chamber vs inlet pressure
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
 | 3.17e-10 | 3.34e-10 | 3.44e-10 | 3.54e-10 | 3.63e-10 | 3.7e-10 | 3.78e-10 | 3.83e-10 | 3.87e-10 | 3.9e-10
Detailed data for each chamber will be loaded to indico

## Slide 25: Pressure data when e-gun is on
Chart: Pressure in the interaction chamber vs inlet pressure
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
 | 2.53e-08 | 2.53e-08 | 2.53e-08 | 2.53e-08 | 2.53e-08 | 2.53e-08 | 2.52e-08 | 2.53e-08 | 2.53e-08 | 2.53e-08
Chart: Pressure in the interaction chamber vs inlet pressure
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
 | 3.24e-09 | 3.24e-09 | 3.25e-09 | 3.25e-09 | 3.25e-09 | 3.26e-09 | 3.26e-09 | 3.27e-09 | 3.28e-09 | 3.28e-09
Detailed data for each chamber will be loaded to indico

## Slide 26: Summary of highlight
A prototype supersonic gas jet monitor based on BIF mode was designed, built and successfully commissioned;
N2 gas jet was carefully measured and has been successfully tested as a working gas using laboratory electron beam source;
Neon gas jet has been proven as working gas.
Vacuum test of a modified System in order to design a LHC compatible gas jet system.

## Slide 27: Future work
Continue to optimize the design and geometry
  Check again the alignment
  E.g. new De Laval nozzle
  Change geometry of skimmers.
Characterize Neon gas jet
Argon used as a working gas
Design and building of v3 gas jet system (LHC compatible)
  Final deliverable for the HL-LHC-UK

## Slide 28: Thank you

## Slide 29: Optimize the gas jet parameter
Inlet pressure
Geometry (Nozzle to skimmer1 distance)
Chart: Inlet Pressure (bar)
| 0.5 | 1 | 2 | 3 | 4 | 5 | 6 | 8
1.78mm | 26930 | 26550 | 20980 | 15260 | 8590 | 5420 | 4229 | 3176
2.78mm | 30300 | 31290 | 32700 | 31450 | 27230 | 22330 | 17220 | 9744
3.78mm | 33692.2743 | 35529.7542 | 43100.238 | 51126.0165 | 57582.8769
4.28mm | 30800 | 33630 | 38500 | 40840 | 42170
6.48mm | 26660 | 28180 | 30230 | 29970 | 30740 | 31140
8.48mm | 24490 | 24630 | 24130 | 22810 | 22420 | 21620
Nozzle to skimmer1 distance

## Slide 30: Skimmer to Nozzle Distance 6.06mm (No 3rd skimmer)
Inlet = 1bar
Inlet = 2bar
Inlet = 3bar
Photon Number Vs Pixel
Pixels

## Slide 31
DN100 gasket with 63 hole
DN160 to DN100 adaptor (200 mm) + DN100 gasket with 63 hole + Hipace300
DN160 to DN100 zero length adaptor + DN100 gasket with 63 hole + Hipace300
Normal DN160 gasket and DN160 to DN100 adaptor (200 mm)
DN100 gasket with 63 mm hole +DN100 to DN160 adaptor (100 mm) + HiPACE700
HiPACE700 here with standard gasket

## Slide 32
DN100 gasket with 63 hole
DN160 to DN100 adaptor (200 mm) + DN100 gasket with 63 hole + Hipace300
DN160 to DN100 zero length adaptor + DN100 gasket with 63 hole + Hipace300
DN160 gasket with 32 mm hole (backflow blocker) + DN160 to DN100 adptor (200 mm)
DN100 gasket with 63 mm hole +DN100 to DN160 adaptor (100 mm) + HiPACE700
HiPACE700 here with standard gasket

## Slide 33
14
2
4
6
8
10
12
1
2
3
4
5
6
7
8
0
1
2
3
4
5
6
7
x 10
8.3 mm
1.1 mm

## Slide 34
0
2
4
6
8
10
12
x 10
14
6.0 mm
0.72 mm

## Slide 35: Plans for 2018
Install new gun in the old setup to further decrease the integration time (in seconds or less)
Scanning gauge working in continuous jet mode
Finish commissioning of the second gas jet setup
  Nozzle and skimmers alignment
  Chamber blackening
  Pumping test
  Bake-out
Experiment tasks for the second gas jet
  Jet image of e-beam, integration time and resolution
  Different gas species, Nitrogen, Neon, Argon
  Nozzle sizes (20um, 30um, 50um) and shape (regular or naval nozzle)
  Jet density measurement
  Nozzle, skimmer distance

## Slide 36: Alignment of the nozzle skimmers assembly
1. Only nozzle is mounted  marker on nozzle position
2. 1st 2nd skimmer mounted
3. Adjust nozzle precisely, according to preparation
4. Mount 2nd skimmer and adjust it precisely, according to the preparation
Notes: The alignment is performed off-line using a imaging system.
The Imaging resolution about 3 um. The alignment is based on the position of the 1st skimmer, Images of Nozzle and 2nd skimmer is taken accordingly and adjust to the laser position.

## Slide 37: Camera system for alignment
Magnification (calibration): 1.35 um/ pixel
Resolution: Group 7 set3: 161.3 lines pair/mm