## Slide 1: Outcome of the bench impedance measurements on the LHC CFC collimator
F. Giordano, N. Biancacci, L. Teofili
Many thanks to: P. Arpaia, S. Antipov, I. Lamas Garcia, B. Salvant and TE-VSC collegues.

## Slide 2: Outline
Brief introduction with results
Measurement setup
Data post-processing
Comparison with IW2D
Quick look at Mo on MoGr

## Slide 3: Why have we measured it?
The CFC is a very anisotropic material [1], some doubt was raised whether it has been cut in the proper direction.
The measurements were needed to verify that the resistivity along the direction of the beam of the CFC in the LHC collimator is as expected ().
Our measurements show that it is cut in the proper direction.
[1]:MoGr and CFC: influence of the orientation on the electrical conductivity. F. Carra, A. Bertarelli, J. Guardia
125th HiLumi WP2 Meeting

## Slide 4: Measurement setup: Resonant method [2]
A wire has been installed inside a spare TCS. It has been welded on one side and linked to a weight (4kg) on the other side in order to minimize the sag of the wire.
Two capacitors have been added at the extremity of the wire to create a resonant structure.
A VNA has been used to measure the S11, S12 and S22 inside the TCS.
The wire has been cleaned with ethanol and handled with gloves in order to not
contaminate the collimator.
Slow and careful installation without touching the jaw of the TCS.
The value of the capacity has been empirically chosen in order to better observe the resonance in the structure
Check wich TCS is
[2]:Longitudinal and Transverse Wire Measurements for the Evaluation of Impedance Reduction Measures on the MKE Extraction Kickers. T. Kroyer, F. Caspers, E. Gaxiola

## Slide 5: Measurement setup
TCS
Capacitor
Capacitor
VNA
4kg weight
4kg

## Slide 6: Measurement setup
The jaws of the collimator have been moved with micrometric precision by a remote controller

## Slide 7: Measured data
Given L the length of the structure a resonance occurs each time we have a wavelength such that :
With n integer.
The Q factor of each resonance peak has been computed
Span from 0 to 1.8GHz with a step of 180 MHz

## Slide 8: Data post-processing
Given the factor at each resonance the has been computed from:
[2]:Longitudinal and Transverse Wire Measurements for the Evaluation of Impedance Reduction Measures on the MKE Extraction Kickers. T. Kroyer, F. Caspers, E. Gaxiola
With:
That can be derived from the log formula at the resonance positions [2] from:
and:
Only for flat beam chamber.
Due to the presence of resonant mode of the structure itself some values were corrupted.

## Slide 9: Data post-processing
Good scaling: 1/gap
Bad scaling
To check which value were good and which not the scaling of the impedance with the gap has been analysed:
Half gap [mm]
Half gap [mm]

## Slide 10: Comparison with IW2D
Considering only the non corrupted measured values and comparing with IW2D the square root fit of the measurements gives:
In agreement with of the impedance model.

## Slide 11: Eddy current measurement on naked CFC jaw:
side
top
front
N. Biancacci
The resistivity of the CFC has been measured on the top, front and side of a naked jaw.

## Slide 12
side
top
front
N. Biancacci
The resistivity along the top side is around 4 in a reasonably good agreement with the resonant wire method and the impedance model.
Eddy current measurement on naked CFC jaw:

## Slide 13: Can we apply the same method to measure the MoMoGr?
Depending on the resistivity of the Mo on MoGr the method has to be applied within the accuracy limit.
Log scale
Linear scale

## Slide 14: Conclusions
Both the resonant wire method and the eddy current one show that the CFC installed in the TCS is cut in the proper direction.
The Mo on MoGr has a factor 10 of improvement w.r.t. the CFC.
In order to get a good measurement of the Mo on MoGr resistivity the resonant method can be applied within the accuracy limit.

## Slide 15: Thank you for the attention!

## Slide 16: Back up slides

## Slide 17: Mo on MoGr

## Slide 18: Q_no_wire vs Q_wire