## Slide 1: Magnetic measurements at cryogenic temperature on MCBRDP1 (D2 corrector prototype)
Lucio Fiscarelli
WP3 meeting, 21/11/2018

## Slide 2: Up to now…
Available measurements
First aperture of the short model at room temperature
16 Jan 2018
Both apertures of the short model at cryogenic temperature
27 Jun 2018 https://indico.cern.ch/event/738022/
Both apertures of the prototype at room temperature
12 Sep 2018 https://indico.cern.ch/event/753441/
One aperture of the prototype at cryogenic temperature
this presentation

## Slide 3: Measurement setup
Reference radius
35 mm
Rotating-coil in the helium bath (vertical cryostat setup in SM18)
Only one shaft was available, one aperture was equipped
  Length = 2.1 m
  5 segments
  Not sufficient for covering the full integral field, we will focus on the central field (3 segments)
The equipped aperture is the one with field in the vertical direction (AP1 in the following)

## Slide 4: Measurement list
3 cycles with valid data:
  Stair-step cycle up to ±ultimate in AP1
  0 A in AP2
  Stair-step cycle up to ±400 A in AP1
  400 A in AP2
  Stair-step cycle up to ±400 A in AP1
  -400 A in AP2
  All cycles were preceded by a pre-cycle

## Slide 5: AP1 up to ±ultimate and AP2 at 0 A
Specified value 6.02 T kA-1

## Slide 6: AP1 up to ±400 and AP2 at 400 A

## Slide 7: Field quality table
 | AP1 at ambient temperature |  |  |  |  | AP1 at 1.9 K and 430 A |  |  |  |  | 
 | NCS | Center | CS | INTEGRAL |  | pos1 | pos2 | pos3 |  |  | 
n | bn |  |  |  |  |  |  |  |  |  | 
2 | 1.37 | -0.68 | -0.06 | -0.18 |  | -2.19 | -3.80 | -5.12 |  |  | units
3 | -10.13 | -9.79 | -11.18 | -10.11 |  | -10.93 | -11.18 | -11.19 |  |  | 
4 | 0.19 | 0.07 | 0.07 | 0.09 |  | -0.17 | -0.13 | -0.13 |  |  | 
5 | 2.57 | 2.26 | 2.10 | 2.29 |  | -0.01 | 0.09 | 0.19 |  |  | 
6 | -0.04 | 0.02 | -0.13 | -0.02 |  | -0.06 | 0.02 | -0.06 |  |  | 
7 | -0.96 | -0.98 | -1.20 | -1.02 |  | -0.38 | -0.31 | -0.33 |  |  | 
8 | 0.07 | 0.01 | -0.02 | 0.02 |  | -0.02 | 0.03 | -0.02 |  |  | 
9 | 0.31 | 0.13 | 0.42 | 0.22 |  | -0.02 | 0.00 | 0.01 |  |  | 
10 | 0.00 | 0.01 | 0.19 | 0.04 |  | 0.00 | 0.01 | 0.00 |  |  | 
n | an |  |  |  |  |  |  |  |  |  | 
2 | 13.07 | -6.34 | -31.48 | -7.41 |  | -4.80 | -3.23 | -4.78 |  |  | units
3 | -4.41 | -0.59 | 2.15 | -0.79 |  | -0.18 | -0.35 | -0.87 |  |  | 
4 | -1.33 | 0.06 | 1.25 | 0.02 |  | -0.35 | -0.26 | -0.28 |  |  | 
5 | -0.28 | 0.25 | 0.33 | 0.17 |  | 0.17 | 0.17 | 0.22 |  |  | 
6 | -0.53 | -0.26 | -0.28 | -0.31 |  | -0.06 | -0.06 | -0.04 |  |  | 
7 | -0.08 | 0.00 | -0.01 | -0.02 |  | 0.04 | -0.01 | 0.01 |  |  | 
8 | 0.06 | 0.01 | -0.24 | -0.02 |  | 0.00 | 0.01 | -0.02 |  |  | 
9 | -0.06 | -0.03 | -0.15 | -0.06 |  | -0.03 | -0.03 | -0.04 |  |  | 
10 | -0.11 | -0.02 | 0.41 | 0.04 |  | 0.00 | 0.01 | 0.00 |  |  | 
b3 in agreement with results at ambient temperature

## Slide 8: Cross talk
AP1 | +400 A |  |  |  |  |  |  |  | 
AP2 | 0 A | +400 A | -400 A |  | 0 A | +400 A | -400 A |  | 
n | bn |  |  |  | an |  |  |  | 
2 | -3.84 | 1.31 | 0.98 |  | -3.39 | -10.61 | 3.96 |  | units
3 | -11.27 | -16.13 | -15.96 |  | -0.33 | 2.12 | -3.26 |  | 
4 | -0.14 | 0.42 | 0.30 |  | -0.27 | -1.30 | 0.87 |  | 
5 | 0.23 | -0.26 | -0.20 |  | 0.18 | 0.63 | -0.30 |  | 
6 | 0.01 | 0.01 | -0.02 |  | -0.05 | -0.21 | 0.12 |  | 
7 | -0.35 | -0.39 | -0.38 |  | -0.01 | 0.05 | -0.08 |  | 
8 | 0.03 | 0.00 | -0.01 |  | 0.01 | -0.01 | 0.03 |  | 
9 | 0.01 | 0.01 | 0.01 |  | -0.03 | -0.02 | -0.04 |  | 
10 | 0.01 | 0.01 | 0.01 |  | 0.01 | 0.00 | 0.01 |  | 
Cross-talking effects are noticeable in contrast with results from the short model.

## Slide 9: Conclusions
From tests at cryogenic temperature (1.9 K)
  Transfer function larger than specifications (table WP3)
  In agreement with results at ambient temperature
  b3 in the straight section is -10 units
  In agreement with measurements at ambient temperature
  Cross-talking effects are noticeable
  They were not visible on the short model
In occasion of the thermal cycle, we are equipping the magnet with two new shafts with suitable length in order to perform additional magnetic measurements

## Slide 10: Additional slides

## Slide 11: MCBRDP1 at room temperature 1
Field repeatable between the two apertures within 2 units
Perpendicularity in the order of 1.5 mrad
Measured TF larger than calculations,
a verification (current generator) is ongoing
 | AP1 |  |  |  |  | AP2 |  |  |  |  | 
 | NCS | Center | CS | INTEGRAL |  | NCS | Center | CS | INTEGRAL |  | 
 | 18.7% | 62.7% | 18.7% | 100.0% |  | 18.7% | 62.6% | 18.7% | 100.0% |  | 
I | 0.4467 | 0.4464 | 0.4466 |  |  | 0.4465 | 0.446 | 0.4463 |  |  | A
B1 | 0.877 | 2.944 | 0.877 |  |  | -0.876 | -2.942 | -0.878 |  |  | mT
TF |  |  |  | 12.672 |  |  |  |  | -12.675 |  | mT m A-1
angle | 0.2045 | 0.2105 | 0.2137 | 0.2100 |  | -1.3642 | -1.359 | -1.3559 | -1.3594 |  | rad

## Slide 12: MCBRDP1 at room temperature 2
Allowed multipoles are larger than expected but field quality is within specifications
 | AP1 |  |  |  |  | AP2 |  |  |  |  | 
 | NCS | Center | CS | INTEGRAL |  | NCS | Center | CS | INTEGRAL |  | 
n | bn |  |  |  |  |  |  |  |  |  | 
2 | 1.37 | -0.68 | -0.06 | -0.18 |  | 2.36 | -1.09 | 1.04 | -0.05 |  | units
3 | -10.13 | -9.79 | -11.18 | -10.11 |  | -9.50 | -9.20 | -8.75 | -9.17 |  | 
4 | 0.19 | 0.07 | 0.07 | 0.09 |  | 0.49 | 0.76 | 0.85 | 0.73 |  | 
5 | 2.57 | 2.26 | 2.10 | 2.29 |  | 2.91 | 2.72 | 2.51 | 2.71 |  | 
6 | -0.04 | 0.02 | -0.13 | -0.02 |  | -0.23 | -0.04 | -0.11 | -0.09 |  | 
7 | -0.96 | -0.98 | -1.20 | -1.02 |  | -0.87 | -0.81 | -0.82 | -0.82 |  | 
8 | 0.07 | 0.01 | -0.02 | 0.02 |  | -0.08 | 0.06 | 0.04 | 0.03 |  | 
9 | 0.31 | 0.13 | 0.42 | 0.22 |  | 0.17 | 0.15 | 0.13 | 0.15 |  | 
10 | 0.00 | 0.01 | 0.19 | 0.04 |  | -0.09 | -0.02 | -0.04 | -0.03 |  | 
n | an |  |  |  |  |  |  |  |  |  | 
2 | 13.07 | -6.34 | -31.48 | -7.41 |  | -22.67 | 0.77 | 25.34 | 0.99 |  | units
3 | -4.41 | -0.59 | 2.15 | -0.79 |  | -3.81 | 0.09 | 2.36 | -0.21 |  | 
4 | -1.33 | 0.06 | 1.25 | 0.02 |  | 1.40 | 0.13 | -0.61 | 0.23 |  | 
5 | -0.28 | 0.25 | 0.33 | 0.17 |  | 0.04 | 0.34 | 0.17 | 0.25 |  | 
6 | -0.53 | -0.26 | -0.28 | -0.31 |  | 0.34 | 0.17 | -0.06 | 0.16 |  | 
7 | -0.08 | 0.00 | -0.01 | -0.02 |  | 0.11 | 0.05 | 0.09 | 0.07 |  | 
8 | 0.06 | 0.01 | -0.24 | -0.02 |  | 0.21 | 0.03 | 0.05 | 0.07 |  | 
9 | -0.06 | -0.03 | -0.15 | -0.06 |  | 0.04 | -0.02 | -0.06 | -0.01 |  | 
10 | -0.11 | -0.02 | 0.41 | 0.04 |  | 0.11 | 0.01 | 0.05 | 0.04 |  | 

## Slide 13: MCBRDP1 at room temperature 3
AP1
AP2

## Slide 14: MCBRDS1b at cryogenic temperature
AP1 vertical field newest-built aperture |  |  | 
TF | 1.695 |  | T m kA-1
n | bn | an | 
2 | -4.1 | -2.97 | units at 35 mm
3 | -3.29 | -0.74 | 
4 | -0.59 | -0.16 | 
5 | -1.09 | -0.08 | 
6 | -0.05 | -0.16 | 
7 | -0.09 | -0.14 | 
8 | 0.01 | -0.05 | 
9 | -0.06 | -0.03 | 
10 | -0.03 | -0.02 | 
at nominal field level and 1.9 K |  |  | 
AP2 horizontal field first-built aperture |  |  | 
TF | 1.698 |  | T m kA-1
n | bn | an | 
2 | -2.05 | 1.7 | units at 35 mm
3 | -2.69 | -0.01 | 
4 | 0.35 | 1.17 | 
5 | -1.77 | 0.23 | 
6 | 0.54 | 0.23 | 
7 | 0.28 | 0.09 | 
8 | 0.28 | 0.05 | 
9 | 0.01 | 0.05 | 
10 | 0.07 | 0.01 | 
at nominal field level and 1.9 K |  |  | 

## Slide 15: AP1 newest built (↕) Above nom.
TF = 1.696 Tm kA-1
MCBRDS1b at cryogenic temperature

## Slide 16: AP2 first built (↔) Above nom.
TF = 1.699 Tm kA-1
MCBRDS1b at cryogenic temperature

## Slide 17: MCBRDS1b at cryogenic temperature AP1 newest built (↕) X-talk
TF = 1.696 Tm kA-1

## Slide 18: MCBRDS1b at cryogenic temperature AP2 first built (↔) X-talk
TF = 1.699 Tm kA-1

## Slide 19: MCBRDS1b at room temperature
First-built aperture at room temperature 1 A |  | 
n | bn | an
2 | 3.34 | 3.84
3 | -7.86 | 2.24
4 | -1.07 | 1.53
5 | -0.84 | -0.12
6 | -0.41 | 0.45
7 | 0.34 | -0.16
8 | -0.29 | 0.14
9 | 0.09 | -0.07
10 | -0.03 | 0.17