## Slide 1: Known instrumentation needs for NA + EA
Lau Gatignon

## Slide 2: Introduction
As stated in the introduction to the mandate, there are limitations in the instrumentation for the North and East Areas. They are due to
Aging of the equipment
Obsolescence of equipment and technology used
Loss of expertise, lack of spares and their availability on the market
Lack of correspondence between performance and needs, e.g. intensity range
Tighter safety and RP requirements, decrease of allowed dose and hence need for better control of losses and activation
Changing requirements from users, e.g. spill quality
And so forth

## Slide 3: North Area – primary beams
Several problems are outstanding in LSS2, TT20, TDC2, TCC2:
Losses at ZS are too high, therefore long cooldown times before repair
Losses along TT20 – need BLM
Uncertainties about intensities due to calibration and aging of SEM foils. Difficult to estimate the transmission
Steering through splitters problematic, hence losses and high dose above H6, as well as tails on T4 profile (therefore lower transmission to T10)
Problems around steering to targets (TBIU positioning)
Absence of instrumentation around T2 ion splitter
More issues from your side?
Synergy with SLAWG working group
Notes: Long

## Slide 4: A recent example
SLAWG et al have made tremendous progress with loss reduction at the ZS extraction septa, using a crystal diffuser.
However, there were some difficulties with steering through TT20, mostly based on split foils
Beamlet has small intensity but sometimes large offset, therefore inadequate input to steering program
Can different type of instrumentation help? (e.g. SEM grids)
Main beam
Deflected beamlet
Split foil

## Slide 5: Other examples
Modern high-rate experiments tend to be very sensitive to spill structure, e.g. NA62 and COMPASS
This includes 50 Hz plus higher harmonics, 43 kHz and even 200 MHz
The present spill monitor cannot easily handle the higher frequencies
The effective duty spill seen by OP is ~4.5 s (for 4.8 s flat top), whereas the experiments report rather spills of order 3 seconds
Most experiments discard the beginning of the spill (0.2 to 0.5 s) to avoid RF-induced pile-up
A fast feedback on effective spill on page-1 would be appreciated by the users and helpful for the operation team.

## Slide 6: North Area – secondary beams
Most of the North Area instrumentation dates from the late 1970’s, although the electronics was replaced in the beginning of the century.
Intensity monitoring is largely based on scintillators, but at intensities above a MHz, the situation is more difficult, e.g. the calibration of XIONs and SEMS can be problematic and dependent on the beam type.
Profile monitoring is in a dramatic situation, as the most used intensity range is not covered adequately. Also the equipment is obsolete and we have lost the tooling and sometimes the expertise to produce spares
Particle identification relies on very old and delicate devices without spares and with disappearing expertise.
Pure electron beams require absence of material on the beam. Can we reduce vacuum interruptions?
We have no instrumentation for wide beams or beam halo
Any more from your side?

## Slide 7: Beam intensity
XSCINT are basically ok, but consolidation required?
Also need such signals as strobe for more complicated detectors (Cerenkovs, XEMC, XDWC, etc)
XION calibration depends on material upstream and needs calibration, e.g. factor 7400 in M2 muon beam but 5300 for M2 hadron beam. Why factor 10,000 in K12 beam?
In P0 only intensity measurements at T4 and very close to T10
SEM response degrades by up to 20 % in beam impact zone. Would need frequent, therefore easier calibration
…..

## Slide 8: Beam profiles
XDWC work up to few 104/spill and are often in very degraded state
MWPC work from 105/spill upward and have tails on the profile. Some suffer seriously from radiation damage (e.g. M2)
FISCs are in general ok, but each profile takes a long time. Fast FISC mode is nice for very narrow beams, but is subject to spill shape for wide beams (normalisation does not work so well)
For wide beams at low intensity we have prepared XBPF detectors. So far only for the EHN1 extension, but they could be useful elsewhere
One proposal was to replace all XDWC and XWCA by (small) XBPFs. An in/out movement would even allow to use them instead of FISCs.
FISC coincidence functionality remains important for CEDAR tuning
No profiles are available at all in P0, other than at T4 and T10 (via OP). Collimator scans are not realistic for full beam (> 2 1012 ppp)

## Slide 9: Particle identification
CEDARS are very nice devices and vital for many experiments, but:
  No spares, in particular for CEDAR-W (NA61, NA62)
  No expertise and spare parts
  Thermal housing has degraded severely
  Gas system obsolete
  Rate capability no longer sufficient
  Need for other gases (different optical components)?
Thresholds are old.
  Mirrors in bad shape, no spares
  Gas system obsolete
  Rate capability no longer sufficient in some cases
XEMC reliability not perfect (often blocked in beam)
Not all energies covered for all particles, E.g. add TOF detectors?

## Slide 10: East Area
The East Area goes through a complete renovation program with completely new beam lines. The new East Area will be more similar to the North Area in terms of beam instrumentation. Therefore most of the arguments cited before also apply here.
For the primary beam instrumentation beam loss monitoring must be defined in detail. Also profile and intensity monitoring of the primary beams would profit from developments in the North Area.
In the secondary beams, XBPF (of 20x20 cm2?) would be a good solution for the profiles. Time of flight is the most realistic tool for hadron identification at the lower energies, whereas threshold counters could cover the higher energy range. The present XDWC are severely degraded.
No profile monitoring exists for T11 (CLOUD).
In the IRRAD and CHARM facility the intensity measurements are difficult, in particular for ion beams.

## Slide 11: Final remarks
These slides are not intended at all to give the full picture, but just to give a few examples of where we are and where we have to go.
The first step is now to complete the picture, following the proposed timeline.
A stated elsewhere, there is synergy with other activities, such as with SLAWG and with the different consolidation projects.
SLAWG has already brought up many ideas which have been discussed with BI and other groups involved.
EACONS work is starting now on the floor and instrumentation is necessary for the start-up in 2021. The same is true for ELENA.
ADCONS is under way too and instrumentation issues have been discussed in AD since a long time.
For the NA a major consolidation project is under discussion and it becomes important and urgent to define/refine scope, costs and resources for instrumentation now.