## Slide 1: AD beam instrumentation

## Slide 2: Systems in use
BPM (Chromox) in front of AD target
GEM
BTV
BCCCA
Scraper ?
New Orbit measurement system: Acquisition system based on VXS card with FMC-ADC, developed BE-RF
SEM’s
LPU (schottki)
Capacitive PU’s (BPM)

## Slide 3: Known issues
BCCCA :
Measurement only possible at end of cycle (due flux-jump at injection)
  Study improvement in slew-rate limitation
  Modify coupling circuit in SQUID cartridge / add EM filter inside cryostat (to be studied)
  Integration with Vistar
Superconducting connection in DC coupling circuit kept only for limited range
  Currently compensated by imposing a constant DC current
  Replace SQUID cartridge with improved niobium wire connections (to be studied)
Cryogenic availability
  LHe circuit becomes contaminated blocking the He gas flow in the closed-circuit
  He diffusion through ceramic breaks degrades insulation vacuum
GEM + BTV: destructive
SEM: technology not fully transferred to CERN
… OP and BI to contribute !