## Slide 1: LEMs for ProtoDUNE-DP (NP02)
Edoardo Mazzucato CEA/Irfu/DPhP

## Slide 2: ProtoDUNE-DP : a 300t LArTPC
Charge Readout
Planes (CRP)
Field Cage
Cathode
PMTs
Ground Grid
6m
6m
6m
Dual phase principle
3m
3m
1 CRP = 36 Anodes + 36 LEMs (50×50 cm2)
+ 3×3 m2 Extraction Grid

## Slide 3: Dual phase concept validation @ 4-ton scale
Bldg. 182
@ CERN
3 m3 TPC
The 311 prototype
Prototype operated from June until November 2017
Proportional scintillation (S2) observed
Extraction of electrons over 3 m2 area achieved
Amplification through LEMs measured
But … HV stability issues with Grid and LEMs encountered
Through-going muon
EM shower
HAD interaction
(B. Aimard et al., JINST 13 (2018) 11, P11003)

## Slide 4: LEM Design for ProtoDUNE-DP
CFR-34 – 311 prototype
CFR-35 – NP02
5 + 10 mm
clearance
2 + 2 mm
clearance
Not necessarily the final design for a DUNE 10 kt detector!
 86% active area
 96% active area

## Slide 5
LEM production
Cleaning + drying
+ polymerization
Soldering HV pins +
glueing MACOR insulation
Ultrasonic bath
LEM Survey
LEM manufacturing by ELTOS (Italy) – requested specifications met (FR4 and copper thicknesses, hole, rim and LEM size)
Characterization and tests @ CEA/Irfu
74 LEMs produced and validated for 2 CRPs
(Jan. – Oct. 2018)

## Slide 6
LEM HV Tests in a High Pressure Chamber
Amplification gain inside the LEM depends on gas density  or P/T :
Geff = Ced with  = Ae-B/E
GAr @ 88°K and 1 bar  T = 20°C and P 3.3bar.
Stack of 6 LEMs
LEM + ANODE
CATHODE
collimated
 source

## Slide 7: LEM HV Tests in Ar @ 3.3 bar
VLEM = 3.5 kV in Ar @ 3.3 bar
- gain > 100 in DLAr (before charging up)
- no PS trip for > 40h
- spark rate : 3/h
CFR-35 – NP02
CFR-34 – 311
- spark rate : 20/h
(> 45% of sparks near edges or corners)
multiple
sparks
VLEM = 3.3-3.5 kV in Ar @ 3.3 bar

## Slide 8: CRP assembly @ CERN (I)
Assembling the 3x3 m2 G10 frame and coupling to the INVAR structure
Mounting the anodes and the LEMs on the G10 frame
Weaving 30 Grid modules and installation on the CRP
Cabling and instrumentation (T probes, level meters, …)
Survey and planarity adjustments
Transport
HV cabling
Anodes
G10 frame
INVAR structure
CRP
Grid module
Combs
Grid wires
36 anodes + 36 LEMs
1 cm spacing
between LEMs
and Grid

## Slide 9: CRP assembly @ CERN (II)
Fully instrumented CRP
Fake CRP
Ground plates
ProtoDUNE-DP will have 2 fully instrumented CRPs (CRP1 & 2) and 2 fake CRPs (CRP3 & 4)
with an Extraction Grid and ground plates to close the E-field
Transport box

## Slide 10: CRP Tests in Cold Box
Perform electrical and mechanical tests of CRPs in nominal thermodynamic dual phase conditions:
  - HV tests of LEMs and Extraction Grid
  - CRP planarity
  - Tensioning of the extraction grid wires
- Ar purity < 100 ppm O2eq ; GAr @ Patm
- LAr evaporation less than 1 mm/h (< 700 W heat input)
- No boiling, no bubbles, LAr level very quiet
- LAr level controlled @ +/- 0.125 mm with constant refilling
All four CRPs tested (07/2018 - 01/2019)
LAr level  5mm above Grid
Insertion of a CRP in
the Cold Box

## Slide 11: CRP1 and CRP2 LEM HV Test Results (Oct. – Nov./2018)
Liquid level in CB stable to within 250µm; TLEM  91°K; VLEM-GRID = 3kV
Effective gain > 20 before charging up achieved. Best HV configuration with VTOP  0.5kV. For larger VTOP values, need to decrease ELEM to reach stable operation.
* PS TRIP time set too short
VTOP (kV) | VBOT (kV) | ELEM (kV/cm) | Time (h) | Spark Rate (h-1) | Patm (mbar) | Estimated Geff (no ch. up)
0.25 | 3.35 | 31.0 | 12 | 1.3 | 968 - 972 | 20
0.50 | 3.55-3.60 | 30.5-31.0 | 13 | 1.3 | 962 - 966 | 24 - 31
0.75 | 3.70 | 29.5 | 42 | 0.6 | 943 - 953 | 20
1.00 | 3.80 | 28.0 | 18 | 2 trips* | 970 - 976 | 9
1.00 | 3.85 | 28.5 | 12 | 3 trips | 936 - 947 | 15
VTOP (kV) | VBOT (kV) | ELEM (kV/cm) | Time (h) | Spark Rate (h-1) | Patm (mbar) | Estimated Geff (no ch. up)
0.10 | 3.15 – 3.20 | 30.5 – 31.0 | 17 | 0.8 | 969 - 973 | 9 - 11
0.25 | 3.34 | 30.9 | 16 | 1.3 | 968 – 970 | 19
0.50 | 3.55 | 30.5 | 11 | 0.9 | 957 – 965 | 24
0.50 | 3.555 | 30.55 | 42 | 0.5 | 962 – 964 | 25
CRP1
CRP2

## Slide 12: Effective gain
VGRID-LEM (V)
Includes probability
for electrons to enter
LEM holes
VLEM_TOP (V)
Geff = extr × GLEM × ind
VLEM (V)
Gushchin et al., Sov. Phys.
JETP 55 (1982) 860-862.
Extraction efficiency
Collection efficiency
C. Cantini et al., JINST 10 (2015) 03, P03017

## Slide 13: CRP1 : VTOP = 0.50kV and VBOT = 3.55-3.60kV
0 trip, 17sparks @ 0.50kV/3.60kV in 13 hrs
ELEM = 31kV/cm
IMON
VMON

## Slide 14: Sparks, trips and … continuous discharges
1 spark
Recovery time < PS TRIP time
2 sparks
Recovery time > PS TRIP time (120s)
No Trip
2 Channels
tripped
Recovery time > TRIP time
Continuously discharging
PS Trip
Can be detected
and stopped
with appropriate
SC monitoring
program
8 minutes

## Slide 15: Inspection of discharging LEMs after tests in Cold Box revealed dark spots in the corners
LEM 16
Points to a weaker regions of the LEM
Issues with a few LEMs (5/72)
Treated with potassium
permanganate @ CERN to
remove carbonized FR4
inside holes.
Copper surface unaffected.
LEMs operational again
LEM 30
LEM 28
LEM 14
After treatment
Risks can be mitigated with
appropriate SC monitoring
Work in progress with CERN
EP-DT-DI

## Slide 16: LEM rim defects?
Rims around holes near the LEM edges/corners appear to be
decentered and containing copper residues
Problem known by Rui de Oliveira (CERN EP-DT-EF)
Due to method used by ELTOS for the micro-etching process
Decentered rims with
copper residues
Bad rims?
Decentered rims with
copper residues
Problem being investigated in collaboration with EP-DT-EF

## Slide 17: ProtoDUNE-DP Schedule
End on detector installation inside cryostat : 11/03
Cryostat closed : 06/05
Purge and cooldown completed : 27/05
Cryostat filled with LAr : 05/07
 Start LAr purification and commissioning next July
LEMs (2019-2020) : improve design for
larger active area and better HV stability
Possibility to run after the LS2 (2021) with a charged particle beam and a complete detector with the final design for DUNE

## Slide 18: Backup Slides

## Slide 19: LEM QA/QC by ELTOS
LEM Specifications
RIM
HOLE
FR4
PCB

## Slide 20: LEM Survey @CEA/Irfu
t1
t2
t1
t2
Cu
Cu
FR4

## Slide 21: LEM HV training in Ar @ 3.3 bar
All LEMs in CRP1 and CRP2 tested up to 3.4 – 3.5 kV with < 1 spark / 20 minutes
LEM A143
LEM A143
Fully automated HV training up to 3.5kV
3 hours @ 3.5 kV
V
I

## Slide 22: CRP1 : VTOP = 1.00kV and VBOT = 3.85kV
3 trips @ 1.0kV/3.85kV
Very unstable conditions.
Coincides with LAr delivery
that caused 2mm rapid
increase of level in CB
Discharges

## Slide 23: CRP2 : VTOP = 0.25kV and VBOT = 3.34kV
0 trip, 20 sparks @ 0.25kV/3.34kV in 16 hrs