## Slide 1: Inventory of present situation
J. Tan, I. Ortega, S. Deschamps, W. Devauchelle, P. Carriere, J. Cenede, G. Tranquille
F. Roncarolo, M. Duraffourg, S. Burger
February 1st 2019
EA-BI-WG

## Slide 2
Courtesy L. Gatignon

## Slide 3: Outline
From F61 to East Area: status by observable
From TT20 to North Area : status by observable
Transfer lines of AD and ELENA : profile monitors
Summary

## Slide 4: F61 line and East Area: Intensity measurement
Instrument | Location | Status | Qty Installed | Spares Monitors + Aqn chain | Remarks
BCGAA (ex. LSD) | F61 | operational | 1 | 0 + HV power supply + PMT | Ionization photons
BCTF | F61 | operational | 1 | 0 + TRIC cards | Very old monitor - Bunched beam
XSEC | F61 – F62 – F63 – T08 | operational | 5 | 1 + 10 scaler boards | SEM
XION | T08 | operational | 1 | 3 + 10 scaler boards | SEM+Ar From 106 to 1010/spill
XTEL | F61 | operational | 2 | 1 + 5 scint. boards | Scintillators
Instrument | Location | Status | Qty Installed | Spares Monitors + Aqn chain | Remarks
XSCI | T09-T10-T11 | operational | 4 | 10 assembled. 5 new paddles. 5 scint. boards | From 1 to 107/spill. In common with NA.
Primary lines after LS2
Secondary lines after LS2

## Slide 5: F61 line and East Area: Profile measurement
Instrument | Location | Status | Qty Installed | Spares Monitors + Aqn chain | Remarks
BTV | F61- F61D-F62 – T08 | operational | 5 | Rad. Hard cameras, push-pull actuator, Al2O3 screens | Upgrade done after LS1 – One BTV w/ bellow in T8
Target TV | F62 – F63 – T08 | operational | 3 | Rad. Hard cameras | 
XWCM | CHARM | operational | 1 | 8 + Aqn boards | Large screen size Common module for XWCA, XWCM and GEM
Instrument | Location | Status | Qty Installed | Spares Monitors + Aqn chain | Remarks
XBPF | T09 - T10 - T11 | In production | 4 | 3 fibre planes + Aqn board, as EHN1 | Upgrade, replace old XDWC
Primary lines after LS2
Secondary lines after LS2

## Slide 6: East Area – secondary: Particle identification
Instrument | Location | Status | Qty Installed | Spares Monitors + Aqn chain | Remarks
XCET | T09 – T10 | Being upgraded by EN | 4 | ?? + 9 boards | EN: Upgrade of monitor + gas valve mechanics BI: Aqn. Boards as NA + gas valve control
Secondary lines after LS2

## Slide 7: N.A. Primary beam diagnostics
Each dot is an instrument or a diagnostics box with more than 1 instrument…
Courtesy F. Roncarolo

## Slide 8: N.A. Primary beam diagnostics: SEM-based monitors
Instrument | Location | Status | Qty | Spares | Remarks
BSI | lines + target lines | operational | 6 + 3 |  | Intensity. Large systematic errors (~ 20%) LS2: partial electronics consolidation (spill sampling)
BSPH-BSMH | lines | operational | 13 + 2 |  | H position, steering
BSPV-BSMV | lines | operational | 8 + 3 |  | V position, steering
BSGH-V | lines | operational | 5 + 13 |  | Profile w/ grids
BBSH-V | lines | operational | 3 + 3 |  | Profile w/ movable strip
BSG | T2 and T4 | operational | 1+1 |  | Dual profile
TT20, TDC2, TCC2
Target boxes
Instrument | Location | Status | Qty | Spares | Remarks
TBIU | T2, T4 and T6 | operational | 3 |  | Renovated in 2014
TBID | T2, T4 and T6 | operational | 3 |  | Renovated in 2014

## Slide 9: N.A. Primary beam diagnostics: BTVs and BLMs
Instrument | Location | Status | Qty | Spares | Remarks
BTV | TT20 | operational | 2 | A few tube-based cameras, screens | LS2:Mech. Consolidation (Malta Cross)
Screen for long. spill | TT20 | operational | 1 | A few tube-based cameras, screens | 
BTV | BA80 | operational | 7 | A few tube-based cameras, 0 screens (40yo) | Very old, yet robust. Information on a couple of screens is lost. LS2: Upgrade to standard BTV device (VME) & application New cables Mech. Consolidation (Malta Cross)
BTV 240119 | Downstream T4 |  |  |  | not used anymore, to be removed during LS2 in agreement with EA
BTV
BLM
Instrument | Location | Status | Qty Installed | Spares | Remarks
BLM | TT20 | operational | 6 |  | SPS type
BLM | TDC2 | Failure in 2018 | ?? |  | LS2: Will install new cables and have old and new in parallel

## Slide 10: N.A. Secondary beam diagnostics
T6
T4
T2
T2 wobbling
T4 wobbling
T10
NA62
NA61
H2
H4
H6
H8
M2
P42
P6
K12
NA64
COMPASS
Courtesy Alexander Gerbershagen

## Slide 11: North Area- secondary: Intensity measurement
Instrument | Location | Status | Qty | Spares | Remarks
XION | M2, K12 | Operational | 3 | 3 , IFC and Scaler modules | Old aqn. chain
XSCINT | H2, H4, H6, H8, P42-62, M2, K12 | Operational | 32 | 10 sets (Scint+PMT) assembled. 5 new paddles. 5 boards | 
XBTF | Neutrino Platform | Operational | 5 | 0 +0 | Produced in 2018 No spares in agreement with users

## Slide 12: North Area – secondary : Profile measurement
Instrument | Location | Status | Qty Installed | Spares Monitors + Aqn chain | Remarks
XWCA | H2, H4, H6, H8 | Operational | 18 | 10 + 35 boards | From 105/spill. Common module for XWCA, XWCM and GEM Some present degraded performance. Unknown status of spare monitors.
XWCM | H4, M2, K12 | Operational | 11 | 8 | Large XWCA (20cm x 20cm). Some present degraded performance. Unknown status of spare monitors.
XDWC | H2, H4, H6, H8 | Operational | 16 | 19 + 6 boards | Up to 104/spill. Unknown status of spare monitors.
XFISC | K12, H2, H4, H6, H8, P41-61, P42-62, M2 | Operational | 56 tanks | 5 + 24 | Mechanics by EN-EA. Spare filaments – OK. Spare PMT – OK.
XBPF | Neutrino Platform | Operational | 10 planes | 0 +0 | Produced in 2018 No spares in agreement with users
XGEM | H6 | Operational | 1 | 5 | removed in 2018, to be replaced by XWCA during LS2

## Slide 13: North Area – secondary : Particle identification
Instrument | Location | Status | Qty Installed | Spares Monitors + Aqn chain | Remarks
XCET | H2, H4, H6, H8, Neutrino Platform | Operational | 10 | ?? + 9 boards + PMT | Mechanics by EN-EA. LS2 : a couple of XCET to be installed in Neutrino Platform (H2)
XCED | H2, H6, M2, K12 | 2 Operational 3 Need Therm. housing + new gas control | 4 | ?? + boards + PMT | Mechanics by EN-EA. Unknown status of spare monitors.
XEMC | H4, H6, H8 | Operational | 3 | 2 + 1 [FISC+PHA] | Lead scintillator calorimeter.

## Slide 14: AD & ELENA : Profile measurement
Instrument | Location | Status | Qty Installed | Spares Monitors + Aqn chain | Remarks
BTV | AD Transfer line 7000 | Operational | 1 | Camera and screen | CCD camera
XGEM | AD Transfer line 7000 and DEM | Operational | 6 | 15 | Destructive – Produced in ~2012 Common module for XWCA, XWCM and GEM
SEM (Hori detector) | AD and ELENA Transfer lines | In production | 8 |  | 95% transparent 3 assembled monitors ready for VSC acceptance tests 5 to be assembled 28 yet to be assembled (missing grids) 1 set of FE boards available only. Missing 43 !!
A total of 44 SEMs to be installed in the AD and ELENA transfer lines

## Slide 15: Summary of the present status
East Area Consolidation is in good shape
  upgrade : XBPF and XCET
  keep the rest
  Complete cable renovation cables
North Area and transfer lines
  40 yo monitors : loss of documentation and expertise
  Splitters and targets areas are very activated.
  Most monitors: obsolete, miss
  TT20 line: Partial consolidation undertaken before and during LS2
Is ELENA (new machine) within the scope of this Working Group ?
Inventory
  Very difficult to make; few written documentation.
  Some monitors have been lent to users and lost track.
  Spare parts dispersed: centralised storage is needed.
  The status of some spares is unknown: control campaign needed.
  Obsolete boards