## Slide 1: ALBA optical fiber BLMs
Angel Olmos
(on behalf of the ALBA and ICCUB/SiUB teams)
DEELS19, ESRF - Grenoble

## Slide 2: MOTIVATION
We suspected to have a Physical Aperture problem in our MPW chamber
Vertical bumps seem to point out towards a chamber misalignment
In order to diagnose, many Bergoz BLMs were placed on the chamber (top/bottom)
Results were contradictory and so we decided to look for other methods
Development of a BLM based on optical fibers, which will provide integrated losses all along the chamber and possibilities of losses location

## Slide 3: PARTS OF THE SYSTEM
Beam lost monitor system based on scintillating optical fibers + SiPM detectors
Possibility to identify the place of the losses
Preliminary tests of remote control done using a skippy-TANGO interface

## Slide 4: PARTS OF THE SYSTEM
4
MUSIC board developed by ICCUB/SiUB team
It takes care of the SiPM reading + digitalization
Provides analog/digital signals to scope (debugging) and data to Raspberry PI

## Slide 5: LAB TESTS
+80V
+5V
SiPM
SiPM
SiPM
SiPM
SCOPE
MUSIC board

## Slide 6: LAB TESTS
Losses location / Delay Distribution
Delay time between the signals from the SiPM at both ends of the fiber, processed with MUSIC ASIC and recorded with Scope.

## Slide 7: LAB TESTS
Histogram summation for all the delay times from the different positions
Enough resolution to distinguish two losses with 20cm of separation in a 2m optical fiber
20cm resolution
Losses location / Delay Distribution

## Slide 8: TESTS WITH BEAM
Setup

## Slide 9: TESTS WITH BEAM
Scope Mode
Debug mode not intended for normal operation, just for calibration
Example of turn by turn losses during a TopUp injection cycle
1 turn
Digitalized pulses width ~ losses

## Slide 10: TESTS WITH BEAM
V threshold [~mV]
Counts #
CH0 CH1 // CH5 CH6
Calibration
SiPM response calibration with beam turned to be too complicated
Difficult to find “good” SiPM thresholds providing same counts for all sensors

## Slide 11: TESTS WITH BEAM
Counting Mode
Example: Losses during normal Beamlines operation
DCCT
CH0 CH1 // CH5 CH6

## Slide 12: TESTS WITH BEAM
Counting Mode
Example: Beam lost and reinjection
DCCT
CH0 CH1 // CH5 CH6

## Slide 13: TESTS WITH BEAM
Counting Mode
Example: Acquired losses during a beam scraping
Jaws inserted
DCCT
CH0 CH1 // CH5 CH6

## Slide 14: TESTS WITH BEAM
Counting Mode
Example: Single Bunch operation
DCCT
CH0 CH1 // CH5 CH6

## Slide 15: TESTS WITH BEAM
Losses Location Mode
Not yet done !!!

## Slide 16: PROBLEMS / To be understood
Too different channels counts even after calibration
3days
Calibration needs to be better understood and better performed

## Slide 17: PROBLEMS / To be understood
Sudden counts increase w/o correlation to other machine parameters
3days
… and it takes days to stabilize back
Is it real?
Seems to be an electronics issue
Orbit/Corrs/Lifetime … checked and no correlation observed

## Slide 18: PROBLEMS / To be understood
Counts fluctuation with time
3days

## Slide 19: PROBLEMS / To be understood
120 storage ring Bergoz BLMs
3days
Counts fluctuation with time
BLM #

## Slide 20: PROBLEMS / To be understood
Comparison wrt Bergoz BLM with higher counts
3days
SiPM BLM
Bergoz BLM
Sometimes they coincide, sometimes they don’t
sometimes they behave opposite (?)
Counts fluctuation with time

## Slide 21: PROBLEMS / To be understood
Comparison wrt Injection Bump
3days
Not clear correlation
SiPM BLM
Kicker01
Kicker03
Kicker04
Counts fluctuation with time
Maybe

## Slide 22: PROBLEMS / To be understood
Spurious counts / glitches
3days

## Slide 23: PROBLEMS / To be understood
Spurious counts / glitches
3days
Comparison wrt closest Bergoz BLM
SiPM BLM
Bergoz BLM
These are real losses
... due to (?)

## Slide 24: SUMMARY
Preliminary results show that it should work as an standard BLM
Extra feature of losses location to be tested
Problems with RaspberryPI hanging  Replaced by NUC-PC seems to work … but it is Windows
Too different channels counts even after calibration
Extremely sensible to fibers placing on the SiPM  company needed to perform proper connectorized setup
Sudden counts increase w/o correlation to other machine parameters
TANGO Device Server not yet fully operational
Many thanks for you attention!!

## Slide 25: EXTRA SLIDES

## Slide 26
SiPM Hamamatsu

## Slide 27
Fiber Kuraray SCSF-78
1mm diameter