## Slide 1: BPM Resolution Studies at PETRA III
Gero Kube
DESY (MDI)
Introduction
Comments on BPM Resolution
„Three BPM“ Correlation & Prinicipal Components Analysis
Resolution Studies at PETRA III
Conclusion
DEELS 2019
ESRF, 3. -5. June 2019

## Slide 2: PETRA III @ DESY
Gero Kube, DESY / MDI
PETRA history
1978 – 1986: e+e- collider (up to 23.3 GeV / beam)
1988 – 2007: pre-accelerator for HERA (p @ 40 GeV, e @12 GeV)
since 2007: dedicated 3rd generation light source, commissioned in 2009 TDR: DESY 2004-035
→ 14 beamlines (15 experimental stations) operating in parallel
from 2014: staged extension project W. Drube et al., 2016 https://doi.org/10.1063/1.4952814
→ up to 12 additional beamlines (presently not all of them in operation)
Max von Laue Hall
Extension Hall North
Paul P. Ewald
Extension Hall East
Ada Yonath
DEELS 2019 Workshop @ ESRF, 4.6.2019

## Slide 3: PETRA IV: Overview
Gero Kube, DESY / MDI
PETRA IV storage ring and pre-accelerators
use of existing accelerator tunnel
→ asymmetric ring structure
additional experimental hall
→ 29 straight ID sections
time line
presently
→ preparation of Conceptual
Design Report
DEELS 2019 Workshop @ ESRF, 4.6.2019

## Slide 4: PETRA IV: Diffraction Limited Light Source
Gero Kube, DESY / MDI
DLS design
courtesy: I. Agapov (DESY)
single bunch / single turn < 20 μm (assuming 0.5 mA in single bunch → 2.5×1010 particles bunch )
closed orbit < 100 nm (rms, 200 mA in 1600 bunches) at 300 Hz BW
BPM resolution requirements
DEELS 2019 Workshop @ ESRF, 4.6.2019

## Slide 5: BPM Resolution
Gero Kube, DESY / MDI
position determination in circular accelerator
depends on
pickup geometry → beam pipe diameter
button size → small correction
x = Kx
(P1+P4) – (P2+P3)
P1+P2+P3+P4
y = Ky
(P1+P2) – (P3+P4)
P1+P2+P3+P4
x
y
beam
: monitor constant
: signal-to-noise ratio
position resolution (small displacements from center)
geometry (button size) → signal strength
infrastructure → cable length & attenuation
read-out electronics
main focus: read-out electronics
goal
performance test → existing Libera Brilliance readout electronics @ PETRA III
DEELS 2019 Workshop @ ESRF, 4.6.2019

## Slide 6: BPM Resolution Measurements
Gero Kube, DESY / MDI
modern ADCs optimized for cw signals
common methods for correlation analysis
„three BPM“ correlation method
Principal Components Analysis (PCA)
BPM resolution with beam generated signals
signal from BPM button → far away from cw signal
BPM signal measurement with beam → 2 kinds of jitter
beam jitter
→ real change of beam angle and position caused by fluctuations in accelerator
(ground motion, energy fluctuation, kicks, …)
→ seen by several / all BPMs simultaneously
(correlation via beam optics)
noise of BPM electronics
→ quantitiy to be measured
→ no correlation between adjacent BPM readings
correlation analysis in order to disentangle both jitter sources
brief review
DEELS 2019 Workshop @ ESRF, 4.6.2019

## Slide 7: „Three BPM“ Correlation Method
Gero Kube, DESY / MDI
DEELS 2019 Workshop @ ESRF, 3.- 5.6.2019
BPM1
BPM2
BPM3
beam
principle setup: 3 adjacent BPMs
connection via transport matrices
no non-linear elements between BPMs
BPM reads position information
difference: measured position vs. expectation
calculate variance (error propagation)
all BPM readings with same error
σΔ → N consecutive position measurements

## Slide 8: „Three BPM“ Correlation Method (2)
Gero Kube, DESY / MDI
procedure
N consecutive position measurements with 3 adjacent BPMs
determination of transfer matrix elements X21, X23
→ straight forward: calculation according to beam optics
→ model independent: Moore-Penrose pseudo inverse
(least-square estimate for X)
BPM resolution
→ evaluate formula
constant
offset ≈ 0
successive application to all BPMs
→ grouping three adjacent BPMs
example: KEK-B M. Arinaga et al., NIM A499 (2003) 100
restrictions
no non-linear elements → difficult at DLS
same error of BPM readings → not possible at PETRA III
sometimes weak correlations with neighbour BPMs
→ large uncertainty in BPM resolution (especially with Moore-Penrose)
phase advance ≠ n/2 π
n = ±1, 3, 5, …
DEELS 2019 Workshop @ ESRF, 4.6.2019

## Slide 9: Principal Components Analysis (PCA)
Gero Kube, DESY / MDI
method of multivariate statistics
conversion of set of correlated variables into set of linearly uncorrelated ones
→ principal components (PC)
cleansing of correlations in data sets
→ structuring of large data sets, compression, …
prinicipal axis determination
orientation of first principal axis
→ axis rotation such that overall data variance is maximized (requires centering of data)
→ alternative: minimize projections (hint: χ2)
orientation of second (third, …) principal axis
→ remove contribution of 1st PC from data
→ repeat rotation and variance maximization
condition: uncorrelated with (i.e., perpendicular to) first principal component
mathematics behind
form covariance matrix C → real & symmetric matrix
diagonalization of C → C = V Λ VT
V : formed by orthonormal eigen vectors
eigen vectors → principal components
eigen values (Λ) → amount of variance for PC
sort eigen vectors according to eigen values
DEELS 2019 Workshop @ ESRF, 4.6.2019

## Slide 10: Principal Components Analysis (2)
Gero Kube, DESY / MDI
alternative numerical method → Singular Value Decomposition (SVD)
DEELS 2019 Workshop @ ESRF, 4.6.2019
instead of diagonalization of covariance matrix
→ SVD of data matrix M
relation between singular and eigen values:
→
courtesy: Wikipedia
(M real: V* → VT)
advantage
→ SVD numerically more stable (formation of MMT can cause loss of precision → Läuchli matrix)
→ benefit: SVD provides additional information (accelerator physics)
orbit (space coordinate)
turns
(time coordinate)
application to BPM data
construction of BPM matrix M
→ BPM data centered
→ normalization:
C.X. Wang, SLAC-R-547 (2003)
exploration of SVD matrix properties
→ U: column vectors contain information about temporal pattern (tune, …)
→ V: column vectors contain information about spatial pattern (orbit / β function, …)
comment: U/V as temporal/spatial vectors → depends on orientation of matrix M

## Slide 11: PCA Example
Gero Kube, DESY / MDI
test measurement @ PETRA III
fill pattern: 960 bunches @ 5.6 mA
single vertical kick with excitation kicker
2048 turns recorded
DEELS 2019 Workshop @ ESRF, 4.6.2019
SVD analysis
dominant modes (singular values) → mode 1 & 2
temporal modes
information about tune
→ FFT of u1, u2
spatial modes
information about beam optics →
BPM resolution
singular values σ1, σ2, …= 0 → calculate cleaned orbit data

## Slide 12: PETRA III BPM System
Gero Kube, DESY / MDI
11 different pickup types:
Max von Laue hall
DEELS 2019 Workshop @ ESRF, 4.6.2019
I. Krouptchenkov et al., Proc. DIPAC 2009, Basel, TUPD03, p. 291
standard octants:
3 different cable types
RFA ½̋, ⅜̋ , ⅞̋ – 50 Ω
cable lengths: 10m … 200m
246 individual settings
normalize to Kx,y = 10 mm
RMSx,y = f(signal power)

## Slide 13: BPM Resolution Studies
Gero Kube, DESY / MDI
DEELS 2019 Workshop @ ESRF, 4.6.2019
resolution studies @ PETRA III
single bunch with Qb varying
Kx,y
normalization
single vertical kick with excitation kicker
2048 turns
plot as function
of signal power
combination of measurements for various Qb
PCA for
beam jitter
removement

## Slide 14: TbT Single Bunch Resolution
Gero Kube, DESY / MDI
machine studies at PETRA III with existing BPM system
Libera Brilliance (Instrumentation Technologies)
Libera Brilliance resolution not sufficient for PETRA IV
min(RMSx,y) ≈ 30 μm
PETRA IV CDR: rely on successor model → Libera Brilliance+
DEELS 2019 Workshop @ ESRF, 4.6.2019

## Slide 15: Test of Read-out Electronics
Gero Kube, DESY / MDI
beam test @ PETRA III: december 2018
BPM_SOL_24
10 dB Attenuator
Signal Combiner/Splitter
MACOM DS-409-4
Libera Brilliance+
Libera Brilliance+
(old)
Libera Brilliance
Libera Spark
Combiner / Splitter
Combiner / Splitter
Combiner / Splitter
Combiner / Splitter
~ 1 dB Attenuation
Signal Combiner/Splitter
MACOM DS-409-4
Libera Brilliance
→ BPM_SOL_24, in use at PETRA III
Libera Brilliance+
→ CDR: PETRA IV system
→ one system bought for DORIS / Olympus
→ one system recently bought
Libera Spark
→ new platform, no long-term stabilization
→ borrowed from I-Tech (thanks to Peter Leban)
BPM TbT resolution determination
orbit data contain contributions due to
→ correlated beam jitter
→ noise of BPM electronics
disentangle contributions
correlation analysis → does not work for single BPM
eliminate correlated jitter → sum & split orbit data
DEELS 2019 Workshop @ ESRF, 4.6.2019

## Slide 16: Resolution Comparison
Gero Kube, DESY / MDI
digital down conversion (DDC)
σ < 10 μm never achieved
σ = 19.8 μm @ IB = 1.549 mA (-29 dB)
→ correct for attenuation: IB = 0.38 mA (≈ -42 dB)
Libera Brilliance+
time domain processing (TDP)
σ = 10.3 μm @ IB = 1.323 mA (-30 dB)
→ correct for attenuation: IB = 0.38 mA
DDC mode < 20 μm (rms) → IB ≈ 0.4 mA
TDP mode ≈ 10 μm (rms) → IB ≈ 0.4 mA
would fulfil requirements
→ but Libera Spark is better…
DEELS 2019 Workshop @ ESRF, 4.6.2019

## Slide 17: Closed Orbit & Stability
Gero Kube, DESY / MDI
long term stability
user operation: 480 bunches in 100 mA, top-up
all Liberas in closed orbit (SA) mode → drift compensation (digital signal conditioning, DSC) on
Spark:
no DSC available
Brilliance+ without DSC
good climatization (hutches, racks) mandatory
DEELS 2019 Workshop @ ESRF, 4.6.2019

## Slide 18: Closed Orbit Resolution
Gero Kube, DESY / MDI
specification:
< 100nm (rms) in brightness mode (200mA in 1600 bunches) @ BW 300 Hz
first 12 hours (before beam dump):
σrms = 20.76 nm in SA mode (BW 4Hz, Kx,y = 10mm; BW 4Hz → see Brilliance+ User Manual, p.34)
→ scaling with band width:
× sqrt(300/4) = 8.66
→ σrms = 180 nm (@ BW 300Hz)
horizontal orbit
DEELS 2019 Workshop @ ESRF, 4.6.2019

## Slide 19: Conclusion
Gero Kube, DESY / MDI
Libera Spark
single bunch resolution much better → closed orbit: needs stabilization
Libera Brilliance+
single bunch resolution → specs fulfilled @ bunch curent IB ≈ 0.4 mA and monitor constant Kx,y = 10 mm
DDC mode < 20 μm (rms), TDP mode ≈ 10 μm (rms)
closed orbit → specs not fulfilled (< 100 nm @ 300 Hz bandwidth)
and Kx,y = 10 mm
σy,rms ≈ 180 nm @ 300 Hz bandwidth
DEELS 2019 Workshop @ ESRF, 4.6.2019
BPM resolution studies
require correlation analysis in order to disentangle noise and beam generated jitter
→ „three BPM“ correlation method & PCA
next step: Independent Component Analysis (ICA)
correlation analysis
„three BPM“ correlation method → not suitable for PETRA III BPM system
PCA → powerful tool, not only for resolution studies (e.g. BPM performance evaluation @ SSRF, …)
→ model independent, but limited by mode mixing
Z.-C. Chen et al., Chinese Physics 38 (2014) 077004
Nucl. Sci and Tech. 25 (2014) 020102
Libera Brilliance
single bunch resolution → specs not fulfilled for PETRA IV