## Slide 1: Constrains on Laser system implementation in the SPS tunnel
Gamma Factory Meeting
Faculty of Physics, Astronomy and Applied Computer Science
Jagiellonian University
Krakow, Poland
28–30 January 2019
By Valentin Fedosseev (CERN, EN department)

## Slide 2: Requirements for laser installation
Space for laser, optics, diagnostics, other electronic devices
Electrical power
Laser cooling => autonomous water circuit preferable
Temperature stability => typically better than +/- 1° C
Air purity => typically ISO 7 or better
Class 4 laser safety => Designated Laser Area with access controls + enclosed beams outside DLA (Class 1)
Specific requirements:
Synchronization to the machine timing
Integration into the CERN control system
General requirements:

## Slide 3: Interaction regions in SPS rings
Starights ections LSS4 and LSS6 were considered

## Slide 4: Straight section LSS4
30m
LSS.416
Passage should be left free for transport.
Practically no space for laser installation next to the beam line.

## Slide 5: Straight section LSS4
10m
Could it be possible to place the laser in neighbouring areas?

## Slide 6: Straight section LSS6
30m
LSS.616
Tunnel is larger

## Slide 7: Straight section LSS6
LSS.616
Space behind the beam line looks free

## Slide 8: SPS Radiation Survey Results
PCMx probe SAPHYMO
(plastic scintillator – 200cm3)
Distance ≈ 1m from beam line
All along the 7km of the SPS Ring (≈ 18000 measurements).
Labview registration application
Survey usually performed after 30 hours of cool-down
Measuring probe
Plastic scintillator Vol: 200cm3
  Measurements taken after 30 hours of cool-down:
      - 2017 Physics run (protons) from 8th May2017 to 23th Oct 2017
      - 2016 Physics run (protons) from 11th Mar 2016 to 14th Nov 2016
    Measurements taken after 2 months of cool-down - Jan 2018

## Slide 9: SPS Radiation Survey Results

## Slide 10: SPS Radiation Survey Results

## Slide 11: SPS Radiation Survey Results

## Slide 12: SPS Radiation Survey Results
For access after 30 h cool-down:
None of the two areas is particularly problematic.
However, both areas have some small areas with higher dose rates.

## Slide 13: Latest Radiation Survey made on 18.12.2018
LSS6 looks slightly better regarding the radiation emitted by beam line equipment

## Slide 14: Radiation on Electronics measurements
Measurements performed at the IR locations in LSS4 and LSS6
Period of measurements: 18/09/2018 – 17/01/2019
  Covers operation with protons and ions
  Beam operation stopped on 10.12 => in total ~ 1700 h of beam time
Measured values
  ΦHEH - High Energy Hadron fluence
  ΦThN - Thermal Neutron fluence
  TID - Total Ionizing Dose ( using passive silicon dosimeters)
Method to measure HEH using BatMon system:
The Single Event Upsets (SEUs) induced by radiation in SRAM memories biased at two different voltages are measured and converted in fluences by solving the system:
Performed by EN/SMM-RME (EDMS document 2080251)

## Slide 15: Radiation on Electronics measurements
Position of detectors in LSS4
Position of detectors in LSS6

## Slide 16: Radiation on Electronics measurements
Number of SEU detected by each BatMon pair
Summary of the fluences and the R factor measured
Long Straight Section | Position | Height | SEU 3V | SEU 5V
LSS4 | A | Beam | 20963 | 10498
 | B | Floor | 10351 | 1177
LSS6 | E | Beam | N/A | 17577
 | F | Floor | 20181 | N/A
Long Straight Section | Position | Height | R-factor | ΦHEH [pp/cm2] | ΦThN [pp/cm2]
LSS4 | A | Beam | 0.1 | 8.88 x 109 | 8.30 x 108
 | B | Floor | 1.9 | 8.80 x 108 | 1.64 x 109
LSS6 | E | Beam | 0.1* | 1.50 x 1010 | 1.40 x 109
 | F | Floor | 1.9* | 2.04 x 109 | 3.81 x 109

## Slide 17: Radiation on Electronics measurements
Summary of the measured TID for each position
Long Straight Section | Position | Height | TID [Gy]
LSS4 | A | Beam | 10.5
 | B | Floor | 4.8
LSS4 | C | Beam | 5.4
 | D | Floor | 3.9
LSS6 | E | Beam | 9.8
 | F | Floor | 4.0
Further data analysis and implications for electronics will be performed by EN/STI-BMI

## Slide 18: Comparison with AWAKE conditions
Simulation Parameters:
3x1011 p+/bunch
1 shot every 30 s
400 GeV
Experimental runs:
  4 weeks per run
  2 runs per year
Annual intensity: 4.8x1016
Fluka simulation by HSE-RP
Laser room
Streak camera
Digital cameras
Digital cameras
UV beam diagnostics
No visible influence of proton beam on the devices
SEUs regularly detected during proton beam time

## Slide 19: Synchronization to the machine timing
The integer harmonic of laser mode locking should be equal to an integer harmonic of the SPS RF frequency
  laser mode locking frequency should be specified accordingly
Stabilized optical link between laser and SPS beam control equipment
Phase stability conditions of the laser-ion interaction

## Slide 20: Synchronization to the machine timing
AWAKE example
H. Damerau et al., “RF Synchronization and Distribution for AWAKE at CERN” - in Proceeding of IPAC2106 [THPMY039]
(for 400 GeV protons)

## Slide 21: Laser beam control and diagnostics
At CERN the accelerator control system is based on the Real-time FESA (Front-End Software Architecture) framework. => FESA classes are required for laser beam elements
Motorization of laser mirrors, translators and other movable elements
Laser shutters and dumps
Laser power/energy meters
Acquisition of laser beam images

## Slide 22: Summary
Location for laser installation should fulfil certain requirements
  Space
  Air quality
  Temperature stability
  Safety (enclosure)
IR location at LSS6 offers some space near the beam line, where a laser installation could be placed,
The residual radiation conditions in both locations are acceptable for human work during the no-beam periods
Risks of laser electronics damage by HEH and thermal neutrons shall by analysed by experts. A qualified advice on possible shielding is expected
Laser synchronization is a separate task to work out.
Control of laser beam elements and diagnostics should be designed in accordance with CERN standards.

## Slide 23
ActiveFiberSystemsGmbH(AFS)islocatedinJena,
knownasthe“cityofphotonics”inGermany.Asa
spin-offfromtheFraunhoferIOFJenaandtheInsti-
tuteofAppliedPhysicsattheuniversityofJenaAFS
representstheexpertiseofinnovativesolid-state-laser
development.