## Slide 1
Development of SOI Monolithic Pixel Detector for Fine Measurement of Space and Time SOFIST
Project supported by
JSPS Grant-in-Aid for Scientific Research on Innovative Areas (2013-2017)
K. Hara (Univ. Tsukuba)
for SOIPIX Collaboration
SOFIST development:
U. Tsukuba: K. Hara, D. Yamauchi, R. Abe, S. Iwanami, H. Murayama, K. Watanabe
KEK: T. Tsuboyama, M. Yamada, S. Ono, Y. Arai, Y. Ikegami, I. Kurachi, M. Togawa, J. Haba
Tohoku U.: Li. Taohan

## Slide 2
OUTLINE
SOI PIXEL Sensor Development – Introduction
Highlights of Achievements
SOFIST
  design & developments
  SOFIST-1 position resolution
  SOFIST-2 timestamp resolution
  SOFIST-3 quick evaluation
SOFIST-4 3D stacking
Summary

## Slide 3
SOI pixel devices
SOI: Silicon-On-Insulator
Material budget
S/N
power dissipation
single event effects
latch up
Operation temp. (0.3K～570K)
speed
cost
Pixel size
CMOS circuitry fabricated on buried oxide(BOX)
LAPIS 0.2um FD-SOI
Features:
Monolithic (no metal bumps)
SOI-CMOS (FETs fully isolated)
Can choose* substrate of optimum resistivity (fully depleted CMOS sensors possible)
*SOITEC SmartCutTM
many excellent features
TID tolerance improved to 1MGy by introducing double SOI wafer (HSTD11)

## Slide 4: Grant-in-aid (fy2013-17) interdisciplinary research on quantum imaging opened with 3D semiconductor detector

## Slide 5: highlight-2 (INTPIX4)
Fe-55
Integration &global shutter
10.2x15.4mm

## Slide 6: highlight-2 (INTPIX4)

## Slide 7
3mm
INTPIX4: Computed Tomography with Syncrotron X-ray
(by R. Nishimura, K. Hirano (KEK)
Notes: 煮干しを0〜180度まで回転させたときの
それぞれの角度での積分画像を1度刻みでを取得しています。
全181枚のデータをRA=255の軸を中心に
1度ずつずらして配置した形に再構成したものが
お送りした動画になります。
再構成時にスケールの修正は特に行っていないとのことですので、
再構成後の画像の1ピクセル＝1um角となります。
なお、
波紋状のもの
については背景減算後に残ったビームプロファイルの痕跡によるものと考えられ
ます。
(すべての画像でほぼ同じ位置に痕跡が残っているので回転像では円形に映りま
す)

## Slide 8: highlight-1 (SOIPIX-PDD)
Fe-55

## Slide 9: highlight-1 (SOIPIX-PDD)
Fe-55

## Slide 10: highlight-3 (FPIX2)
U Tsukuba/KEK
mm
mm

## Slide 11: highlight-3 (FPIX2)
U Tsukuba/KEK
1.9

## Slide 12: highlight-4 (XRPIX5)

## Slide 13: highlight-4 (XRPIX5)

## Slide 14: highlight-4 (XRPIX5)
Energy resolution as good as CCD’s
Much faster!
Muon veto possible

## Slide 15: sofist SOI Monolithic Pixel Detector for Fine Measurement of Space and Time
Allows time and space point measurement
Design specified to meet the requirements for the ILC experiment
Main requirements:
Position resolution <3um
Low material <0.2%Xo/layer
TID~1kGy/year
NIEL ~1011neq/cm2/year
Readout between trains
bunch ID highly preferred
Power dissipation: <50mW/cm2
The goal is a sensor with both good spatial resolution of <3um and time resolution <1us.
pixel size ~20um-sq with analog readout
timestamp recording the ramping voltage
store the data allowing multiple hits on the same pixel
on-chip ADCs &zero-suppression for fast data transfer

## Slide 16: Sofist pixel architecture
Main requirements:
Position resolution <3um
Low material <0.2%Xo/layer
TID~1kGy/year
NIEL ~1011neq/cm2/year
Readout between trains
bunch ID highly preferred
Start Vramp per beam train
Signals exceeding Vth
shift-register chooses* memory cell to record
analog signal recorded
timestamp voltage recorded
Recorded memories are digitized by column ADCs (8bit has been implemented)
Zero-suppression applicable

## Slide 17: SOFIST – in development
FNAL
Beamtested
(only analog)
FNAL
Beamtested
(both but in separate pixels)
delivered
Under 3D
Under evaluation (both in same pixel)

## Slide 18: soFist-1 spatial resolution
1.37 ± 0.04 mm
1.50 ± 0.05 mm
1.38 ± 0.05 mm
1.35 ± 0.04 mm
Residual X
Residual Y
SOFIST#1(BPW14x14)
SOFIST#2(BPW16x16)
SOFIST residual to FPIX track (s_track~0.57/0.65mm)
Bias=130V (~500um depletion) =>15V (~200mm depletion)
Readout: external 12-b ADCs =>on-chip 8-b ADCs
1.49 ± 0.06 mm
1.57 ± 0.08 mm
1.55 ± 0.08 mm
1.38 ± 0.05 mm
1.54 ± 0.04 mm
1.58 ± 0.05 mm
1.33 ± 0.03 mm
1.32 ± 0.03 mm
plots for “black case”
S/N~300 (130V)
~120 (15V)
Ono (KEK)

## Slide 19: soFist-1 spatial resolution
Compatible with DEFFET ([22])
Track uncertainty (0.6um) subtracted
~200um
~500um
depletion
η correction may improve further
Ono (KEK)
~400um
1000
500

## Slide 20: soFist-2 beam test 2018mar
(2mm□)
INTPIX4
RoI(2,3,,mm□)
adjustable
scint.+MPPC
event tag
120GeV p
BUSY
SOFIST
(17um□, 8.7x14.1mm)
Beamline Setup
1ms
Vramp
spill
STORE
4.2s
500ms
BUSY
SOFIST
event tag
RO
RO
space&time
INTPIX4
RO
RO
space
timestamp
Timing Chart
RO with external 12-bit ADCs

## Slide 21: soFist-2 timestamp calibration
Ttest=250us
linear correlation
assumed
1ms
All pixels
Ono (KEK)
Ono (KEK)
1V
conducted on site
conducted on site

## Slide 22: soFist-2 timestamp correlation
Correlation of timestamps in two SOFISTs
Overlapping events in same 500us gate
Periodic clusters?
Ono (KEK)
Ono (KEK)

## Slide 23: Size of FNAL Main injector (MI)
Beam structure due to bunch rotation in the MI: 11us→3.3km (3319m actual)
Yamada (KEK)
Yamada (KEK)

## Slide 24: SOFIST-2 timestamp precision
Time spread for 500us delay pulses
conducted on testbench
resolution
~0.75us(~2.9ADC)
No notable position dependence
(0,0)
linearity
ADC counts
Test pulse delay [us]
Murayama (Tsukuba)
ΔADC/Δdelay[us]
Slight row dependence

## Slide 25: SOFIST-2 analog signal
25um
p
BNW
8
67
Cluster charge
(HV dependence)
-5V
-40V
-20V
Yamada (KEK)
Yamada (KEK)

## Slide 26: Sofist-2 full depletion
Yamada (KEK)
S>700 ADC (N~5 ADC) for 67um thickness
Full depletion expected ~30V for 9x1012/cm3 (0.5kΩcm) concentration (measured)

## Slide 27: Sofist-2 cluster charge
Mostly one-cluster events
Vth ~1/2 mip
Seed>10σ
3x3 cluster
peak~300ADC: X-talk at Comparator ON
events
Seed>3σ
Vth ~1/4 mip should be possible…
Yamada (KEK)
Yamada (KEK)

## Slide 28
Tracker for SOFIST testbeam
Using 4 plaes of large-area INTPIX4
Pixel size : 17x17um
Sensitive area (8.7x14.1mm)
64x512
64x512
64x512
512x832 pixels
-> 300 fps
faster by RoI selection
13 parallel readout
SOI Testbeam@FNAL
(2018 Feb/Mar)
Hit residuals of 2nd to tracks (1st,3rd,4th)
σ=1.77um
σint=1.37um
（山内, 修論)

## Slide 29: SOFIST2 spatial resolution
Due to a high Vth applied in the testbeam, hits are with ~single cluster ⇒25um/√12~7.2um expected
7.03 6.73 7.50
7.03 7.30 7.16

## Slide 30: SOFIST-3 – quick test
S. Ono(KEK)
Both functions implemented on 30um pixels
Under testbeam @ FNAL

## Slide 31: SOFIST-3 – pixel layout
Yamada (KEK)
30um
analog
SF
..
..
MIM(50fF, 1.5fF/um2)
1 for CDS
2 for comparator
X2 for timestamp

## Slide 32: SOFIST Ver.4
Lower pixel
Upper pixel
Up to comparator
Shift-registers
3 Memory cells each

## Slide 33: SOFIST-4: pixel layout
Lower pixel
Pre-amp
Comparator
Upper pixel
Shift-register
Analog memory x3
Timestamp x3
20um
Cone-Bump pad(5x5um)
20um
Au cone bump
Reported LAST YEAR
3 bumps/pixel

## Slide 34: 3D is coming (SOFIST-4)
Cylinder-type is more reliable
Upper&lower bumps are stacked well
(peel off one sensor layer)
slight mis-alignment to UBM ⇒corrected
Bonding<200degC
T-Micro

## Slide 35: Remaining processes (SOFIST-4)
(e) Pad patterning
and passivation
Back gate adjust electrode
Passivation
Bond Pad
High R-Si
etching
(d) Bulk-Si removal
High R-Si
8μm
T-Micro
glue injection

## Slide 36: summary
5-year Grant-in-Aid (2013-2017) accelerated development/application of various SOI devices.
SOFIST has been tested in FNAL beam
SOFISTv1 <20um pixel size>
spatial resolution:
1.3 um(500um)-1.4 um (200um) with no h-correction applied
SOFISTv2 <25um pixel size>
time resolution: 1.89us (linear calibration)
good S/N~120 for 67um thickness
cluster size~1 (Vth was set slightly high)
- spatial resolution ~8um as expected from pixel si
SOFISTv3 is beam tested
SOFISTv4 fabrication is underway

## Slide 37: spares

## Slide 38: International Linear Collider
ILC Experiment
  e+e− linear collider
  Center of mass energy: 250 - 500 GeV
  Precise measurement of the Higgs boson
  Search for beyond the Standard Model
ILC detector concept (ILD)
Vertex detector system

## Slide 39: Double SOI wafer
Courtesy of Lapis semiconductor

## Slide 40: Issues to overcome
p-
-
BNW
E-field
+++++++++++
Tied to pixel node conventionally
-> area need to be optimized
( coverage of FETs/capacitance )
DSOI solves all
In application of SOI to pixel devices …..

## Slide 41: Wafer types
Lapis Semiconductor: 0.20um FD-SOI process
On-going MPW
Optimum wafer selectable (single p/n, various resisitivities):
One type for double: getting better!