## Slide 1
Kaonic Deuterium X-Ray Measurements with the SIDDHARTA-2 Apparatus at DAΦNE
Marlene Tüchler
(on behalf of the SIDDHARTA-2 collaboration)
Joint Meeting of the APS and SPS, Zürich
August 27th, 2019

## Slide 2: OUTLINE
Motivation: Kaonic Atoms
The SIDDHARTA-2 Apparatus
  Target Cell
  X-Ray Detection System
  Veto System
Conclusion
Outlook
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 3: Motivation: Kaonic Atoms
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 4: FORMATION OF KAONIC ATOMS
n=1
p
e-
K-
X-ray
K- slows down
Cascade:
Auger electrons, X-rays
Ground state
n=25
Captured by atom
Absorption
by nucleus
n=2
n=1
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 5: MOTIVATION: KAONIC ATOMS
Direct study of low-energy QCD with strangeness
Strong interaction: energy shift ε1s and width Γ1s of ground state - directly observable
ΔE
ΔE = Emeasured - Eem
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 6: MOTIVATION: KAONIC ATOMS
Direct study of low-energy QCD with strangeness
Strong interaction: energy shift ε1s and width Γ1s of ground state - directly observable
SIDDHARTA:
  Precise experimental data for kaonic hydrogen
SIDDHARTA-2:
  Kaonic deuterium measurement needed to extract isospin-dependent scattering lengths , - chiral symmetry breaking
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 7: SCATTERING LENGTHS
Improved Deser-Trueman:
Isospin-dependent scattering lengths ,
𝛂 … fine structure constant
µC … reduced mass of K-p system
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 8: The SIDDHARTA-2 Apparatus
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 9: DAΦNE COLLIDER AT LNF
e+-e-
collider
LINAC
Accu.
operates at the centre-of-mass
energy of the  meson
mass m = 1019.413 ± 0.008 MeV
width  = 4.43 ± 0.06 MeV
 produced via e+e- collision
(e+e- → ) ~ 5 µb
→ monochromatic kaon beam
(127 MeV/c)
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 10: MOTIVATION
Kaonic deuterium:
  Yield lower than K-p yield
  Width approx. 2× K-p width
  Large area X-ray detectors with good energy resolution
  Lightweight cryogenic gaseous target
  Veto system:
  Veto-1: Timing information
  Veto-2: Timing & spatial correlation of SDD-hit and Veto-2 hit
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 11: DETECTOR SYSTEM SIDDHARTA-2
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 12: TARGET CELL
M. Tüchler – APS & SPS Meeting, Zürich 2019
Lightweight - filled with gaseous deuterium
Kapton sidewalls
Operated at cryogenic temperatures (30 K at 0.3 MPa working pressure)
High X-ray transmission
Sufficiently high gas density

## Slide 13: X-RAY DETECTORS - SDDs
M. Tüchler – APS & SPS Meeting, Zürich 2019
SIDDHARTA
  JFET implemented on anode
  Active/total area ~ 25%
  Drift times ~ 800 ns
SIDDHARTA-2
  CUBE preamplifier on ceramic
  Active/total area ~ 75%
  Drift times ~ 300 ns

## Slide 14: X-RAY DETECTORS - SDDs
M. Tüchler – APS & SPS Meeting, Zürich 2019
SIDDHARTA-2
  48 arrays of 8 cells
  Each cell 8 × 8 mm2 active area
  Energy resolution < 180 eV at 6 keV
55Fe spectrum
123.0 eV FWHM
SDD characteristics:
area/cell = 64 mm2
total area = 512 mm2
T = - 100°C
drift time < 500 ns
Counts
Energy (keV)

## Slide 15: VETO-1 SYSTEM
12 detector units
Scintillator and light guide with photomultiplier read-out
Particle hit-position independent
M. Tüchler – APS & SPS Meeting, Zürich 2019
Distinction between “real” and “false” kaon stops using timing information
Time resolution of < 500 ps required

## Slide 16: VETO-2 SYSTEM
π
large signal
SDD
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 17: VETO-2 SYSTEM
π
π
large signal
signal in region of interest
SDD
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 18: VETO-2 SYSTEM
π
π
π
large signal
signal in region of interest
delta-ray (knock-on e-)
SDD
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 19: VETO-2 SYSTEM
π
π
π
π
large signal
signal in region of interest
delta-ray (knock-on e-)
secondary X-ray or e-
from setup
SDD
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 20: VETO-2 SYSTEM
Detector Setup:
  96 Silicon photomultipliers
  96 plastic scintillators
  3D-printed ABS boxes
Characterisation:
  Time resolution (293 ± 45) ps
  Detection efficiency (99.1 ± 0.2)%
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 21: CONCLUSION & OUTLOOK
Kaonic atom studies at LNFrascati:
  Installation and commissioning at Frascati ongoing
  SIDDHARTINO: kaonic hydrogen and helium, background studies
  SIDDHARTA-2: Kaonic deuterium 2020
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 22: Attachment
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 23: FORMATION OF KAONIC ATOMS
Properties:
  mK- : 493.677 MeV
  Bohr energy levels and Bohr orbit:
  Particle captured in highly excited state:
α… fine structure constant
μ…. reduced mass of moving particle
n… principal quantum number
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 24: SCATTERING LENGTHS
Theoretical Predictions:
  Expected values for kaonic deuterium
ε1s (eV) | Γ1s (eV) | (fm) | Reference
 | 1016 |  | Weise et al. (2017)
 | 757 |  | Mizutani et al. (2013)
 | 1011 |  | Shevchenko (2012)
 | 650 |  | Meißner et al. (2011)
 | 674 |  | Gal (2007)
ε1s (eV) | Γ1s (eV) |  | Reference
 | 1016 |  | Weise et al. (2017)
 | 757 |  | Mizutani et al. (2013)
 | 1011 |  | Shevchenko (2012)
 | 650 |  | Meißner et al. (2011)
 | 674 |  | Gal (2007)
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 25: TIMING VETO-1
Time response of detector
M. Tüchler – APS & SPS Meeting, Zürich 2019
Green: Events after detection of K+ in bottom scintillator of kaon monitor
Red: Events after detection of K- in bottom scintillator of kaon monitor
Blue: Neither K- nor K+ detected in bottom kaon monitor
Black: Sum of all events
Time (ns)
Counts
Direct stops in wall material
Products of K- absorption in gas

## Slide 26: DETECTOR SETUP - SiPMs
  Array of Geiger-mode avalanche photodiodes (G-APDs)
  Geiger-mode: high S/N-ratio
  Active quenching resistors
Hamamatsu
microcell structure
M. Tüchler – APS & SPS Meeting, Zürich 2019
Notes: https://www.ketek.net/sipm/technology/
https://www.ketek.net/sipm/technology/microcell-construction/
R_quench = 200 – 800 kOhm

## Slide 27: DETECTOR SETUP - SiPMs
SiPM vs PMT:
 | SiPM | PMT
Supply Voltage | 20 – 100 V | 1000 - 3000 V
Gain | 105 - 107 | 104 - 109
Quantum Efficiency | max. 80% | 25% - 40%
Rise Time | < 1 ns | 0.5 – 3.5 ns
Sensitive to B-fields | no | yes
M. Tüchler – APS & SPS Meeting, Zürich 2019
Notes: https://www.ketek.net/sipm/technology/
https://www.hamamatsu.com/resources/pdf/etd/PMT_handbook_v3aE.pdf

## Slide 28: DETECTOR SETUP - SiPMs
M. Tüchler – APS & SPS Meeting, Zürich 2019
Notes: https://www.sensl.com/downloads/ds/TN%20-%20Intro%20to%20SPM%20Tech.pdf

## Slide 29: VETO-2 DETECTOR SETUP
SiPMs
  FBK-AdvanSiD (Trento)
  Near-UV operation (420 nm)
  Active area 4 x 4 mm2
Scintillators
  EJ-200 plastic scintillators
  50 x 12 x 4 mm3
  Wavelength of maximum emission 425 nm
M. Tüchler – APS & SPS Meeting, Zürich 2019
Notes: Hier lassen oder in den Anhang?

## Slide 30: VETO-2 DETECTOR SETUP
Amplification Boards
  12 channels with analogue (DS) and digital output (LVDS)
  Digital: Time over Threshold
  Gain and threshold: remotely adjustable
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 31: AMPLIFIER BOARDS
Sauerzopf C. et al. (2016) Nucl. Instrum. Meth. Phys. Res. A 819 163–166
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 32: VETO-2 CHARACTERISATION
Time Resolution
  Diode laser (450 nm) on three or five positions
  Time delay between laser pulse – SiPM signal
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 33: VETO-2 CHARACTERISATION
Time Resolution
  No systematic trend in position dependency
  Average local time resolution (FWHM): (54 ± 2) ps
  Average time resolution of Veto-2 system (FWHM): (293 ± 45) ps
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 34: VETO-2 CHARACTERISATION
Detection Efficiency
  Cosmic rays
  Stack of four scintillator-SiPM tiles
  Pulse height spectra
  Coincidence between outer detectors
SiPM
Scintillator
Cosmic Ray
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 35: VETO-2 CHARACTERISATION
Detection Efficiency
  Detection efficiency of (99.1 ± 0.2)%
Cosmic data + MC simulation
MC simulation: K-N decay products
M. Tüchler – APS & SPS Meeting, Zürich 2019
Notes: Muons normally distributed between 150 & 700 MeV/c bzw. Σ±π ∓, Σ0π 0 , Λπ−, Σ0π− and Σ−π0  n, p, π, γ

## Slide 36: VETO-2 CHARACTERISATION
Temperature studies of Detectors
  SiPMs in climate chamber with Na-22 source
  Amplification board outside at constant temperature
  Pulse height spectra from T = -40°C to +40°C
Compton edge
M. Tüchler – APS & SPS Meeting, Zürich 2019
Notes: Na-22: 511 keV, 1275 keV γ

## Slide 37: VETO-2 CHARACTERISATION
Temperature studies of Detectors
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 38: VETO-2 CHARACTERISATION
Temperature studies of Detectors
M. Tüchler – APS & SPS Meeting, Zürich 2019
Notes: 0.5 V difference between -40 and 20 degrees

## Slide 39: VETO-2 CHARACTERISATION
Detection Efficiency
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 40: VETO-2 CHARACTERISATION
Detection Efficiency
  Multiplicity: (98.8 ± 0.2)% for four detectors
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 41: VETO-2 CHARACTERISATION
Temperature studies of Detectors
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 42: VETO-2 CHARACTERISATION
Temperature studies of Detectors
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 43: VETO-2 CHARACTERISATION
Temperature studies of Detectors
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 44: VETO-2 CHARACTERISATION
Temperature studies of Detectors
AdvanSiD. NUV SiPMs Datasheet (Rev.9). www.advansid.com/products/product-detail/asd-rgb-nuv-4s-p (Accessed 12.03.2019)
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 45: VETO-2 CHARACTERISATION
Temperature studies of Amplification Boards
  Amplification board in climate chamber
  SiPM with Na-22 source outside at constant temperature
  Pulse height spectra at T = (10, 20, 30, 40)°C
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 46: VETO-2 CHARACTERISATION
Temperature studies of Amplification Boards
Comparison Position of First Compton Edge
Comparison Position of Second Compton Edge
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 47: VETO-2 CHARACTERISATION
Correlation of ToT and Amplitude
M. Tüchler – APS & SPS Meeting, Zürich 2019

## Slide 48: REFERENCES
Griffiths, D. J.: Introdution to elementary particles. Weinheim, Wiley-VCH, 2008
Backenstoss, G.: Pionic Atoms. Annual Review of Nuclear Science, 1970, 201(1), 467-508
Gotta, D. et al.: Pionic Hydrogen. Physics Procedia, 2011, 17, 69-76
Meißner U.-G., Raha U., Rusetsky A.: Spectrum and Decays of Kaonic Hydrogen. Eur. phys. J. C, 2004, 35, 349
Bazzi, M. et al.: Preliminary study of kaonic deuterium X-rays by the SIDDHARTA experiment at DAFNE. Nuclear Physics A, 2013, 907, 69-77
www.ketek.net/sipm/technology/working-principle/ (15.08.2018)
www.ketek.net/sipm/technology/device-parameters/ (15.08.2018)
M. Tüchler – DK-PI, Vienna 2019